## CALL_REPORTING_DISTRICT_WISE
Sno | District | Hospital | Asset_No | Description | U_ReqName | Contact_No | WorkOrder_No | Call_recd_Date | Call_Recd_Time | Call_Responded_date | Call_Responded_Time | Call_Completed_Date | Call_Completed_Time | Problem_Reported | Action_Taken | Status | Serial_No | Criticality | downtime | Aging_Days
1 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | DORTHI SCOTT | 9764971169 | 5909 | 04-Aug-2017 | 12:04 |  |  | 05-Aug-2017 | 15:00 | NOT WORKING | REPLACED NEW INLET FILTER, NOW FILTRATION PROCESS DONE PROPERLY, MACHINE IS WORKING CONDITION | Closed | 2SXAE669 | HIGH | 26:56:0 | 2
2 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | STOTT NURSE | 9764971169 | 14683 | 07-Apr-2018 | 09:15 |  |  | 09-Apr-2018 | 09:59 | NOT WORKING | REPLACED NEW INLET FILTER,NOW MACHINE IS WORKING CONDITION. | Closed | 2SXAE669 | HIGH | 48:44:0 | 3
3 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | SARIKA VASNIK | 8446603844 | 21964 | 11-Nov-2018 | 08:40 |  |  | 15-Nov-2018 | 10:53 | NOT WORKING | CHECK AND FOUND CONDUCTIVITY AND FLOW ISSUE IS PRESENT. CLEANED THE A AND PART CHAMBER AND CONDUCTIVITY PROBLEM IS RESOLVED. DIASAFE CLIP IS FIXED INSTEAD OF DIASAFE AND FLOW PROBLEM IS RESOLVED. NOW DIALYSIS MACHINE IS WORKING. | Closed | 2SXAE669 | HIGH | 98:13:0 | 5
4 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 39319 | 11-Feb-2020 | 08:55 |  |  | 11-Feb-2020 | 21:39 | NOT WORKING | CHECK AND FOUND DURING CALIBRATION OF FLOW AT 800,FLOW FLUCTUATES AND WATER ALARM IS SHOWN. CLEANED THE WATER INLET VALVE AND ALL THE FILTERS. CARBON BRUSH IS CHECKED AND DEGASSING AND FLOW PUMP IS CLEANED. NOW, CONDUCTIVITY DOESN'T FLUCTUATES AND FLOW ALARM IS ALSO SORTED OUT. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 12:44:0 | 1
5 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | Hema Borjar | 9767558503 | 41384 | 31-Mar-2020 | 08:10 |  |  | 31-Mar-2020 | 17:29 | NOT WORKING | CHECK AND FOUND WATER ALARM IS SHOWN DURING TEST AND CONDUCTIVITY TEST COULD NOT BE PASSED. CLEANED THE INLET WATER VALVE, FLOAT SWITCH AND ALL INTERNAL FILTERS. FLOW AND DEGASSING PUMP IS OPENED UP AND CLEANED. CALIBRATION OF FLOW IS BEING DONE.GEAR HEAD IS BEING INTERCHANGED. PROBLEM IS RESOLVED. IN FUTURE GEAR HEAD NEED TO BE CHANGED AND MR HAS BEEN SENT. AT PRESENT DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 9:19:0 | 1
6 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | Hema Borjhar: | 9767558503 | 42327 | 17-May-2020 | 09:00 |  |  | 17-May-2020 | 11:37 | NOT WORKING | CHECK AND FOUND FLOW ALARM IS DISPLAYED WHILE MACHINE IS IN TEST AND FLOW 800 & 500 IS VERY HIGH WHEN CHECKED IS CALIBRATION MODE. CLEANED ALL THE INLET FILTERS. FRONT CLEANING AND HOT DISINFECTION IS BEING GIVEN. CALIBRATION OF FLOW AND BLOOD LEAK IS BEING DONE. PROBLEM OF FLOW IS BEING RESOLVED. MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 2:37:0 | 1
7 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA GURJAR | 9767558503 | 7696 | 02-Oct-2017 | 08:25 |  |  | 04-Oct-2017 | 11:42 | NOT WORKING | CLEANED ALL MICRON FILTERS AND REPLACED NEW INLET FILTER THEN DONE CALIBRATION PROCESS, NOW MACHINE IS WORKING CONDITION. | Closed | 2SXAE669 | HIGH | 51:17:0 | 3
8 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | Hema Borjhar: | 9767558503 | 43119 | 15-Jun-2020 | 08:10 |  |  | 15-Jun-2020 | 16:30 | NOT WORKING | PROBLEM REPORTED IS OF FLOW ALARM DURING DIALYSIS OF PATIENT. FIXED THE NEW TUBING ACROSS VALVE NUMBER 33 & 35 OF BALANCING CHAMBER AS ITS WAS HAVING LEAKAGE. CLEANED ALL THE FILTERS. CALIBRATION OF FLOW IS BEING DONE. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 8:20:0 | 1
9 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA SISTER | 9767558503 | 21212 | 10-Oct-2018 | 08:17 |  |  | 28-Oct-2018 | 12:46 | NOT WORKING | CHECK AND FOUND OPTICAL DETECTOR TEST FAILS.ULTRASONIC SENSOR IS ALSO NOT WORKING. CHANGED THE NEW LP 450M BOARD AND ULTRASONIC SENSOR. ALL TEST IS PASSED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 436:29:0 | 19
10 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | SUNIL | 9158418368 | 24895 | 11-Feb-2019 | 09:09 |  |  | 12-Feb-2019 | 13:56 | NOT WORKING | CHECK AND FOUND MACHINE BLOOD LEAK,NEED TO CALIBRATION. CALIBRATION IS DONE NOW MACHINE IS WORKING | Closed | 2SXAE669 | HIGH | 28:47:0 | 2
11 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | Hema Borjar | 9767558503 | 41600 | 11-Apr-2020 | 09:30 |  |  | 15-Apr-2020 | 17:23 | NOT WORKING | CHECK AND FOUND WATER ALARM ISSUE AND FLUCTUATIONS IN FLOW WITH FAILURE OF CONDUCTIVITY TEST. FIXED THE NEW DIAPHRAGM OF BALANCING CHAMBER. HEATER BLOCK IS BEING CLEANED ALONG WITH FLOAT SWITCH. NEW TMP SENSOR IS BEING FIXED. CALIBRATION OF FLOW IS BEING DONE. PROBLEM IS RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 103:53:0 | 5
12 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 35537 | 21-Nov-2019 | 09:08 |  |  | 22-Nov-2019 | 17:30 | NOT WORKING | CHECK AND FOUND HIGH TMP(TRANS-MEMBRANE PRESSURE) ALARM SHOWS UP DURING DIALYSIS OF PATIENT. CALIBRATION OF TMP IS DONE. NOW, NO 'HIGH TMP' ALARM SHOWS UP DURING DIALYSIS OF PATIENT. NOW, DIALYSIS IS DONE WITHOUT ANY HINDRANCE. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 32:22:0 | 2
13 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | SUNIL | 9158418368 | 18871 | 06-Aug-2018 | 08:55 |  |  | 06-Aug-2018 | 18:30 | NOT WORKING | CHECK AND FOUND MACHINE SHOWING AIR DIRECTOR ALARM CONTINUES. REPLACED ULTRASONIC SENSOR THE SAME. NOW MACHINE IS IN WORKING CONDITION. | Closed | 2SXAE669 | HIGH | 9:35:0 | 1
14 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | Akshay | 8888591294 | 42101 | 08-May-2020 | 08:00 |  |  | 11-May-2020 | 13:27 | NOT WORKING | CHECK AND FOUND HEATER BLOCK IS BURNT AND HAVING LEAKAGE THROUGH WHICH WATER IS SPILLING OUT. FIXED THE NEW HEATER BLOCK. CALIBRATION OF FLOW IS BEING DONE ALONG WITH DISINFECTION PROCESS. PROBLEM IS RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 77:27:0 | 4
15 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA BHORJKAR | 9767558503 | 11230 | 03-Jan-2018 | 08:10 |  |  | 05-Jan-2018 | 19:57 | NOT WORKING | CLEAN THE DUST FROM MOTOR AND FILTER | Closed | 2SXAE669 | HIGH | 59:47:0 | 3
16 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 26385 | 29-Mar-2019 | 08:45 |  |  | 03-Apr-2019 | 11:41 | NOT WORKING | CHECK AND FOUND MACHINE NIBP CUFF IS DEFECTIVE. NEED TO CHANGE . CHANGE THE NIBP CUFF NOW MACHINE IS WORKING GOOD 9767558503 | Closed | 2SXAE669 | HIGH | 122:56:0 | 6
17 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA BORJAR | 9767558503 | 37285 | 30-Dec-2019 | 09:10 |  |  | 31-Dec-2019 | 10:09 | NOT WORKING | CHECK AND FOUND WATER ALARM IS ALWAYS COMING AND TEST COULD NOT BE PASSED. CLEANED THE HEATER CHAMBER, INLET VALVES, A & B PART SUCTION TUBES, UF PUMP TUBES AND DIASAFE TUBES. NOW, WATER ALARM ISSUE IS BEING SOLVED. CALIBRATION OF BLOOD LEAK, 500 & 800 FLOW IS DONE. NOW, DIALYSIS MACHINE IS WORKING. | Closed | 2SXAE669 | HIGH | 24:59:0 | 2
18 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | NILESH | 9860704598 | 3269 | 21-Apr-2017 | 09:30 | 26-Apr-2017 | 18:47 | 26-Apr-2017 | 14:36 | NOT WORKING | CLEANED PART B MOTOR SECTION AND CALIBRATE THE MACHINE | Closed | 2SXAE669 | HIGH | 125:6:0 | 6
19 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | SARIKA WASNIK | 8446603844 | 16602 | 12-Jun-2018 | 08:15 |  |  | 12-Jun-2018 | 14:48 | NOT WORKING | CLEANED PART A AND B FILTER ALSO MOTOR SPRING | Closed | 2SXAE669 | HIGH | 6:33:0 | 1
20 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | BHARATI | 9764971169 | 6828 | 08-Sep-2017 | 08:56 |  |  | 12-Sep-2017 | 15:19 | NOT WORKING | CALIBRATION PROCESS DONE PROPERLY ON BLOOD LEAK,FLOW FUNCTIONS . MACHINE IS WORKING CONDITION | Closed | 2SXAE669 | HIGH | 102:23:0 | 5
21 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | SHUBHANGI NURSE | 7798024792 | 20995 | 02-Oct-2018 | 08:33 |  |  | 03-Oct-2018 | 14:05 | NOT WORKING | CHECK AND FOUND FLOW ALARM IS COMING. DIASAFE ISSUE IS BEING DETECTED. CHANGED THE NEW DIASAFE. FRONT CLEANING IS DONE. BLOOD LEAK PROBLEM IS ALSO BEING SOLVED. DIALYSIS MACHINE IS FULLY WORKING. | Closed | 2SXAE669 | HIGH | 29:32:0 | 2
22 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 35308 | 18-Nov-2019 | 08:58 |  |  | 18-Nov-2019 | 15:07 | NOT WORKING | CHECK AND FOUND 'FLOW ALARM' ISSUE. ERROR OF 147 IS DISPLAYED DURING CALIBRATION OF 500 & 800.PROBLEM FOUND IN DIALYSATE COUPLING PATH. BLOCKAGE IS BEING REMOVED FROM IT. ALL THE VALVES WERE CHECKED I.E 24,24B,30,31,33,34,35,36,37,38,87 IN DIAGNOSTIC MODE AND VALVES ARE FUNCTIONAL WITH NO BLOCKAGE. FLOW PUMP CARBON BRUSH IS ALSO CHANGED. BASIC CALIBRATION OF FLOW IS DONE. 'FLOW ALARM' ERROR IS REMOVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 6:9:0 | 1
23 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 4828 | 27-Jun-2017 | 09:55 |  |  | 01-Jul-2017 | 12:20 | NOT WORKING | REPLACED NEW UF MOTOR AND NEW ULTRASONIC SENSOR ,NOW ALL TEST GET PASSED AND MACHINE IS WORKING CONDITION | Closed | 2SXAE669 | HIGH | 98:25:0 | 5
24 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 11444 | 09-Jan-2018 | 09:20 |  |  | 10-Jan-2018 | 11:39 | NOT WORKING | CALIBRATION HAS DONE MACHINE IS WORKING CONDITION | Closed | 2SXAE669 | HIGH | 26:19:0 | 2
25 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | AKSHAY SAVARKAR | 8888591294 | 20183 | 08-Sep-2018 | 08:45 |  |  | 11-Sep-2018 | 12:04 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING | Closed | 2SXAE669 | HIGH | 75:19:0 | 4
26 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | SEEMA BORJAD | 9767558503 | 27910 | 25-May-2019 | 09:02 |  |  | 29-May-2019 | 09:28 | NOT WORKING | CHECK AND FOUND PROBLEM FOUND IN SPRING SET OF BIC PUMP DUE TO WHICH CONDUCTIVITY PROBLEM ALWAYS USE TO ARISE. CHANGE THE OLD DEFECTIVE SPRING SET WITH THE NEW ONE. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 96:26:0 | 5
27 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | Hema Borjhar | 9767558503 | 43554 | 27-Jun-2020 | 08:13 |  |  | 27-Jun-2020 | 11:30 | NOT WORKING | PROBLEM REPORTED OF ULTRAFILTRATION (UF) ERROR AS WASTE WATER EXCRETED IS LESS AS SET UF VOLUME. CLEANED ALL THE INLET FILTERS. FRESH PATH AND DRAIN PATH IS BEING CHECKED AND FOUND TO BE OK. CHECKED THE PULSE OF UF PUMP FOR 60 STROKE IN 60 SECONDS AND ITS OK. CALIBRATION OF FLOW, UF, BLOOD LEAK & TMP IS BEING DONE. DISINFECTION IS BEING GIVEN. DONE THE DUMMY DIALYSIS AND FOUND EXACT ML OF WATER IS EXCRETED AS SET ML OF UF VOLUME. DIALYSIS OF PATIENT IS BEING DONE WITH NO ISSUES. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 3:17:0 | 1
28 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | MISS SCOTTE | 9764971169 | 18490 | 27-Jul-2018 | 08:05 |  |  | 30-Jul-2018 | 14:30 | NOT WORKING | CHECK AND FOUND CONDUCTIVITY PROBLEM. REPLACED BY PART , NOW CONDUCTIVITY PROBLEM IS SOLVED. | Closed | 2SXAE669 | HIGH | 78:25:0 | 4
29 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 37959 | 15-Jan-2020 | 08:00 |  |  | 16-Jan-2020 | 16:29 | NOT WORKING | CHECK AND FOUND FLOW ALARM IS COMING DURING DIALYSIS OF PATIENT. CLEANED ALL THE INLET FILTERS. CLEANED THE DEGASSING AND FLOW PUMP. CALIBRATION OF FLOW AT 500 AND 800 IS DONE. PROBLEM RECTIFIED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 32:29:0 | 2
30 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 23664 | 07-Jan-2019 | 08:07 |  |  | 07-Jan-2019 | 15:50 | NOT WORKING | CHECK AND FOUND NO CONDUCTIVITY IS COMING. PROBLEM FOUND IN BIC PUMP. FIXED THE SPRING SET OF BIC PUMP, HYPO IS BEING GIVEN AND BLOOD LEAK CALIBRATION IS BEING DONE. CALIBRATION OF CONDUCTIVITY IS ALSO BEING DONE. NOW, PROBLEM OF CONDUCTIVITY IS BEING SOLVED AND PROPER CONDUCTIVITY IS COMING. | Closed | 2SXAE669 | HIGH | 7:43:0 | 1
31 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00001 | DIALYSIS MACHINE | HEMA | 9767558503 | 41176 | 21-Mar-2020 | 08:08 |  |  | 21-Mar-2020 | 23:26 | NOT WORKING | CHECK AND FOUND 'FLOW ALARM' IS SHOWN DURING TEST AND CONDUCTIVITY TEST FAILS. CLEANED ALL THE INTERNAL TUBINGS AND FILTERS. DEGASSING AND FLOW PUMP IS CLEANED. CALIBRATION OF FLOW IS BEING DONE. HOT-DISINFECTION AND FRONT CLEANING IS GIVEN.FLOW PROBLEM IS RESOLVED. CONDUCTIVITY PROBLEM IS RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE669 | HIGH | 15:18:0 | 1
32 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | DORTHI SCOTT | 9764971169 | 5910 | 04-Aug-2017 | 12:08 |  |  | 05-Aug-2017 | 14:30 | NOT WORKING | REPLACED NEW TUBE FIXING FOR BLOOD PUMP,NOW MACHINE IS WORKING CONDITION | Closed | 2SXAE07 | HIGH | 26:22:0 | 2
33 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEEMA | 9727558503 | 29489 | 05-Jul-2019 | 09:09 |  |  | 07-Jul-2019 | 23:35 | NOT WORKING | CHECK AND FOUND WATER ALARM IS ALWAYS COMING AND TEST COULD NOT BE PASSED. CLEANED THE HEATER CHAMBER, INLET VALVES, A & B PART SUCTION TUBES, UF PUMP TUBES AND DIASAFE TUBES. NOW, WATER ALARM ISSUE IS BEING SOLVED. CALIBRATION OF BLOOD LEAK, 500 & 800 FLOW IS DONE. NOW, DIALYSIS MACHINE IS WORKING NORMALLY. | Closed | 2SXAE07 | HIGH | 62:26:0 | 3
34 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA BHORGAR | 9767558503 | 5689 | 27-Jul-2017 | 09:05 |  |  | 28-Jul-2017 | 14:35 | NOT WORKING | AFTER CALIBRATION AND INSERT WATER INLET FILTER MACHINE WORKING PROPERLY | Closed | 2SXAE07 | HIGH | 29:30:0 | 2
35 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA | 9767558503 | 10233 | 08-Dec-2017 | 09:15 |  |  | 13-Dec-2017 | 12:30 | NOT WORKING | AFTER CHECKED MACHINE AN POWER SUPPLY BOARD GET BURNT, SO WE REPLACED NEW POWER SUPPLY UNIT INTO AN MACHINE.NOW MACHINE IS WORKING CONDITION. | Closed | 2SXAE07 | HIGH | 123:15:0 | 6
36 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 35309 | 18-Nov-2019 | 08:58 |  |  | 21-Nov-2019 | 16:46 | NOT WORKING | CHECK AND FOUND WATER EXCRETED OUT FROM THE BODY IS VERY LESS AS COMPARED TO SET VALUE OF UF. WHEN DONE THE DUMMY DIALYSIS BY TAKING 500 ML OF SALINE BOTTLE, ONLY 400 ML OF WATER IS EXCRETED OUT. UF PUMP PROBLEM IS FOUND. CHANGED THE OLD UF PUMP WITH THE NEW ONE. NOW, EXACT AMOUNT OF SET UF WATER IS EXCRETED DURING DIALYSIS BY ANALYSING THE WEIGHT. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE07 | HIGH | 79:48:0 | 4
37 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | SARIKA WASNIK | 8446603844 | 20983 | 02-Oct-2018 | 09:25 |  |  | 06-Oct-2018 | 12:11 | NOT WORKING | CHECK AND FOUND BLOOD PUMP HARNESH CABLE WAS DEFECTIVE,REPAIRE IT THE CABLE FIXED IN THE MACHINE NOW MACHINE IS WORKING | Closed | 2SXAE07 | HIGH | 98:46:0 | 5
38 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | MANISH NURSE | 9049837042 | 16950 | 20-Jun-2018 | 08:15 |  |  | 20-Jun-2018 | 11:23 | NOT WORKING | CLEANED THE PART B MOTOR AND IT'S SPRING | Closed | 2SXAE07 | HIGH | 3:8:0 | 1
39 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 7470 | 25-Sep-2017 | 08:00 |  |  | 26-Sep-2017 | 11:55 | NOT WORKING | CLEANED BICARBONATE MOTOR SPRING PROPERLY,NOW TEST GET PASSED MACHINE IS WORKING CONDITION. | Closed | 2SXAE07 | HIGH | 27:55:0 | 2
40 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | Hema Borjhar: | 9767558503 | 42566 | 23-May-2020 | 09:09 |  |  | 23-May-2020 | 14:30 | NOT WORKING | PROBLEM REPORTED OF FLOW ALARM AND NO CONDUCTIVITY DURING DIALYSIS OF PATIENT. CHECK AND FOUND THERE IS ERROR CODE OF 147 DURING CALIBRATION OF 500 FLOW. FIXED THE NEW GEAR PUMP ON BOTH DEGASSING AND FLOW PUMP AND NEW CARBON BRUSH IN FLOW PUMP. ON FURTHER, FOUND DIAPHRAGM OF BALANCING CHAMBER IS RUPTURED AND FIXED THE NEW DIAPHRAGM AS WELL. PROBLEM OF FLOW FLUCTUATION AND NO CONDUCTIVITY IS RESOLVED. DIALYSIS MACHINE IS NOW WORKING WITH NO PROBLEM DURING DIALYSIS OF PATIENT. | Closed | 2SXAE07 | HIGH | 5:21:0 | 1
41 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | AKSHAY SAVARKAR | 8888591294 | 17658 | 09-Jul-2018 | 08:31 |  |  | 16-Jul-2018 | 05:45 | NOT WORKING | CHECK AND FOUND CONDUCTIVITY AND FLOW ALARM IS SHOWING IN MACHINE. CHECK FLOW MOTOR FOUND CARBON CONTAIN IN THE MOTOR. CLEAN THE MOTOR AND INSTALL, DID ALL TEST FOUND PASS. NOW DIALYSIS MACHINE WORKING SATISFACTORY. | Closed | 2SXAE07 | HIGH | 165:14:0 | 8
42 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | KHEME SISTER | 9767558503 | 20296 | 12-Sep-2018 | 08:08 |  |  | 18-Sep-2018 | 01:10 | NOT WORKING | CHECK AND FOUND FLOW ALARM, WATER ALARM AND BLOOD LEAK PROBLEM DETECTED.ABOVE MENTION PROBLEM IS SOLVE NOW MACHINE IS WORKING | Closed | 2SXAE07 | HIGH | 137:2:0 | 7
43 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | SARIKA | 8446603844 | 21922 | 08-Nov-2018 | 08:00 |  |  | 13-Nov-2018 | 08:11 | CONDUCTIVITY ISSUE | CHECK AND FOUND PROBLEM OF LOW CONDUCTIVITY AND FLOW EXIST. A PART AND B PART CHAMBER IS CLEANED AND FIXED. DIASAFE CLIP IS FIXED TO SOLVE FLOW PROBLEM. NOW DIALYSIS MACHINE IS WORKING. | Closed | 2SXAE07 | HIGH | 120:11:0 | 6
44 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | SUNIL | 9158418368 | 24896 | 11-Feb-2019 | 09:09 |  |  | 12-Feb-2019 | 13:58 | NOT WORKING | CHECK AND FOUND MACHINE CONDUCTIVITY IS LOW. NEED TO CLEAN THE FILTER AND SPRING SET. CLEAN THE FILTER AND SPRING SET, NOW MACHINE IS WORKING | Closed | 2SXAE07 | HIGH | 28:49:0 | 2
45 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA GODJAD | 9767558503 | 23606 | 04-Jan-2019 | 09:06 |  |  | 07-Jan-2019 | 10:09 | CONDATIVITY ISSUE | CHECK AND FOUND ANSTANDBILITY OF CONDUCTIVITY WHILE PATIENT IS ON DIALYSIS.DEGASSING MOTOR CARBON BRUSH IS DEGRADED AND IT HAS CHANGED. CALIBRATION OF CONDUCTIVITY IS BEING DONE. NOW CONDUCTIVITY IS MAINTAINED PROPERLY. DIALYSIS MACHINE IS NOW WORKING | Closed | 2SXAE07 | HIGH | 73:3:0 | 4
46 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | Hema Borjhar: | 9767558503 | 42328 | 17-May-2020 | 09:00 |  |  | 18-May-2020 | 15:02 | NOT WORKING | PROBLEM REPORTED OF NO ACCURACY IN ULTRA FILTRATION (UF) PROCESS. CLEANED ALL THE INLET 5 FILTERS INCLUDING UF FILTER. UF PUMP FUNCTIONING IS BEING CHECKED IN CALIBRATION MODE WITH SET COUNT PULSE OF 60 AND EXACT 60 PULSE COUNT OF DIALYSATE IS EJECTED OUT. CALIBRATION OF BLOOD LEAK, FLOW IS BEING DONE. FRONT CLEANING AND HOT DISINFECTION IS BEING GIVEN. DUMMY DIALYSIS IS BEING DONE FOR 1/2 HOUR AND EXACT 500 ML OF UF SET VOLUME OF NS IS EXCRETED OUT THROUGH THE FLOW AND UF PUMP. MACHINE IS WORKING WITH NO ISSUE OF ULTRAFILTRATION PROCESS. DIALYSIS OF PATIENT IS BEING DONE LATER WITH NO COMPLAINT. | Closed | 2SXAE07 | HIGH | 30:2:0 | 2
47 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA BHORJKAR | 9767558503 | 11231 | 03-Jan-2018 | 08:10 |  |  | 05-Jan-2018 | 20:01 | NOT WORKING | CLEANING THE MOTOR BEARING AND FILTER AND REPLACED WATER INLET FILTER | Closed | 2SXAE07 | HIGH | 59:51:0 | 3
48 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | STOTT NURSE | 9764971169 | 14684 | 07-Apr-2018 | 09:13 |  |  | 09-Apr-2018 | 13:05 | NOT WORKING | REPLACED NEW INLET FILTER , NOW MACHINE IS WORKING CONDITION. | Closed | 2SXAE07 | HIGH | 51:52:0 | 3
49 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | MANISA NANAWARE | 9049837042 | 17347 | 30-Jun-2018 | 08:36 |  |  | 02-Jul-2018 | 06:32 | NOT WORKING | FOUND CONDUCTIVITY NOT WORKING. RECTIFIED THE PROBLEM. NOW WORKING SATISFACTORY WITHOUT ANY FALIOUR. | Closed | 2SXAE07 | HIGH | 45:56:0 | 3
50 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA BHORJAR | 9767558503 | 23229 | 21-Dec-2018 | 08:00 |  |  | 26-Dec-2018 | 19:40 | CONDUCTIVY AND AIR DETECTOR SENSSOR ISSUE | CHECK AND FOUND CONDUCTIVITY COULD NOT BE MAINTAINED AND THERE IS ABRUPT CHANGE IN IT. BIC PUMP SPRING SET IS DEFECTIVE. FIXED THE NEW BIC PUMP SPRING SET. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE07 | HIGH | 131:40:0 | 6
51 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 26386 | 29-Mar-2019 | 08:45 |  |  | 01-Apr-2019 | 15:19 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS LEAK. CHANGED THE DEFECTIVE WITH THE NEW ONE. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE07 | HIGH | 78:34:0 | 4
52 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA | 9767558503 | 41177 | 21-Mar-2020 | 08:09 |  |  | 21-Mar-2020 | 16:31 | NOT WORKING | CHECK AND FOUND UPPER FLOW ALARM ERROR IS DISPLAYED DURING DIALYSIS OF PATIENT. CHANGED THE CARBON BRUSH OF DEGASSING PUMP AS IT IS DEGRADED. CLEANED THE FLOW MOTOR. DIASAFE FILTER IS BEING CHANGED. CALIBRATION OF FLOW IS BEING DONE. UPPER FLOW ALARM PROBLEM IS BEING RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE07 | HIGH | 8:22:0 | 1
53 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | Hema Borjhar: | 9767558503 | 43800 | 03-Jul-2020 | 08:42 |  |  | 04-Jul-2020 | 10:50 | NOT WORKING | PROBLEM REPORTED OF "RINSE FAILURE F02" ERROR DURING STARTING OF HOT-DISINFECTION. CLEANED THE VALVE NUMBER 84 AND UF FILTER. CITRO-STERIL CAN TUBE IS CLEANED. "RINSE FAILURE F02" ERROR IS RESOLVED. NOW, HOT-DISINFECTION IS BEING DONE WITH NO ERROR. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE07 | HIGH | 26:8:0 | 2
54 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA GHORJAR | 9767558503 | 3716 | 13-May-2017 | 10:00 |  |  | 15-May-2017 | 03:30 | NOT WORKING | CLEANED MOTOR SECTION CONDUCTIVITY PROBLEM SOLVED,NOW MACHINE HANDOVER IN GOOD CONDITION. | Closed | 2SXAE07 | HIGH | 41:30:0 | 3
55 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | AKSHAY SAVARKAR | 8888591294 | 22536 | 01-Dec-2018 | 09:01 |  |  | 02-Dec-2018 | 11:57 | CONDUCTIVITY ISSUE | CHECK AND FOUND DIALYSIS MACHINE CONDUCTIVITY IS VERY LOW.CHECK THE B PART SPRING SET AND FIXED IT PROPERLY AFTER FIXED PROPERLY MACHINE IS WORKING GOOD | Closed | 2SXAE07 | HIGH | 26:56:0 | 2
56 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA BORJAR | 9767558503 | 28352 | 08-Jun-2019 | 08:07 |  |  | 09-Jun-2019 | 14:28 | NOT WORKING | CHECKED THE MACHINE AND FOUND PROBLEM IN HOT DISINFECTION PROCESS. SOLVED IT AND MACHINE IS IN WORKING CONDITION AND HANDED OVER TO THE USER. | Closed | 2SXAE07 | HIGH | 30:21:0 | 2
57 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 40919 | 16-Mar-2020 | 08:30 |  |  | 16-Mar-2020 | 11:41 | NOT WORKING | CHECK AND FOUND ULTRASONIC SENSOR IS NOT WORKING AND DUE TO WHICH AIR DETECTOR ALARM IS SHOWN WITH BLOOD PUMP NOT ROTATING. FIXED THE NEW ULTRASONIC SENSOR AND AIR DETECTOR ERROR IS REMOVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE07 | HIGH | 3:11:0 | 1
58 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00002 | DIALYSIS MACHINE | HEMA GHODAZAD | 9767558503 | 22821 | 11-Dec-2018 | 09:06 |  |  | 14-Dec-2018 | 11:35 | AIR DITAKE ISSUE | CHECK AND FOUND ULTRASONIC SENSOR IS NOT WORKING PROPERLY. REFIXED THE ULTRASONIC SENSOR IN PROPER WAY NOW MACHINE IS WORKING GOOD | Closed | 2SXAE07 | HIGH | 74:29:0 | 4
59 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA GODJAD | 9767558503 | 23607 | 04-Jan-2019 | 09:06 |  |  | 07-Jan-2019 | 06:13 | CONDATIVITY ISSUE | CHECK AND FOUND MACHINE CONDUCTIVITY IS NOT STABLE WHILE PATIENT IS ON DIALYSIS.PROBLEM IS FOUND OUT DEGASSING MOTOR AND IT'S CARBON BRUSH IS CHANGED. FLOW MOTOR IS CLEAN AND FIXED.CALIBRATION OF CONDUCTIVITY IS DONE. NOW CONDUCTIVITY IS MAINTAINED PROPERLY. DIALYSIS MACHINE IS NOW WORKING | Closed | 2SXAB438 | HIGH | 69:7:0 | 4
60 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA SISTER | 9767558503 | 21213 | 10-Oct-2018 | 08:18 |  |  | 12-Oct-2018 | 13:15 | NOT WORKING | CHECK AND FOUND NIBP CUFF FAULTY. CHANGED THE NEW NIBP CUFF. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAB438 | HIGH | 52:57:0 | 3
61 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | AKSHAY SOWRAKAR | 8888591294 | 17455 | 03-Jul-2018 | 09:33 |  |  | 05-Jul-2018 | 07:30 | NOT WORKING | FOUND CONDUCTIVITY PROBLEM. CLEAN ALL INTERNAL FILTERS. NOW IT'S WORKING WITHOUT ANY FALIOUR | Closed | 2SXAB438 | HIGH | 45:57:0 | 3
62 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | SARIKA VASNIK | 8446603844 | 21965 | 11-Nov-2018 | 08:40 |  |  | 13-Nov-2018 | 10:15 | NOT WORKING | CHECK AND FOUND PROBLEM OF LOW CONDUCTIVITY AND FLOW EXIST. A PART AND B PART CHAMBER IS CLEANED AND FIXED. DIASAFE CLIP IS FIXED TO SOLVE FLOW PROBLEM. NOW DIALYSIS MACHINE IS WORKING. | Closed | 2SXAB438 | HIGH | 49:35:0 | 3
63 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | KHEME SISTER | 9767558503 | 20297 | 12-Sep-2018 | 08:08 |  |  | 17-Sep-2018 | 10:04 | NOT WORKING | CHECK AND FOUND DIALYSIS MACHINE UF PUMP AND OUTPUT BORD WAS NOT FIXED PROPERLY.FIXED THE UF PUMP AND OUTPUT BORD NOW MACHINE IS WORKING | Closed | 2SXAB438 | HIGH | 121:56:0 | 6
64 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA GHORJAR | 9767558503 | 3717 | 13-May-2017 | 10:05 |  |  | 15-May-2017 | 02:30 | NOT WORKING | CLEANED BLOOD MOTOR SECTION AND ALSO DONE CALIBRATION OF VOLTAGE DIM,NOW MACHINE HAND OVER USER GOOD CONDITION | Closed | 2SXAB438 | HIGH | 40:25:0 | 3
65 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA GURJAR | 9767558503 | 7776 | 04-Oct-2017 | 09:28 |  |  | 05-Oct-2017 | 08:30 | NOT WORKING | CALIBRATION PROCESS DONE ON BLD, DIMNESS AND ALSO REPLACED NEW INLET FILTER, NOW ERROR PROBLEM GET RECTIFIED, MACHINE IS WORKING CONDITION. | Closed | 2SXAB438 | HIGH | 23:2:0 | 2
66 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | SARIKA WASNIK | 8446603844 | 16603 | 12-Jun-2018 | 08:10 |  |  | 12-Jun-2018 | 15:04 | NOT WORKING | DONE THE ROUTINE MAINTENANCE. NOW MACHINE WORKING | Closed | 2SXAB438 | HIGH | 6:54:0 | 1
67 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 38596 | 28-Jan-2020 | 08:00 |  |  | 27-Feb-2020 | 22:52 | NOT WORKING | CHECK AND FOUND FLUCTUATIONS IN FLOW AND CONDUCTIVITY. CHECKED THE FLOW AND DEGASSING PUMP AND CLEANED IT. DIAPHRAGM OF BALANCING CHAMBER IS CHECKED. CHECKED THE FUNCTIONING OF FLOAT SWITCH. CLEANED THE VALVE NUMBER 78,76,43 AND 26. CLEANED THE AIR SEPARATION PUMP AND CONDUCTIVITY CELL. CALIBRATION OF MIXING SYSTEM IS DONE. CALIBRATION OF VENOUS PRESSURE IS DONE ALONG WITH FLOW. PROBLEM IS RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAB438 | HIGH | 734:52:0 | 31
68 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | Hema Borjhar | 9767558503 | 41599 | 11-Apr-2020 | 09:30 |  |  | 15-Apr-2020 | 16:15 | NOT WORKING | CHECK AND FOUND FLUCTUATIONS IN FLOW AND CONDUCTIVITY. ALSO NEGATIVE AND POSITIVE PRESSURE TEST FAILS. CHANGED THE NEW DIAPHRAGM OF BALANCING CHAMBER. NEW TMP SENSOR IS FIXED. BOTH THE GEAR PUMP IS BEING NEWLY FIXED. CALIBRATION OF TMP AND FLOW IS BEING DONE. PROBLEM IS RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAB438 | HIGH | 102:45:0 | 5
69 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA GODJAR | 9767558503 | 25453 | 27-Feb-2019 | 09:13 |  |  | 02-Mar-2019 | 16:43 | NOT WORKING | CHECK AND FOUND TEMPERATURE COULD NOT BE MAINTAINED AT 37 DEGREE CELSIUS AND IT AUTOMATICALLY SHIFTS TO 42 DEGREE CELSIUS CAUSING HEATING OF TUBES AND TEMPERATUTE TEST COULD NOT BE PASSED. PROBLEM FOUND IS OF SMPS. OLD DEFECTIVE SMPS BOARD IS BEING REPLACED WITH THE NEW REPAIRED SMPS BOARD.NOW, TEMPERATURE IS MAINTAINED AT 37 DEGREE CELSIUS AND TEMPERATURE TEST IS ALSO PASSED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAB438 | HIGH | 79:30:0 | 4
70 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | D V STOTT | 9764971169 | 14979 | 17-Apr-2018 | 09:10 |  |  | 21-Apr-2018 | 10:33 | NOT WORKING | BLOOD LEVEL DETECTOR REPLACED. NOW THE MACHINE IS IN WORKING CONDITION | Closed | 2SXAB438 | HIGH | 97:23:0 | 5
71 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA BHORJAR | 9767558503 | 23282 | 22-Dec-2018 | 08:03 |  |  | 26-Dec-2018 | 21:30 | CONDUCTIVITY ISSUE | CHECK AND FOUND CONDUCTIVITY COULD NOT BE MAINTAINED AND THERE IS ABRUPT CHANGE IN IT. BIC PUMP SPRING SET IS DEFECTIVE. FIXED THE NEW BIC PUMP SPRING SET. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAB438 | HIGH | 109:27:0 | 5
72 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 31634 | 22-Aug-2019 | 09:03 |  |  | 24-Aug-2019 | 15:28 | NOT WORKING | CHECK AND FOUND MOTHER BOARD ISSUE. ALSO BLOOD PUMP IS NOT WORKING. MR IS BEING RAISED FOR MOTHER BOARD AND BLOOD PUMP STEPPER MOTOR. MATERIAL IS IN TRANSIT AND AS SOON AS PARTS WILL BE RECEIVED, COMPLAINT WILL BE RAISED AGAIN AND MACHINE WILL BE MADE OK. AT PRESENT, PRESENT COMPLAINT IS BEING CLOSED. | Closed | 2SXAB438 | HIGH | 54:25:0 | 3
73 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | NILESH TAROKAR | 9860704598 | 5847 | 03-Aug-2017 | 09:31 |  |  | 05-Aug-2017 | 09:30 | NOT WORKING | REPLACED NEW INLET FILTER,NOW MACHINE IS WORKING CONDITION | Closed | 2SXAB438 | HIGH | 47:59:0 | 3
74 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | NILESH PALEKAR | 986070498 | 16996 | 20-Jun-2018 | 10:10 |  |  | 21-Jun-2018 | 14:26 | NOT WORKING | CLEANED THE PART A AND B FILTER ALSO IT'S MOTOR | Closed | 2SXAB438 | HIGH | 28:16:0 | 2
75 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | SUNIL GAJBHIYE | 9158418368 | 10416 | 12-Dec-2017 | 10:26 |  |  | 16-Dec-2017 | 09:21 | NOT WORKING | REPLACED NEW TUBING SILICON REINF.6.35X2.65MM 3M.NOW MACHINE IS WORKING CONDITION. | Closed | 2SXAB438 | HIGH | 94:55:0 | 5
76 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 11427 | 09-Jan-2018 | 09:00 |  |  | 09-Jan-2018 | 23:45 | NOT WORKING | REPLACED THE INLET FILTER AND CLEANING THE MOTOR BEARING AND FILTER SECTION | Closed | 2SXAB438 | HIGH | 14:45:0 | 1
77 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA BORJAR | 9767558503 | 28353 | 08-Jun-2019 | 08:07 |  |  | 12-Jun-2019 | 11:46 | NOT WORKING | CHECK AND FOUND LOW CONDUCTIVITY PROBLEM. ALSO RINSE FAILURE F02 IS COMING. CLEANED THE A AND B PART TUBING THROUGH BIC PUMP AND CONDUCTIVITY PROBLEM IS SOLVED. FOR RINSE FAILURE F02, NEW UF PUMP IS REQUIRED. REQUISITION IS BEING SENT FOR THE NEW PUMP AS SOON AS NEW PUMP IS RECEIVED, CALL WILL BE RAISED AGAIN AND NEW PUMP WILL BE FIXED AND RINSE FAILURE PROBLEM WOULD BE SORTED OUT. | Closed | 2SXAB438 | HIGH | 99:39:0 | 5
78 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA BORZER | 9767558503 | 27242 | 02-May-2019 | 09:05 |  |  | 06-May-2019 | 12:31 | NOT WORKING | CHECK AND FOUND CONDUCTIVITY COULD NOT BE OBTAINED. BLOCKAGE IN B PART SUCTION TUBE DETECTED. CLEANED THE B PART TUBE AND ALSO THE BICARBONATE PUMP AND ULTRA FILTRATE FILTER. CONDUCTIVITY PROBLEM RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAB438 | HIGH | 99:26:0 | 5
79 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA BORJHAR | 9767558503 | 37649 | 08-Jan-2020 | 09:05 |  |  | 08-Jan-2020 | 16:19 | NOT WORKING | CHECK AND FOUND THERE IS PROBLEM OF UPPER FLOW ALARM. 'RINSE FAILURE F02 & V84 FAULTINESS' ERROR ARE DISPLAYED DURING HOT DISINFECTION AND AFTER TEST IS PASSED. CLEANED ALL THE INLET FILTERS. DEGASSING AND FLOW PUMP IS CLEANED. CALIBRATION OF FLOW AT 500 AND 800 IS DONE. VALVE NUMBER 84 IS CLEANED. PROBLEM IS RESOLVED OF UPPER FLOW ALARM, RINSE FAILURE & VALVE 84 FAULT. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAB438 | HIGH | 7:14:0 | 1
80 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | MANISHA | 9049837042 | 18439 | 26-Jul-2018 | 08:00 |  |  | 27-Jul-2018 | 11:56 | NOT WORKING | LOW CONDUCTIVITY PROBLEM IS BEING REPORTED. CHECK AND FOUND B-PART CONCENTRATION IS NOT ACCURATE AND CHANGED THE B-PART. ALL FILTERS ARE CLEANED. MACHINE IS NOW WORKING. | Closed | 2SXAB438 | HIGH | 27:56:0 | 2
81 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | SARIKA VASE | 8446603844 | 19040 | 10-Aug-2018 | 09:00 |  |  | 14-Aug-2018 | 15:30 | NOT WORKING | CHECK AND FOUND WATER ALARM IS THEIR. REPLACED CARBON BRUSH OF MOTOR. NOW MACHINE IS IN WORKING CONDITION. | Closed | 2SXAB438 | HIGH | 102:30:0 | 5
82 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | ANIRRUDHA | 9403242582 | 17810 | 12-Jul-2018 | 08:30 |  |  | 13-Jul-2018 | 12:30 | NOT WORKING | CHECK AND FOUND BLOOD LEAK TEST IS NOT PASSED . DID CALIBRATION OF BLOOD LEAK AND DIMNESS. NOW BLOOD LEAK TEST IS PASS AND MACHINE WORKING SATISFACTORY. | Closed | 2SXAB438 | HIGH | 28:0:0 | 2
83 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA BORJAR | 9767558503 | 40359 | 03-Mar-2020 | 09:01 |  |  | 07-Mar-2020 | 16:28 | NOT WORKING | CHECK AND FOUND THERE IS FLUCTUATIONS IN FLOW AND DUE TO WHICH CONDUCTIVITY ALSO FLUCTUATES. THIS CALL WAS ASSIGNED TO FRESENIUS ENGINEER AND AFTER HIS VISIT HE CHANGED THE DIAPHRAGM AND 4 VALVES OF BALANCING CHAMBER. FLOW FLUCTUATIONS PROBLEM WAS RESOLVED BUT LATER HIGH CONDUCTIVITY AND POSITIVE & NEGATIVE TEST FAILURE ISSUE ARISES. FRESENIUS ENGINEER HAVE CONFIRMED HIS NEXT VISIT AFTER LOCKDOWN WILL BE OVER I.E 14TH APRIL. LATER AGAIN THE CALL WILL BE RAISED AFTER FRESENIUS ENGINEER VISIT AND PROBLEM WILL BE RESOLVED. | Closed | 2SXAB438 | HIGH | 103:27:0 | 5
84 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 26387 | 29-Mar-2019 | 08:45 |  |  | 03-Apr-2019 | 11:44 | NOT WORKING | CHECK AND FOUND MACHINE NIBP CUFF IS DEFECTIVE. NEED TO CHANGE . CHANGE THE NIBP CUFF NOW MACHINE IS WORKING GOOD | Closed | 2SXAB438 | HIGH | 122:59:0 | 6
85 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | NILESH KAROKAR | 9860704598 | 20035 | 04-Sep-2018 | 08:40 |  |  | 05-Sep-2018 | 08:40 | NOT WORKING | CHECK AND FOUND NO CONDUCTIVITY ISSUE.FIXED THE A&B PART CHEMBER AND SPRING ATTACHED.CONDUCTIVITY ISSUE RESOLVED.DIALYSIS MACHINE IS WORKING NOW | Closed | 2SXAB438 | HIGH | 24:0:0 | 2
86 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | HEMA BURJAR | 9767558503 | 21690 | 27-Oct-2018 | 08:08 |  |  | 28-Oct-2018 | 17:56 | NOT WORKING | CHECK AND FOUND OPTICAL DETECTOR TEST FAILS. ULTRASONIC SENSOR IS ALSO NOT WORKING. NEW LP 450M CONTROLLER BOARD IS FIXED ALONG WITH NEW ULTRASONIC SENSOR. VENOUS PRESSURE TEST ISSUE BEING RESOLVED.ALL TEST IS PASSED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAB438 | HIGH | 33:48:0 | 2
87 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00003 | DIALYSIS MACHINE | NILESH DAROKAR | 9860704598 | 6850 | 08-Sep-2017 | 09:35 |  |  | 12-Sep-2017 | 14:30 | NOT WORKING | RECONNECTED THE DISPLAY CONNECTOR AND REASSEMBLE ALL CPU BOARDS.NOW MACHINE IS WORKING CONDITION. | Closed | 2SXAB438 | HIGH | 100:55:0 | 5
88 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA | 9767558503 | 9590 | 23-Nov-2017 | 09:55 |  |  | 27-Nov-2017 | 09:42 | NOT WORKING | REPLACED NEW INLET FILTER, UF-PUMP FILTER,FILTER 5,0MM AND FILTER 7,0MM.NOW MACHINE IS WORKING CONDITION. | Closed | 2SXAE667 | HIGH | 95:47:0 | 5
89 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 29612 | 08-Jul-2019 | 08:27 |  |  | 10-Jul-2019 | 14:20 | NOT WORKING | CHECK AND FOUND DISPLAY SCREEN IS DEFECTIVE. M. R HAS BEEN SENT AND AS SOON AS NEW DISPLAY SCREEN WILL BE RECEIVED AGAIN CALL WILL BE RAISED AND MACHINE WILL BE MADE FUNCTIONAL. AT PRESENT CLOSING THIS PRESENT CALL. | Closed | 2SXAE667 | HIGH | 53:53:0 | 3
90 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | Hema Borjhar | 9767558503 | 43248 | 19-Jun-2020 | 08:10 |  |  | 19-Jun-2020 | 16:30 | NOT WORKING | PROBLEM REPORTED OF UPPER FLOW ALARM DURING DIALYSIS OF PATIENT. CLEANED ALL THE FILTERS. CALIBRATION OF FLOW ALONG WITH BLOOD LEAK IS DONE. HOT DISINFECTION AND FRONT CLEANING IS BEING GIVEN. PROBLEM IS RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE667 | HIGH | 8:20:0 | 1
91 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA GHORJAR | 9767558503 | 3718 | 13-May-2017 | 10:10 |  |  | 15-May-2017 | 10:30 | NOT WORKING | CLEANED BLOOD MOTOR SECTION AND ALSO CALIBRATION PROCEDURE IS DONE PROPERLY NOW BLOOD LEAKAGE PROBLEM SOLVED,MACHINE HANDOVER IN GOOD WORKING CONDITION. | Closed | 2SXAE667 | HIGH | 48:20:0 | 3
92 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | Hema Borjar | 9767558503 | 41550 | 09-Apr-2020 | 08:00 |  |  | 09-Apr-2020 | 14:15 | NOT WORKING | CHECK AND FOUND BIBAG OPTION IS NOT SHOWN WHILE FIXING BIBAG INSTEAD OF B'PART CAN. FIXED TH NEW BIBAG SENSOR AS EARLIER ONE WAS FAULTY AND WAS UNPLUCKED. DONE THE NECESSARY FLOW CALIBRATION. NOW, BIBAG CAN BE USED AND DIALYSIS IS BEING DONE. | Closed | 2SXAE667 | HIGH | 6:15:0 | 1
93 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA BORZER | 9767558503 | 27239 | 02-May-2019 | 09:01 |  |  | 05-May-2019 | 17:34 | NOT WORKING | CHECK AND FOUND MACHINE DOESN'T TURN ON. HEATER BOARD ISSUE IS DETECTED DUE TO WHICH VOLTAGE COULD NOT BE SUPPLIED TO SMPS BOARD. REPAIRED THE HEATER BOARD AND POWER SUPPLY ISSUE IS RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE667 | HIGH | 80:33:0 | 4
94 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA BORJHAR | 9767558503 | 37650 | 08-Jan-2020 | 09:08 |  |  | 08-Jan-2020 | 10:50 | NOT WORKING | CHECK AND FOUND PROBLEM OF HIGH CONDUCTIVITY WITH REPEATED FLUCTUATIONS, FLOW AND WATER ALARM ISSUE. VALVE NUMBER 41 AND CONDUCTIVITY CELL IS REPLACED WITH THE NEW ONE. ALSO, SPRING SET OF BIC PUMP AND ACETATE PUMP IS REPLACED WITH THE NEW ONE. FLOW & WATER ALARM PROBLEM ARE RESOLVED. REGARDING HIGH CONDUCTIVITY, BIBAG SENSOR FAULT IS FOUND AND IT IS UNPLUCKED FROM MACHINE. DUMMY DIALYSIS IS DONE AND ALSO DIALYSIS OF PATIENT IS DONE. PROBLEM IS RESOLVED. MACHINE IS WORKING. | Closed | 2SXAE667 | HIGH | 1:42:0 | 1
95 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | SHUBHANGI NURSE | 8888291294 | 16815 | 16-Jun-2018 | 09:01 |  |  | 16-Jun-2018 | 13:08 | NOT WORKING | CHANGED THE PART B SOLUTION. | Closed | 2SXAE667 | HIGH | 4:7:0 | 1
96 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | JYOTI CHAUDHARI | 7588768076 | 4656 | 20-Jun-2017 | 10:00 | 22-Jun-2017 | 10:19 | 24-Jun-2017 | 08:50 | NOT WORKING | REPLACE NEW UF PUMP | Closed | 2SXAE667 | HIGH | 94:50:0 | 5
97 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | SEEMA BORJAD | 9767558503 | 27911 | 25-May-2019 | 09:02 |  |  | 25-May-2019 | 17:17 | NOT WORKING | CHECK AND FOUND INTERNAL WATER LEAKAGE. ALSO NO CONDUCTIVITY AND BLOOD LEAKAGE ISSUE FOUND. FIXED THE INTERNAL TUBING AND LEAKAGE ISSUE IS RESOLVED. BLOCKAGE IN B'PART BIC PUMP IS RESOLVED AND CONDUCTIVITY ISSUE SORTED OUT. CALIBRATION OF BLOOD LEAK IS DONE. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE667 | HIGH | 8:15:0 | 1
98 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA BORJAR | 9767558503 | 28354 | 08-Jun-2019 | 08:07 |  |  | 09-Jun-2019 | 14:35 | NOT WORKING | CHECKED THE MACHINE AND PROBLEM FOUND IN HOT DISINFECTION PROCESS AND V39 FAILURE ERROR. CLEAN ALL FILTER AND SOLVED CONDUCTIVITY PROBLEM. NOW MACHINE IS WORKING FINE AND HANDED OVER TO USER. | Closed | 2SXAE667 | HIGH | 30:28:0 | 2
99 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 28790 | 19-Jun-2019 | 09:08 |  |  | 19-Jun-2019 | 18:32 | NOT WORKING | CHECK AND FOUND "RINSE FAILURE FO2" ERROR IS DISPLAYED WHILE DOING HOT DISINFECTION. UF PUMP IS FOUND TO BE DEFECTIVE. CHANGED THE OLD UF PUMP WITH THE NEW ONE. NOW, HOT DISINFECTION IS DONE WITHOUT ANY ERROR.DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE667 | HIGH | 9:24:0 | 1
100 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | D V STOTT | 9764971169 | 14980 | 17-Apr-2018 | 09:10 |  |  | 21-Apr-2018 | 10:50 | NOT WORKING | REPLACED THE BLOOD LEVEL DETECTOR. NOW THE MACHINE IS IN WORKING CONDITION | Closed | 2SXAE667 | HIGH | 97:40:0 | 5
101 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | Hema Borjhar | 9767558503 | 43610 | 29-Jun-2020 | 08:22 |  |  | 29-Jun-2020 | 10:45 | NOT WORKING | PROBLEM REPORTED OF FLOW ALARM DURING DISINFECTION AND DIALYSIS OF PATIENT. CALIBRATION OF FLOW IS DONE. FILTERS ARE CLEANED. HOT-DISINFECTION & FRONT CLEANING IS GIVEN. PROBLEM IS RESOLVED AND DIALYSIS OF PATIENT IS BEING DONE WITH NO ISSUE AND ANY ALARM. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE667 | HIGH | 2:23:0 | 1
102 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 35592 | 22-Nov-2019 | 08:00 |  |  | 23-Nov-2019 | 16:19 | NOT WORKING | CHECK AND FOUND HIGH CONDUCTIVITY ISSUE. DIALYSIS COULD NOT BE DONE AS CONDUCTIVITY IS ALWAYS HIGH AND FLOW STOPS. CROSS-CHECKED THE FLOW AND DEGASSING PUMP, BIC PUMP SPRING SET IS CHANGED AND HAS BEEN CROSS-CHECKED. INTERNAL TUBINGS ARE CLEANED. PROBLEM IS PARTIALLY RESOLVED. MACHINE IS UNDER OBSERVATION. | Closed | 2SXAE667 | HIGH | 32:19:0 | 2
103 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA SISTER | 9767558503 | 21211 | 10-Oct-2018 | 08:17 |  |  | 28-Oct-2018 | 10:37 | NOT WORKING | CHECK AND FOUND POSITIVE PRESSURE TEST FAILS. LEAK IN SHUNT DIASAFE TUBING FOUND. FIXED THE TUBING LEAK. BLOOD LEAK PROBLEM RESOLVED. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE667 | HIGH | 434:20:0 | 19
104 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 6780 | 08-Sep-2017 | 08:01 |  |  | 09-Sep-2017 | 14:25 | NOT WORKING | SOLDERED THE SHUT BUS CABLE WIRE PROPERLY, NOW MACHINE IS WORKING CONDITION | Closed | 2SXAE667 | HIGH | 30:24:0 | 2
105 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA BOJDAR | 9767559503 | 36438 | 11-Dec-2019 | 09:22 |  |  | 13-Dec-2019 | 17:31 | NOT WORKING | CHECK AND FOUND HIGH CONDUCTIVITY ISSUE DURING DIALYSIS OF PATIENT. ENGINEER FROM FRESENIUS HAS VISITED FOR THIS BREAKDOWN BUT PROBLEM COULD NOT BE SOLVED. FURTHER SENIOR ENGINEER VISIT FROM FRESENIUS IS TO BE MADE.AT PRESENT THIS BREAKDOWN COMPLIANT IS BEING CLOSED. | Closed | 2SXAE667 | HIGH | 56:9:0 | 3
106 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 10251 | 09-Dec-2017 | 08:00 |  |  | 12-Dec-2017 | 14:54 | NOT WORKING | REPLACED NEW TFT DISPLAY INTO AN MACHINE,AS PER THE REIMBURABLE WORK IT HAS BEEN DONE. | Closed | 2SXAE667 | HIGH | 78:54:0 | 4
107 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA VARDHA | 9767558503 | 6172 | 16-Aug-2017 | 10:20 |  |  | 18-Aug-2017 | 11:13 | NOT WORKING | REPLACED DISPLAY BOARD | Closed | 2SXAE667 | HIGH | 48:53:0 | 3
108 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | SCOTT NURSE | 9764971169 | 17803 | 12-Jul-2018 | 08:10 |  |  | 12-Jul-2018 | 11:30 | NOT WORKING | CHECK AND FOUND BLOOD LEAK TEST NOT PASS. DID CALIBRATION OF BLOOD LEEK AND DIMNESS. NOW MACHINE IS WORKING SATISFACTORY | Closed | 2SXAE667 | HIGH | 3:20:0 | 1
109 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | SUNIL | 9158418368 | 18872 | 06-Aug-2018 | 08:57 |  |  | 10-Aug-2018 | 17:30 | NOT WORKING | CHECK AND FOUND, MACHINE GIVING WATER ALARM. REPLACED THE CARBON BRUSH OF MOTOR. NOW MACHINE IS IN WORKING CONDITION. HAD OVER TO USER | Closed | 2SXAE667 | HIGH | 104:33:0 | 5
110 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | DORTHI SCOTT | 9764971169 | 5911 | 04-Aug-2017 | 12:10 |  |  | 05-Aug-2017 | 14:25 | NOT WORKING | REPLACED NEW TUBE FIXING FOR BLOOD PUMP,NOW MACHINE HANDOVER TO USER AS GOOD WORKING CONDITION. | Closed | 2SXAE667 | HIGH | 26:15:0 | 2
111 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA BHORJKAR | 9767558503 | 11209 | 02-Jan-2018 | 09:00 |  |  | 03-Jan-2018 | 20:05 | NOT WORKING | REPLACED THE FILTERS AND CLEAN THE MOTOR BEARING AND FILTER | Closed | 2SXAE667 | HIGH | 35:5:0 | 2
112 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA BHORGAR | 9767558503 | 5688 | 27-Jul-2017 | 09:05 |  |  | 28-Jul-2017 | 14:40 | NOT WORKING | PROPERLY CLEAN THE DISPLAY CABLE AND ASSEMBLE WATER INLET FILTER | Closed | 2SXAE667 | HIGH | 29:35:0 | 2
113 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 25752 | 12-Mar-2019 | 08:09 |  |  | 12-Mar-2019 | 15:43 | UPPER FLOW ALARM ISSUE | CHECK AND FOUND FLOW ALARM IS DISPLAYED ON THE SCREEN. DEGASSING MOTOR CARBON BRUSH IS DEGRADED. INSERTED THE NEW CARBON BRUSH AND FLOW IS CHECKED. FLOW CALIBRATION IS DONE. NOW,FLOW PROBLEM IS RESOLVED. DIALYSIS MACHIN IS NOW WORKING. | Closed | 2SXAE667 | HIGH | 7:34:0 | 1
114 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA BORZAR | 9767558503 | 35310 | 18-Nov-2019 | 08:58 |  |  | 18-Nov-2019 | 17:12 | NOT WORKING | CHECK AND FOUND MACHINE COULD NOT TURN 'ON' AND THERE IS ISSUE OF MOTHER BOARD. DEGASSING PUMP IS ALSO NOT WORKING. CONDUCTIVITY FLUCTUATES DURING DIALYSIS OF PATIENT AND ISSUE OF BIC PUMP SPRING SET IS FOUND. FIXED THE NEW MOTHER BAORD, NEW DEGASSING PUMP AND NEW BIC PUMP SPRING SET. CALIBRATION OF CONDUCTIVITY IS DONE. DIALYSIS MACHINE IS NOW WORKING. | Closed | 2SXAE667 | HIGH | 8:14:0 | 1
115 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | HEMA NURSE | 9767558503 | 26388 | 29-Mar-2019 | 08:45 |  |  | 03-Apr-2019 | 11:43 | NOT WORKING | CHECK AND FOUND MACHINE NIBP CUFF IS DEFECTIVE. NEED TO CHANGE . CHANGE THE NIBP CUFF NOW MACHINE IS WORKING GOOD | Closed | 2SXAE667 | HIGH | 122:58:0 | 6
116 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | AKSHAY SAVARKAR | 8888591294 | 17611 | 06-Jul-2018 | 09:57 |  |  | 06-Jul-2018 | 16:15 | NOT WORKING | CHECK AND FOUND CUNDITIVITY NOT COMMING. REPLACED BY PART , NOW CONDUCTIVITY COMMING, WORKING SATISFACTORY. | Closed | 2SXAE667 | HIGH | 6:18:0 | 1
117 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00004 | DIALYSIS MACHINE | NILESH KAROKAR | 9860704598 | 20036 | 04-Sep-2018 | 08:41 |  |  | 05-Sep-2018 | 06:07 | NOT WORKING | CHECK AND FOUND NO CONDUCTIVITY ISSUE.FIXED THE A&B PART CHEMBER AND SPRING ATTACHED.CONDUCTIVITY ISSUE RESOLVED.DIALYSIS MACHINE IS WORKING NOW. | Closed | 2SXAE667 | HIGH | 21:26:0 | 2
118 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00007 | RADIANT WARMER | CHAUDHARI | 7588768076 | 35097 | 13-Nov-2019 | 09:08 |  |  | 14-Nov-2019 | 15:06 | NOT WORKING | CHECK AND FOUND SIDE CRADLE IS BROKEN. MR HAS BEEN SENT AND MATERIAL IS ON TRANSIT. AT PRESENT THIS BREAKDOWN COMPLAINT IS CLOSED. AS SOON AS SIDE CRADLE WILL BE RECEIVED, AGAIN THE CALL WILL BE RAISED AND MACHINE WILL BE MADE FUNCTIONAL. | Closed | RHW2101A10413/223 | HIGH | 29:58:0 | 2
119 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00007 | RADIANT WARMER | JOYTI CHOUDHARI | 7588768076 | 38850 | 01-Feb-2020 | 09:11 |  |  | 01-Feb-2020 | 14:45 | NOT WORKING | CHECK AND FOUND FEMALE CABLE CONNECTOR IS BURNT AND DUE TO WHICH THERE IS POWER SUPPLY FAILURE ISSUE. FIXED THE NEW FEMALE CABLE CONNECTOR. POWER SUPPLY FAILURE ISSUE IS RESOLVED. CUT-OFF OF HEATER IS CHECKED. RADIANT WARMER IS NOW WORKING. | Closed | RHW2101A10413/223 | HIGH | 5:34:0 | 1
120 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00007 | RADIANT WARMER | COUDHRE NURSE | 7588768076 | 7934 | 10-Oct-2017 | 09:07 |  |  | 11-Oct-2017 | 15:35 | NOT WORKING | REPLACE THE TRANSFORMER | Closed | RHW2101A10413/223 | HIGH | 30:28:0 | 2
121 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00007 | RADIANT WARMER | CHOUDHARI NURSE | 7588768076 | 18061 | 18-Jul-2018 | 08:16 |  |  | 18-Jul-2018 | 15:45 | NOT WORKING | CHECK AND FOUND POWER FUSE IS FAULTY. REPLACED GLASS FUSE . NOW WARMER IS WORKING SATISFACTORY. | Closed | RHW2101A10413/223 | HIGH | 7:29:0 | 1
122 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00008 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 37564 | 06-Jan-2020 | 08:25 |  |  | 06-Jan-2020 | 17:48 | NOT WORKING | CHECK AND FOUND AIR PROBE TEMPERATURE SENSED IS VERY HIGH AND DUE TO WHICH HEATER COULD NOT TURN 'ON'. FIXED THE LOOSE CONNECTION OF FEMALE AIR PROBE CONNECTOR THROUGH DISPLAY CONTROLLER BOARD. NOW, AIR PROBE TEMPERATURE SENSED IS NORMAL IN RANGE AND HEATER IS TURNED ON. CUT-OFF OF HEATER IS CHECKED. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 9:23:0 | 1
123 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00008 | RADIANT WARMER | JOYTI CHOUDHARI | 7588768076 | 30636 | 31-Jul-2019 | 09:10 |  |  | 01-Aug-2019 | 16:45 | NOT WORKING | CHECK AND FOUND DISPLAY BOARD AND R3 BOARD IS DEFECTIVE. CHANGED THE OLD DISPLAY BOARD AND R3 BOARD WITH THE NEW ONE. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 31:35:0 | 2
124 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00008 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 16871 | 19-Jun-2018 | 08:02 |  |  | 20-Jun-2018 | 02:30 | NOT WORKING | CHECK AND FOUND ON / OFF SWITCH, EXAMINATION LIGHT NOT WORKING. REPLACED BOTH, NOW IT'S WORKING SATISFACTORY. | Closed | NO INFO | HIGH | 18:28:0 | 2
125 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00009 | RADIANT WARMER | JYOTI BULBULE | 94212718810 | 19565 | 24-Aug-2018 | 09:35 |  |  | 30-Aug-2018 | 12:00 | NOT WORKING | CHECK AND FOUND SIDE CRADLE IS BROKAN CHANGED THE NEW SIDE CRADLE.RADIANT WARMER IS NOW WORKING | Closed | RHW2101A10413/224 | HIGH | 146:25:0 | 7
126 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00009 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 5542 | 21-Jul-2017 | 09:58 |  |  | 22-Jul-2017 | 10:59 | NOT WORKING | REPLACE POWER CARD AND FUSE | Closed | RHW2101A10413/224 | HIGH | 25:1:0 | 2
127 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00009 | RADIANT WARMER | PRADEEP | 9423605762 | 808 | 07-Jan-2017 | 11:45 |  |  | 09-Jan-2017 | 12:45 | NOT WORKING | FUSE CHANGE, NOW MACHINE IS WORKING | Closed | RHW2101A10413/224 | HIGH | 49:0:0 | 3
128 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00009 | RADIANT WARMER | CHAUDHARI NURSE | 7588768076 | 13700 | 12-Mar-2018 | 08:54 |  |  | 15-Mar-2018 | 12:50 | NOT WORKING | REPLACED NEW BABY CRADLE BIG SUPPORT AND SIDE SUPPORT HOLDER.NOW CRADLE PROBLEM GET RECTIFIED, MACHINE IS WORKING CONDITION. | Closed | RHW2101A10413/224 | HIGH | 75:56:0 | 4
129 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00010 | RADIANT WARMER | PRADEEP | 9423605762 | 809 | 07-Jan-2017 | 12:00 |  |  | 10-Jan-2017 | 10:30 | NOT WORKING | FUSE CHANGED | Closed | RHW2101A10413/225 | HIGH | 70:30:0 | 4
130 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00010 | RADIANT WARMER | JYOTI CHOUDHARI | 7588768076 | 37601 | 07-Jan-2020 | 09:05 |  |  | 08-Jan-2020 | 15:10 | NOT WORKING | CHECK AND FOUND HEATER IS NOT TURNING 'ON'. CROSS-CHECKED THE SMPS BOARD, R3 BOARD, DISPLAY CONTROLLER BOARD & HEATER AND FOUND THAT POWER RESISTOR IS FAULTY.FIXED THE NEW POWER RESISTOR. CUT-OFF OF HEATER AT HIGH AND LOW BODY TEMPERATURE IS CHECKED. RADIANT WARMER IS NOW WORKING. | Closed | RHW2101A10413/225 | HIGH | 30:5:0 | 2
131 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00011 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 5540 | 21-Jul-2017 | 09:55 |  |  | 22-Jul-2017 | 14:06 | NOT WORKING | REPLACE NEW POWER CARD AND FUSE | Closed | RHW12101A/0413/220 | HIGH | 28:11:0 | 2
132 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00011 | RADIANT WARMER | JOYTI CHOUDHARI | 7588768076 | 10373 | 12-Dec-2017 | 09:32 |  |  | 16-Dec-2017 | 11:24 | NOT WORKING | REPLACED THE TRANSFORMER | Closed | RHW12101A/0413/220 | HIGH | 97:52:0 | 5
133 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00011 | RADIANT WARMER | PRADEEP WASNIK | 9423605761 | 835 | 09-Jan-2017 | 12:52 |  |  | 16-Jan-2017 | 10:05 | NOT WORKING | CHANGED POWER SUPPLY BOARD | Closed | RHW12101A/0413/220 | HIGH | 165:13:0 | 8
134 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00012 | RADIANT WARMER | PRADIP WASNIK | 9423605761 | 944 | 13-Jan-2017 | 10:05 |  |  | 13-Jan-2017 | 02:30 | NOT WORKING | CHANGED POWER SUPPLY BOARD | Closed | RHW12101A/0A131/1022 | HIGH | -7:-35:0 | 1
135 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00012 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 16870 | 19-Jun-2018 | 08:00 |  |  | 20-Jun-2018 | 04:30 | NOT WORKING | CHECK AND FOUND HEATER IS NOT WORKING. REPLACED HEATER, NOW IT'S WORKING. | Closed | RHW12101A/0A131/1022 | HIGH | 20:30:0 | 2
136 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00013 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 35854 | 28-Nov-2019 | 08:02 |  |  | 29-Nov-2019 | 11:09 | NOT WORKING | CHECK ANF FOUND DISPLAY IS NOT WORKING. FIXED THE LOOSE CONNECTION THROUGH DISPLAY CONTROLLER BOARD. DISPLAY IS NOW WORKING. CUT-OFF OF HEATER IS CHECKED. RADIANT WARMER IS NOW WORKING. | Closed | RHW12101A13/1092 | HIGH | 27:7:0 | 2
137 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00014 | RADIANT WARMER | CHOUDHARI NURSE | 7588768076 | 18062 | 18-Jul-2018 | 08:17 |  |  | 18-Jul-2018 | 15:15 | NOT WORKING | CHECK AND FOUND HEATER CONNECTOR IS LOOSE. FIX THE CONNECTOR PROBLEM. NOW WARMER WORKING SATISFACTORY. | Closed | RHW12101A10413/225 | HIGH | 6:58:0 | 1
138 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00017 | SYRINGE PUMP | JYOTI CHAUDHARI | 7588768076 | 35853 | 28-Nov-2019 | 08:02 |  |  | 29-Nov-2019 | 11:16 | NOT WORKING | CHECK AND FOUND POWER CORD IS FAULTY. REPLACE THE OLD POWER CORD WITH THE NEW ONE. SYRINGE PUMP IS NOW WORKING. | Closed | 1004134.5EX1ST09 | MEDIUM | 27:14:0 | 2
139 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00018 | SYRINGE PUMP | JYOTI BULBULE | 94212718810 | 19563 | 24-Aug-2018 | 09:35 |  |  | 27-Aug-2018 | 12:30 | NOT WOKING | CHECK AND FOUND SYRINGE PUMP IS WORKING | Closed | 150216B4.5RY02 | MEDIUM | 74:55:0 | 4
140 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00019 | SYRINGE PUMP | JOYTI CHOUDHARI | 7588768076 | 13788 | 14-Mar-2018 | 09:30 |  |  | 14-Mar-2018 | 14:55 | NOT WORKING | REPLACED NEW 3A FUSE, NOW MACHINE IS WORKING CONDITION | Closed | 150216B4.5AJ02Z | MEDIUM | 5:25:0 | 1
141 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00019 | SYRINGE PUMP | JYOTI CHAUDHARI | 7588768076 | 7549 | 26-Sep-2017 | 09:45 |  |  | 27-Sep-2017 | 16:00 | NOT WORKING | REPLACED NEW 3AMPS FUSE AND OCCLUSION PROCESS IS DONE | Closed | 150216B4.5AJ02Z | MEDIUM | 30:15:0 | 2
142 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00020 | SYRINGE PUMP | JOYTI CHOUDHARI | 7588768076 | 13789 | 14-Mar-2018 | 09:30 |  |  | 14-Mar-2018 | 14:58 | NOT WORKING | SOLDERED THE PLUNGER WIRE NOW ERROR PROBLEM GET RECTIFIED, MACHINE IS WORKING CONDITION. | Closed | 150216B4.5DY02Z | MEDIUM | 5:28:0 | 1
143 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00020 | SYRINGE PUMP | Jyoti Chaudhari | 7588768076 | 42390 | 16-May-2020 | 09:59 |  |  | 16-May-2020 | 11:22 | NOT WORKING | CHECK AND FOUND PLUNGER SET IS BROKEN. NEW PLUNGER SET IS BEING FIXED. MACHINE IS CHECKED AT DIFFERENT RATES. SYRINGE PUMP IS NOW WORKING. | Closed | 150216B4.5DY02Z | MEDIUM | 1:23:0 | 1
144 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00021 | PULSE OXIMETER LINE POWERED | JYOTI BULBULE | PULSE OXIMETER | 19564 | 24-Aug-2018 | 09:35 |  |  | 27-Aug-2018 | 12:00 | NOT WORKING | CHECK AND FOUND PULSE OXIMETER IS WORKING | Closed | 413186 | MEDIUM | 74:25:0 | 4
145 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00021 | PULSE OXIMETER LINE POWERED | PATIL NURSE | 7758058770 | 7684 | 29-Sep-2017 | 09:42 |  |  | 03-Oct-2017 | 22:07 | NOT WORKING | REPLACED THE NEW SPO2 PROBE | Closed | 413186 | MEDIUM | 108:25:0 | 5
146 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00022 | PULSE OXIMETER LINE POWERED | JYOTI CHAUDHARI | 7588768076 | 14951 | 13-Apr-2018 | 11:15 |  |  | 14-Apr-2018 | 03:46 | NOT WORKING | REPLACED NEW PROBE | Closed | 4131863 | MEDIUM | 16:31:0 | 2
147 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00022 | PULSE OXIMETER LINE POWERED | CHOUDHARI SISTER | 8788509702 | 16662 | 13-Jun-2018 | 08:41 |  |  | 14-Jun-2018 | 14:13 | NOT WORKING | REINSERTED THE POWER CHORD RIGHTLY. THE MACHINE IS WORKING | Closed | 4131863 | MEDIUM | 29:32:0 | 2
148 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00022 | PULSE OXIMETER LINE POWERED | JYOTI CHAUDHARI | 7588768076 | 24429 | 30-Jan-2019 | 08:05 |  |  | 01-Feb-2019 | 09:30 | NOT WORKING | CHECK AND FOUND POWER SUPPLY ISSUE IS THERE. FIXED THE POWER SUPPLY ISSUE. PULSE OXIMETER IS NOW WORKING. | Closed | 4131863 | MEDIUM | 49:25:0 | 3
149 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00023 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 10034 | 04-Dec-2017 | 09:50 |  |  | 05-Dec-2017 | 11:41 | NOT WORKING | REPLACED UV TUBE | Closed | NO INFO | HIGH | 25:51:0 | 2
150 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00023 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 14952 | 13-Apr-2018 | 11:15 |  |  | 16-Apr-2018 | 12:19 | NOT WORKING | REPLACED WITH NEW 18 WATTS BLUE TUBE LIGHT,MACHINE IS WORKING CONDITION | Closed | NO INFO | HIGH | 73:4:0 | 4
151 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00024 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 37573 | 06-Jan-2020 | 08:29 |  |  | 07-Jan-2020 | 14:04 | NOT WORKING | CHECK AND FOUND FEMALE CABLE CONNECTOR IS BURNT AND DUE TO WHICH THERE IS POWER SUPPLY FAILURE ISSUE. FIXED THE NEW FEMALE CABLE CONNECTOR. POWER SUPPLY ISSUE RESOLVED. RADIANT WARMER IS NOW WORKING. | Closed | 4222 | HIGH | 29:35:0 | 2
152 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00025 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 37560 | 06-Jan-2020 | 08:20 |  |  | 06-Jan-2020 | 16:43 | NOT WORJKING | CHECK AND FOUND HEATER ROD WIRE IS OPEN AND DUE TO WHICH HEATER COULD NOT TURN 'ON'. FIXED THE HEATER ROD WIRE. NOW, HEATER IS WORKING AND CUT-OFF OF HEATER AT LOW AND HIGH BODY TEMPERATURE IS CHECKED AND FOUND TO BE OK. | Closed | 4228 | HIGH | 8:23:0 | 1
153 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00032 | SYRINGE PUMP | SCOTT NURSE | 9764971169 | 32400 | 13-Sep-2019 | 08:05 |  |  | 14-Sep-2019 | 15:21 | NOT WORKING | CHECK AND FOUND POWER ON/OFF BUTTON IS NOT WORKING. FIXED THE PROBLEM OF POWER 'ON/OFF' BUTTON. SYRINGE PUMP IS NOW WORKING. | Closed | 904134.5EX1ST09 | MEDIUM | 31:16:0 | 2
154 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00032 | SYRINGE PUMP | CHOUDHARI SISTER | 8788509702 | 16660 | 13-Jun-2018 | 08:37 |  |  | 13-Jun-2018 | 14:15 | NOT WORKING | REPLACED THE FUSE 3AMP | Closed | 904134.5EX1ST09 | MEDIUM | 5:38:0 | 1
155 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00032 | SYRINGE PUMP | CHAUDARY NURSE | 7588768076 | 6979 | 12-Sep-2017 | 09:42 |  |  | 13-Sep-2017 | 19:29 | NOT WORKING | RESET THE SETTING AND GUIDED THE TECHNICIAN | Closed | 904134.5EX1ST09 | MEDIUM | 33:47:0 | 2
156 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00032 | SYRINGE PUMP | JYOTI CHAUDHARI | 7588768076 | 24428 | 30-Jan-2019 | 08:00 |  |  | 01-Feb-2019 | 09:05 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE IS THERE. REPAIRED THE SMPS OF THE MACHINE.SYRINGE PUMP IS NOW WORKING. | Closed | 904134.5EX1ST09 | MEDIUM | 49:5:0 | 3
157 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00032 | SYRINGE PUMP | SCOTT NURSE | 9764971169 | 41513 | 07-Apr-2020 | 08:10 |  |  | 09-Apr-2020 | 11:50 | NOT WORKING | CHECK AND FOUND ALARM OF PLUNGER DISENGAGE POPS UP DURING STARTING THE DOSE. FIXED THE SPANNER AND DRAW ROD. PROBLEM IS RESOLVED. SYRINGE PUMP IS NOW WORKING. | Closed | 904134.5EX1ST09 | MEDIUM | 51:40:0 | 3
158 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00032 | SYRINGE PUMP | JYOTI BULBULE | 94212718810 | 19562 | 24-Aug-2018 | 09:35 |  |  | 27-Aug-2018 | 01:15 | NOT WORING | CHECK AND FOUND SYRING PUMP IS WORKING | Closed | 904134.5EX1ST09 | MEDIUM | 63:40:0 | 4
159 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00033 | SYRINGE PUMP | CHOUDHARI SISTER | 8788509702 | 16661 | 13-Jun-2018 | 08:38 |  |  | 16-Jun-2018 | 15:50 | NOT WORKING | CHECK AND FOUND MOTHER BOARD IS FAULTY, ALSO SYRINGE HOLDER IS MISSING AND PLUNGER NOT WORKING. | Closed | 904134.5EXAST07 | MEDIUM | 79:12:0 | 4
160 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00035 | RADIANT WARMER | JYOTI CHOUDHARI | 7588768076 | 37602 | 07-Jan-2020 | 09:05 |  |  | 07-Jan-2020 | 17:08 | NOT WORKING | CHECK AND FOUND TEMPERATURE SENSED BY AIR PROBE IS HIGH AND DUE TO WHICH HEATER COULD NOT TURN 'ON'. FIXED THE AIR PROBE FEMALE CONNECTOR THROUGH DISPLAY CONTROLLER BOARD. NOW, TEMPERATURE SENSED BY AIR PROBE IS NORMAL AND HEATER IS TURNING 'ON'. CUT-OFF OF HEATER IS CHECKED. RADIANT WARMER IS NOW WORKING. | Closed | 1082 | HIGH | 8:3:0 | 1
161 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00035 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 7550 | 26-Sep-2017 | 09:45 |  |  | 27-Sep-2017 | 16:54 | NOT WORKING | REPLACED NEW 3AMPS FUSE, NOW MACHINE IS WORKING CONDITION | Closed | 1082 | HIGH | 31:9:0 | 2
162 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00036 | RADIANT WARMER | JOYTI CHOUDHARI | 7588768076 | 16733 | 14-Jun-2018 | 09:14 |  |  | 14-Jun-2018 | 11:50 | NOT WORKING | REPLACED THE FUSE 3AMP | Closed | 221 | HIGH | 2:36:0 | 1
163 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00036 | RADIANT WARMER | CHAVAN SISTER | 7588768076 | 21444 | 17-Oct-2018 | 08:32 |  |  | 23-Oct-2018 | 15:34 | NOT WORKING | CHECK AND FOUND CONTROLLER BOARD PROBLEM. REPAIRED THE CONTROLLER BOARD. RADIANT WARMER IS NOW WORKING. | Closed | 221 | HIGH | 151:2:0 | 7
164 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00037 | RADIANT WARMER | CHAVAN SISTER | 7588768076 | 21445 | 17-Oct-2018 | 08:32 |  |  | 23-Oct-2018 | 15:07 | NOT WORKING | CHECK AND FOUND CONTROLLER BOARD PROBLEM. REPAIRED THE CONTROLLER BOARD. RADIANT WARMER IS NOW WORKING. | Closed | 1086 | HIGH | 150:35:0 | 7
165 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00038 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 10033 | 04-Dec-2017 | 09:50 |  |  | 05-Dec-2017 | 11:37 | NOT WORKING | RECONNECTED ALL DISPLAY CONNECTION | Closed | 212 | HIGH | 25:47:0 | 2
166 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00038 | RADIANT WARMER | JOYTI CHOUDHARI | 7588768076 | 30637 | 31-Jul-2019 | 09:10 |  |  | 31-Jul-2019 | 16:45 | NOT WORKING | CHECK AND FOUND SKIN AND AIR PROBE IS SENSING TEMPERATURE AUTOMATICALLY ABOVE 43 DEGREE CELSIUS AND HEATER COULD NOT TURN ON. PROBLEM FOUND IS FROM DISPLAY CONTROLLER BOARD. FIXED THE PROBLEM FROM DISPLAY CONTROLLER BOARD. NOW, BOTH THE SENSORS ARE WORKING AND SENSING TEMPERATURE DESIRABLY. RADIANT WARMER IS NOW WORKING. | Closed | 212 | HIGH | 7:35:0 | 1
167 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00038 | RADIANT WARMER | SM KHAIRKAR | 8055208992 | 40307 | 02-Mar-2020 | 08:20 |  |  | 02-Mar-2020 | 15:46 | NOT WORKING | CHECK AND FOUND DIODE OF THE R3 BOARD IS SHORT AND DUE TO WHICH IT ALWAYS BLOWS AWAY THE FUSE WHEN POWER IS TURNED 'ON'. FIXED THE NEW R3 BOARD. POWER SUPPLY FAILURE ISSUE RESOLVED. RADIANT WARMER IS NOW WORKING. | Closed | 212 | HIGH | 7:26:0 | 1
168 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00040 | RADIANT WARMER | JOYTI CHOUDHARI | 7588768076 | 10372 | 12-Dec-2017 | 09:32 |  |  | 16-Dec-2017 | 11:26 | NOT WORKING | REPLACED THE TRANSFORMER | Closed | 1091 | HIGH | 97:54:0 | 5
169 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00040 | RADIANT WARMER | JYOTI CHOUDHARI | 7388768076 | 38453 | 24-Jan-2020 | 09:11 |  |  | 25-Jan-2020 | 11:08 | NOT WORKING | CHECK AND FOUND R3 BOARD & POWER RESISTOR IS FAULTY. HEATER CABLE CONNECTOR ALONG WITH FEMALE CONNECTOR IS BURNT. FIXED THE NEW R3 BOARD AND POWER RESISTOR. ALSO NEW HEATER CABLE CONNECTOR ALONG WITH FEMALE CONNECTOR IS FIXED. CUT-OFF OF HEATER IS CHECKED. RADIANT WARMER IS NOW WORKING. | Closed | 1091 | HIGH | 25:57:0 | 2
170 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00040 | RADIANT WARMER | SM KHAIRKAR | 8055208992 | 40308 | 02-Mar-2020 | 08:20 |  |  | 02-Mar-2020 | 16:51 | NOT WORKING | CHECK AND FOUND DISPLAY CONTROLLER BOARD IS FAULTY. R3 BOARD IS ALSO FAULTY AS CUT-OFF OF HEATER AT SET TEMPERATURE IS NOT TAKING PLACE. FIXED THE NEW DISPLAY CONTROLLER BOARD AND R3 BOARD. RADIANT WARMER IS NOW WORKING WITH CUT-OFF OF HEATER AT SET TEMPERATURE. | Closed | 1091 | HIGH | 8:31:0 | 1
171 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00040 | RADIANT WARMER | JYOTI CHAVDHARI | 7588768076 | 22451 | 28-Nov-2018 | 09:12 |  |  | 04-Dec-2018 | 12:57 | NOT WORKING | CHECK AND FOUND CONTROLLER BOARD IS DEFECTIVE. REPAIRED THE CONTROLLER BOARD. RADIANT WARMER IS NOW WORKING. | Closed | 1091 | HIGH | 147:45:0 | 7
172 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00040 | RADIANT WARMER | JOYTI CHOUDHARI | 7588768076 | 16734 | 14-Jun-2018 | 09:15 |  |  | 16-Jun-2018 | 03:45 | NOT WORKING | CHECK AND FOUND FUSE IS NOT WORKING. REPLACED FUSE, NOW IT'S WORKING. | Closed | 1091 | HIGH | 42:30:0 | 3
173 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00041 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 35996 | 30-Nov-2019 | 08:55 |  |  | 30-Nov-2019 | 22:35 | NOT WORKING | CHECK AND FOUND FEMALE CABLE CONNECTOR IS BURNT. FIXED THE NEW FEMALE CABLE CONNECTOR. CUT-OFF OF HEATER AT LOW AND HIGH BODY TEMPERATURE IS CHECKED AND ITS OK. RADIANT WARMER IS NOW WORKING. | Closed | 4230 | HIGH | 13:40:0 | 1
174 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00041 | RADIANT WARMER | JYOTI CHAVDHARI | 7588768076 | 22452 | 28-Nov-2018 | 09:12 |  |  | 04-Dec-2018 | 14:00 | NOT WORKING | CHECK AND FOUND CONTROLLER BOARD IS DEFECTIVE.ALSO POWER CORD IS NOT WORKING. REPAIRED THE CONTROLLER BOARD. CHANGED THE NEW POWER CORD. RADIANT WARMER IS NOW WORKING. | Closed | 4230 | HIGH | 148:48:0 | 7
175 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00041 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 5541 | 21-Jul-2017 | 09:58 |  |  | 22-Jul-2017 | 10:02 | NOT WORKING | REPLACE POWER CARD AND FUSE | Closed | 4230 | HIGH | 24:4:0 | 2
176 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00042 | PULSE OXIMETER LINE POWERED | PATIL NURSE | 7758058770 | 7683 | 29-Sep-2017 | 09:41 |  |  | 03-Oct-2017 | 22:11 | NOT WORKING | REPLACED THE 2A FUSE | Closed | 4131859 | MEDIUM | 108:30:0 | 5
177 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00042 | PULSE OXIMETER LINE POWERED | JYOTI CHAUDHARI | 7588768076 | 24430 | 30-Jan-2019 | 08:05 |  |  | 01-Feb-2019 | 09:15 | NOT WORKING | CHECK AND FOUND D-TYPE SPO2 PROBE IS NOT WORKING. CHANGED THE OLD WITH NEW D-TYPE SPO2 NEONATAL PROBE.PULSE OXIMETER IS NOW WORKING. | Closed | 4131859 | MEDIUM | 49:10:0 | 3
178 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00042 | PULSE OXIMETER LINE POWERED | CHOUDHARI SISTER | 8788509702 | 16663 | 13-Jun-2018 | 08:42 |  |  | 13-Jun-2018 | 17:57 | NOT WORKING | NO PROBLEM THE MACHINE IS FINE | Closed | 4131859 | MEDIUM | 9:15:0 | 1
179 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00044 | SYRINGE PUMP | JYOTI CHAUDHARI | 7588768076 | 24427 | 30-Jan-2019 | 08:00 |  |  | 01-Feb-2019 | 10:12 | NOT WORKING | CHECK AND FOUND POWER SUPPLY ISSUE IS THERE. REPAIRED THE SMPS OF THE MACHINE. SYRINGE PUMP IS NOW WORKING. | Closed | 120216B4.5DY03Z | MEDIUM | 50:12:0 | 3
180 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00044 | SYRINGE PUMP | JYOTI CHAUDHARI | 7588768076 | 35852 | 28-Nov-2019 | 08:02 |  |  | 29-Nov-2019 | 17:14 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE ISSUE. FUSE IS BURNT AND REPLACE WITH THE NEW ONE. SMPS IS CHECKED. POWER SUPPLY ISSUE IS RESOLVED. SYRINGE PUMP IS NOW WORKING. | Closed | 120216B4.5DY03Z | MEDIUM | 33:12:0 | 2
181 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00045 | SYRINGE PUMP | CHAUDARY NURSE | 7588768076 | 6981 | 12-Sep-2017 | 09:42 |  |  | 13-Sep-2017 | 19:25 | NOT WORKING | RESET THE SETTING AND GUIDED THE TECHNICIANS | Closed | 0904135.5EX1ST08 | MEDIUM | 33:43:0 | 2
182 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00046 | PHOTOTHERAPHY UNIT | JYOTI CHOUDHARI | 7588768076 | 38010 | 16-Jan-2020 | 09:18 |  |  | 16-Jan-2020 | 16:19 | NOT WORKING | CHECK AND FOUND ALL THE TUBELIGHT IS FUSED AND STARTER IS NOT WORKING. FIXED THE OLD TUBELIGHT AND STARTER WITH THE NEW ONE. PHOTOTHERAPY UNIT IS NOW WORKING. | Closed | PH306116 | HIGH | 7:1:0 | 1
183 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00046 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 10036 | 04-Dec-2017 | 09:50 |  |  | 05-Dec-2017 | 11:44 |  | REPLACED UV TUBE | Closed | PH306116 | HIGH | 25:54:0 | 2
184 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00046 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 6783 | 08-Sep-2017 | 08:05 | 09-Sep-2017 | 19:33 | 11-Sep-2017 | 19:46 | NOT WORKING | REPLACE THE 18W FLUORESCENT LAMP | Closed | PH306116 | HIGH | 83:41:0 | 4
185 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00047 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 6782 | 08-Sep-2017 | 08:05 | 09-Sep-2017 | 19:32 | 11-Sep-2017 | 12:08 | NOT WORKING | REPLACE 18W FLUORESCENT LAMP | Closed | PH306164 | HIGH | 76:3:0 | 4
186 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00047 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 14953 | 13-Apr-2018 | 11:15 |  |  | 16-Apr-2018 | 11:17 | NOT WORKING | REPLACED WITH NEW 18 WATTS BLUE TUBE LIGHT, MACHINE IS WORKING CONDITION. | Closed | PH306164 | HIGH | 72:2:0 | 4
187 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00052 | RADIANT WARMER | CHOUDHARI NURSE | 7588768076 | 18063 | 18-Jul-2018 | 08:17 |  |  | 18-Jul-2018 | 16:35 | NOT WORKING | CHECK AND FOUND WARMER PUT IN MANUAL MOOD. PUT THE WARMER IS AUTO MODE. NOW WARMER WORKING SATISFACTORY. | Closed | 4232 | HIGH | 8:18:0 | 1
188 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00059 | SHADOWLESS LAMP STAND TYPE | SNEHA GAZABE | 8605158010 | 37785 | 10-Jan-2020 | 09:27 |  |  | 10-Jan-2020 | 11:54 | NOT WORKING | CHECK AND FOUND HALOGEN LIGHT IS FUSED. FIXED THE NEW HALOGEN LIGHT OF RATING 12V,50W. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 2:27:0 | 1
189 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00061 | SHADOWLESS LAMP STAND TYPE | SNEHA GAZABE | 8605158010 | 37786 | 10-Jan-2020 | 09:27 |  |  | 10-Jan-2020 | 11:56 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. FIXED THE NEW HALOGEN BULB OF RATING 12V,50W. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 2:29:0 | 1
190 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00062 | SHADOWLESS LAMP STAND TYPE | SNEHA GAZABE | 8605158010 | 37787 | 10-Jan-2020 | 09:27 |  |  | 10-Jan-2020 | 13:57 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. FIXED THE NEW HALOGEN BULB OF RATING 12V,50W. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 4:30:0 | 1
191 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00064 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 6784 | 08-Sep-2017 | 08:05 | 09-Sep-2017 | 19:33 | 11-Sep-2017 | 19:49 | NOT WORKING | REPLACE THE 18W FLUORESCENT LAMP | Closed | NO INFO | HIGH | 83:44:0 | 4
192 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00064 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 10035 | 04-Dec-2017 | 09:50 |  |  | 05-Dec-2017 | 11:46 | NOT WORKING | REPLACED UV TUBE | Closed | NO INFO | HIGH | 25:56:0 | 2
193 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00065 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 6785 | 08-Sep-2017 | 08:05 | 09-Sep-2017 | 19:35 | 11-Sep-2017 | 19:42 | NOT WORKING | REPLACE THE NEW 18W FLUORESCENT LAMP | Closed | PH306145 | HIGH | 83:37:0 | 4
194 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00065 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 10039 | 04-Dec-2017 | 09:50 |  |  | 05-Dec-2017 | 11:57 | NOT WORKING | REPLACED UV TUBE | Closed | PH306145 | HIGH | 26:7:0 | 2
195 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00066 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 10038 | 04-Dec-2017 | 09:50 |  |  | 05-Dec-2017 | 11:54 | NOT WORKING | REPLACED UV TUBE | Closed | PH2K70600/30 | HIGH | 26:4:0 | 2
196 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00066 | PHOTOTHERAPHY UNIT | JYOTI | 8588768076 | 31862 | 27-Aug-2019 | 09:12 |  |  | 29-Aug-2019 | 15:30 | NOT WORKING | CHECK AND FOUND 1 CHOKE IS BURNT AND DUE TO WHICH IT ALWAYS MAKE THE TUBELIGHT FUSE. ALSO 3 TUBELIGHT IS DEFECTIVE. CHANGED THE OLD DEFECTIVE CHOKE AND TUBELIGHT WITH THE NEW ONE. PHOTOTHERAPY UNIT IS NOW WORKING. | Closed | PH2K70600/30 | HIGH | 54:18:0 | 3
197 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00066 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 6786 | 08-Sep-2017 | 08:05 | 09-Sep-2017 | 19:35 | 11-Sep-2017 | 19:38 | NOT WORKING | REPLACE 18W FLUORESCENT LAMP | Closed | PH2K70600/30 | HIGH | 83:33:0 | 4
198 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00069 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | RAJANI PATIL | 8698396820 | 16992 | 20-Jun-2018 | 09:57 |  |  | 27-Jun-2018 | 15:55 | NOT WORKING | CHECK AND FOUND BATTERY PROBLEM, REPLACED BATTERY. NOW MACHINE WORKING SATISFACTORY. | Closed | NO INFO | MEDIUM | 173:58:0 | 8
199 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00070 | BP APPARATUS ANEROID | MANJUSHA NIVARE | 9423646267 | 11233 | 03-Jan-2018 | 08:25 |  |  | 04-Jan-2018 | 11:41 | NOT WORKING | REPLACED THE BLADDER AND TUBING | Closed | 76253 | LOW | 27:16:0 | 2
200 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00070 | BP APPARATUS ANEROID | DORTHI SCOTT | 9764971169 | 23665 | 07-Jan-2019 | 08:10 |  |  | 08-Jan-2019 | 08:05 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK. CHANGED THE NEW BLADDER. B. P APPARATUS IS NOW WORKING. | Closed | 76253 | LOW | 23:55:0 | 2
201 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00070 | BP APPARATUS ANEROID | SCOTT NURSE | 9764971169 | 32798 | 20-Sep-2019 | 08:40 |  |  | 23-Sep-2019 | 14:56 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT. THERE IS LEAKAGE IN VALVE AND THE CONNECTING TUBE. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B.P APPARATUS IS NOW WORKING. | Closed | 76253 | LOW | 78:16:0 | 4
202 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00070 | BP APPARATUS ANEROID | AKSHAY SAVARKAR | 8888591294 | 17610 | 06-Jul-2018 | 09:56 |  |  | 06-Jul-2018 | 15:45 | NOT WORKING | CHECK AND FOUND BULBS PUNCHER. REPLACED BULB AS SAME, NOW WORKING SATISFACTORY. | Closed | 76253 | LOW | 5:49:0 | 1
203 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00070 | BP APPARATUS ANEROID | PRADEEP WASNIK | 9423605761 | 1347 | 01-Feb-2017 | 10:30 |  |  | 03-Feb-2017 | 10:30 | NOT WORKING | MERCURY REPLACED IN BP APPARATUS | Closed | 76253 | LOW | 48:0:0 | 3
204 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00071 | ECG UNIT 3 CHANNEL PORTABLE | SCOTT NURSE | 9764971169 | 29098 | 26-Jun-2019 | 09:00 |  |  | 27-Jun-2019 | 16:41 | NOT WORKING | CHECK AND FOUND PRINT IS NOT COMING AS MODE IS CHANGED. MODE AND SETTING IS CHECKED AND RESET THE INITIAL SETTING. NOW, PRINT IS COMING AS USUAL. E. C. G MACHINE IS NOW WORKING | Closed | DDTD2E5785 | MEDIUM | 31:41:0 | 2
205 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00071 | ECG UNIT 3 CHANNEL PORTABLE | UJWALA SISTER | 8530322606 | 14950 | 16-Apr-2018 | 09:07 |  |  | 18-Apr-2018 | 09:53 | NOT WORKING | REPLACED NEW PROBE | Closed | DDTD2E5785 | MEDIUM | 48:46:0 | 3
206 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00071 | ECG UNIT 3 CHANNEL PORTABLE | MANISHA NURSE | 9423646267 | 16086 | 24-May-2018 | 09:05 |  |  | 29-May-2018 | 10:37 | NOT WORKING | PRINTER COVER IS CRACKED. REPLACED WITH NEW PRINTER COVER. NOW THE MACHINE IS IN WORKING | Closed | DDTD2E5785 | MEDIUM | 121:32:0 | 6
207 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00071 | ECG UNIT 3 CHANNEL PORTABLE | S.R.SHINDE | 9373476576 | 30092 | 18-Jul-2019 | 09:10 | 18-Jul-2019 | 17:45 | 20-Jul-2019 | 15:35 | LEAD ISSUE | CHECK AND FOUND CLAMP IS BROKEN AND BULB IS WORN OUT. CHANGED THE OLD PART WITH THE NEW ONE. E.C.G MACHINE IS NOW WORKING. | Closed | DDTD2E5785 | MEDIUM | 54:25:0 | 3
208 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00071 | ECG UNIT 3 CHANNEL PORTABLE | UJAWALA GHOKHALE | 8530322606 | 11151 | 30-Dec-2017 | 09:15 |  |  | 30-Dec-2017 | 21:10 | NOT WORKING | REPLACED THE ECG ELECTRODE AND BULB | Closed | DDTD2E5785 | MEDIUM | 11:55:0 | 1
209 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00071 | ECG UNIT 3 CHANNEL PORTABLE | LAXMI APURKAR | 9423417228 | 7845 | 05-Oct-2017 | 09:17 |  |  | 07-Oct-2017 | 19:02 | NOT WORKING | REPLACED THE POWER CABLE | Closed | DDTD2E5785 | MEDIUM | 57:45:0 | 3
210 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00071 | ECG UNIT 3 CHANNEL PORTABLE | TRUPTI BOBHATE | 8600356619 | 15693 | 11-May-2018 | 09:00 |  |  | 11-May-2018 | 15:45 |  | ONE PLASTIC ELECTRODE IS BROKEN, REPLACED WITH NEW ONE NOW MACHINE WORKING FINE | Closed | DDTD2E5785 | MEDIUM | 6:45:0 | 1
211 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00071 | ECG UNIT 3 CHANNEL PORTABLE | DOIIY SCOTT | 9464971169 | 33890 | 15-Oct-2019 | 09:06 |  |  | 15-Oct-2019 | 16:38 | NOT WORKING | CHECK AND FOUND ECG GRAPH OBTAINED IS SIMPLY DASH LINE. CHECKED THE ECG CABLE ON DIFFERENT ECG MACHINE AND FOUND ECG CABLE IS DEFECTIVE. CHANGED THE DEFECTIVE ECG CABLE WITH THE NEW ONE. NOW, GRAPH OBTAINED IS SATISFACTORY. ECG MACHINE IS NOW WORKING. | Closed | DDTD2E5785 | MEDIUM | 7:32:0 | 1
212 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00072 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | D. V Scott | 9764971169 | 43871 | 06-Jul-2020 | 08:22 |  |  | 06-Jul-2020 | 11:45 | NOT WORKING | PROBLEM REPORTED OF LOUD SOUND WHEN EXTERNAL COMPRESSOR SYSTEM IS CONNECTED TO VENTILATOR DURING ANY MODES. FOUND HINGE SUPPORTING THE COMPRESSOR MOTOR IS BROKEN AND DUE TO WHICH IT MAKES LOUD SOUND WHEN TURNED 'ON'. FIXED THE HINGE OF THE MOTOR AND CLEANED THE FILTER. EXTERNAL COMPRESSOR SYSTEM IS NOW WORKING WITH NO DISTURBING SOUND. | Closed | ASED0123 | HIGH | 3:23:0 | 1
213 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00072 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | UJWALA POKALE | 8530322606 | 4902 | 29-Jun-2017 | 09:45 | 30-Jun-2017 | 22:04 | 03-Jul-2017 | 14:46 | NOT WORKING | REPLACE THE STAND BY COMPRESSOR | Closed | ASED0123 | HIGH | 101:1:0 | 5
214 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00072 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | Mrs. D V Scott | 9764971169 | 42247 | 14-May-2020 | 09:20 |  |  | 14-May-2020 | 11:15 | NOT WORKING | CHECK AND FOUND CALIBRATION OF OXYGEN FAILS AS OXYGEN SENSOR IS DEFECTIVE. FIXED THE NEW OXYGEN SENSOR. CALIBRATION OF OXYGEN IS SUCCESSFULLY DONE WITH MEASURED AMOUNT OF OXYGEN IN VARIOUS MODES IS NOW DISPLAYED IN FIO2 COLUMN. VENTILATOR IS CHECKED IN ALL MODES. VENTILATOR IS NOW WORKING. | Closed | ASED0123 | HIGH | 1:55:0 | 1
215 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00074 | MONITOR | LAXMI | 9423417228 | 8640 | 01-Nov-2017 | 10:00 |  |  | 02-Nov-2017 | 14:12 | NOT WORKING | REPLACED NEW NIBP CUFF,NOW MACHINE IS WORKING CONDITION. | Closed | 068E00 4546 | MEDIUM | 28:12:0 | 2
216 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00074 | MONITOR | PRADIP WASNIK | 9423605761 | 1565 | 09-Feb-2017 | 09:58 |  |  | 10-Feb-2017 | 14:35 | NOT WORKING | WAITING FOR SPARES PARTS | Closed | 068E00 4546 | MEDIUM | 28:37:0 | 2
217 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00074 | MONITOR | UJWALA POKALE | 8530322606 | 5201 | 11-Jul-2017 | 09:00 | 14-Jul-2017 | 10:58 | 17-Jul-2017 | 19:03 | NOT WORKING | REPLACE SPO2 CABLE | Closed | 068E00 4546 | MEDIUM | 154:3:0 | 7
218 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00074 | MONITOR | SCOTT NURSE | 9764971169 | 32977 | 24-Sep-2019 | 09:12 |  |  | 25-Sep-2019 | 15:17 | NOT WORKING | CHECK AND FOUND BLOOD PRESSURE COULD NOT BE MEASURED AS THERE IS LEAKAGE IN THE NIBP CUFF. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. NOW, B. P IS ACCURATELY MEASURED. MONITOR IS NOW WORKING. | Closed | 068E00 4546 | MEDIUM | 30:5:0 | 2
219 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00074 | MONITOR | SCOTT NURSE | 9764971169 | 37857 | 13-Jan-2020 | 08:05 |  |  | 14-Jan-2020 | 10:42 | NOT WORKING | CHECK AND FOUND B.P COULD NOT BE MONITORED AS NIBP CUFF IS HAVING LEAKAGE. FIXED THE NEW NIBP CUFF. B.P IS NOW MONITORED. MONITOR IS NOW WORKING. | Closed | 068E00 4546 | MEDIUM | 26:37:0 | 2
220 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00074 | MONITOR | SCOTT NURSE | 9764971169 | 30497 | 27-Jul-2019 | 09:30 | 27-Jul-2019 | 14:33 | 29-Jul-2019 | 15:34 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS NOT WORKING. CHANGED THE OLD DEFECTIVE PROBE WITH THE NEW ONE. MONITOR IS NOW WORKING. | Closed | 068E00 4546 | MEDIUM | 54:4:0 | 3
221 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00074 | MONITOR | UJAWALA GHOKHALE | 8530322606 | 11148 | 30-Dec-2017 | 09:15 |  |  | 31-Dec-2017 | 18:41 | NOT WORKING | REPLACED THE NIBP CUFF | Closed | 068E00 4546 | MEDIUM | 33:26:0 | 2
222 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00074 | MONITOR | LAXMI APURKAR | 9423417228 | 19422 | 21-Aug-2018 | 09:00 |  |  | 06-Nov-2018 | 11:15 | NOT WORKING | CHECK AND FOUND MONITOR SPO2 PROBE IS DEFECTIVE . CHANGE THE SAME SPO2 PROB NOW MACHINE IS WORKING | Closed | 068E00 4546 | MEDIUM | 1850:15:0 | 78
223 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00074 | MONITOR | SCOTT NURSE | 9764971169 | 21939 | 10-Nov-2018 | 09:02 |  |  | 13-Nov-2018 | 13:37 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING. PROBLEM IS BEING REPORTED OF HANGING OF MONITOR BUT CHECKED AND ITS BEING CONFIRMED MACHINE IS WORKING. | Closed | 068E00 4546 | MEDIUM | 76:35:0 | 4
224 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00075 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | UJWALA POKLEY | 8530322606 | 5185 | 10-Jul-2017 | 10:05 |  |  | 15-Jul-2017 | 13:12 | NOT WORKING | KEYPAD LOCK WAS UNLOCK AND SWITCH OFF AND ON FUNCTION PROPERLY WORKING | Closed | ASJC0120 | HIGH | 123:7:0 | 6
225 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00075 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | UJWALA KHOKARE | 8530322606 | 4874 | 28-Jun-2017 | 10:10 |  |  | 29-Jun-2017 | 17:09 | NOT WORKING | RECOMMENDED TO PURCHASE NEW UPS | Closed | ASJC0120 | HIGH | 30:59:0 | 2
226 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00075 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | SCOTT NURSE | 9764971169 | 41429 | 03-Apr-2020 | 08:00 |  |  | 21-Apr-2020 | 17:58 | NOT WORKING | CHECK AND FOUND MACHINE AUTOMATICALLY SHUTS DOWN WHEN PATIENT IS ON VENTILATION AND IT RESTARTS IN SAME MODE. IT ALWAYS HAPPENS IN 10-15 MINUTE DURATION. PROBLEM FOUND IS OF SMPS AND DRIVE CONTROLLER BOARD. FIXED THE NEW SMPS AND DRIVE CONTROLLER BOARD. AUTOMATICALLY SHUTS DOWN PROBLEM IS RESOLVED. CHECKED THE VENTILATOR AT DIFFERENT MODES. INTENSIVE CARE VENTILATOR IS NOW WORKING. | Closed | ASJC0120 | HIGH | 441:58:0 | 19
227 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00076 | SUCTION PUMP PORTABLE, ELECTRONIC | MANJUSHA NEVARE | 9423646267 | 14957 | 16-Apr-2018 | 09:15 |  |  | 19-Apr-2018 | 10:15 | NOT WORKING | RECOMMEND FOR CONDEMNATION PROCESS. | Closed | NO INFO | MEDIUM | 73:0:0 | 4
228 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00076 | SUCTION PUMP PORTABLE, ELECTRONIC | MAJUSHA NIWARE | 9423646267 | 6397 | 23-Aug-2017 | 08:55 | 24-Aug-2017 | 15:30 | 25-Aug-2017 | 14:30 | NOT WORKING | CLEANED ROTATOR SHAFT COUPLING AND POURED OIL IN IT ROTAR SECTION,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 53:35:0 | 3
229 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00076 | SUCTION PUMP PORTABLE, ELECTRONIC | DOLLY SCOTT | 9764971169 | 40819 | 12-Mar-2020 | 09:40 |  |  | 12-Mar-2020 | 15:58 | NOT WORKING | CHECK AND FOUND POWER PLUCK IS FAULTY AND POWER CABLE IS DEFECTIVE. FIXED THE NEW POWER PLUCK AND POWER CABLE IS REPAIRED. HYDRAULIC OIL IS BEING FILLED AS SUCTION PRESSURE IS LOW. POWER SUPPLY FAILURE ISSUE IS RESOLVED AND DESIRED SUCTION PRESSURE IS GENERATED. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 6:18:0 | 1
230 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00076 | SUCTION PUMP PORTABLE, ELECTRONIC | MANJUSHA NURSE | 9423646267 | 13152 | 23-Feb-2018 | 09:05 |  |  | 26-Feb-2018 | 12:36 | NOT WORKING | REPLACED NEW BRASS IRON INTO AN COMPRESSOR SECTION,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 75:31:0 | 4
231 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00076 | SUCTION PUMP PORTABLE, ELECTRONIC | MANJUSHA NEWARE | 9423646267 | 5955 | 08-Aug-2017 | 09:05 |  |  | 09-Aug-2017 | 14:14 | NOT WORKING | ADJUST THE VACUUM CONTROL AND RECONNECTED THE SUCTION ASSEMBLY . NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 29:9:0 | 2
232 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00077 | SYRINGE PUMP | MAJUSHA NIWARE | 9423646267 | 6395 | 23-Aug-2017 | 08:53 | 25-Aug-2017 | 11:47 | 01-Sep-2017 | 09:51 | NOT WORKING | REPLACE THE SYRIANGE PLUNGER AND HOLDER | Closed | NO INFO | MEDIUM | 216:58:0 | 10
233 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00078 | MONITOR | SCOTT NURSE | 9764971169 | 33120 | 26-Sep-2019 | 09:30 |  |  | 26-Sep-2019 | 14:21 | NOT WORKING | CHECK AND FOUND BLOOD PRESSURE COULD NOT BE MEASURED AS THERE IS LEAKAGE IN NIBP CUFF. CHANGED THE OLD DEFECTIVE CUFF WITH THE NEW ONE. B. P IS NOW ACCURATELY MEASURED. MONITOR IS NOW WORKING. | Closed | 068E00 4549 | MEDIUM | 4:51:0 | 1
234 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00078 | MONITOR | SCOT NURSE | 9764971169 | 34521 | 01-Nov-2019 | 08:15 |  |  | 01-Nov-2019 | 16:16 | NOT WORKING | CHECK AND FOUND MACHINE DOESN'T TURNS ON AS THERE IS POWER SUPPLY FAILURE ISSUE. FIXED THE POWER SUPPLY ISSUE. FUSE HAS BEEN REPLACED. SMPS BOARD HAS BEEN CHECKED. PROBLEM IS RESOLVED. MONITOR IS NOW WORKING. | Closed | 068E00 4549 | MEDIUM | 8:1:0 | 1
235 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00078 | MONITOR | LAXMI APURKAR | 9423417228 | 19423 | 21-Aug-2018 | 09:00 |  |  | 06-Nov-2018 | 11:08 | NOT WORKING | CHECK AND FOUND MONITOR SPO2 PROBE IS DEFECTIVE. CHANGE THE NEW SPO2 PROB NOW MACHINE IS WORKING | Closed | 068E00 4549 | MEDIUM | 1850:8:0 | 78
236 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00078 | MONITOR | UJWALA POKALE | 8530322606 | 5204 | 11-Jul-2017 | 09:05 | 14-Jul-2017 | 11:00 | 17-Jul-2017 | 10:52 | NOT WORKING | REPLACE SPO2 CABLE | Closed | 068E00 4549 | MEDIUM | 145:47:0 | 7
237 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00078 | MONITOR | SCOTT NURSE | 9764971169 | 33645 | 10-Oct-2019 | 09:05 |  |  | 12-Oct-2019 | 14:21 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS FAULTY AND THERE IS LEAKAGE IN NIBP CUFF. CHANGED THE DEFECTIVE PARTS WITH THE NEW ONE. NOW, SPO2 IS BEING MONITORED AND ALSO THE BLOOD PRESSURE. MONITOR IS NOW WORKING. | Closed | 068E00 4549 | MEDIUM | 53:16:0 | 3
238 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00078 | MONITOR | PRADIP WASNIK | 9423605761 | 1566 | 09-Feb-2017 | 09:58 |  |  | 10-Feb-2017 | 14:26 | NOT WORKING | WAITING FOR SPARES PARTS | Closed | 068E00 4549 | MEDIUM | 28:28:0 | 2
239 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00079 | MONITOR | PRAVIN MADGURE | 9021293547 | 14691 | 07-Apr-2018 | 09:21 |  |  | 07-Apr-2018 | 15:52 | NOT WORKING | REPLACED NEW 3AMPS FUSE, NOW MACHINE IS WORKING CONDITION. | Closed | 068E00 4548 | MEDIUM | 6:31:0 | 1
240 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00079 | MONITOR | UJWALA POKALE | 8530322606 | 5203 | 11-Jul-2017 | 09:05 |  |  | 17-Jul-2017 | 10:48 | NOT WORKING | REPLACE SPO2 CABLE | Closed | 068E00 4548 | MEDIUM | 145:43:0 | 7
241 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00079 | MONITOR | PRADIP WASNIK | 9423605761 | 1567 | 09-Feb-2017 | 10:00 |  |  | 10-Feb-2017 | 14:33 | NOT WORKING | WAITING FOR SPARES PARTS | Closed | 068E00 4548 | MEDIUM | 28:33:0 | 2
242 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00079 | MONITOR | SCOTT NURSE | 9764971169 | 30498 | 27-Jul-2019 | 09:45 | 27-Jul-2019 | 13:33 | 29-Jul-2019 | 16:55 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS DEFECTIVE. CHANGED THE OLD DEFECTIVE PROBE WITH THE NEW ONE. MULTIPARAMONITOR IS NOW WORKING. | Closed | 068E00 4548 | MEDIUM | 55:10:0 | 3
243 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00079 | MONITOR | UJAWALA GHOKHALE | 8530322606 | 11149 | 30-Dec-2017 | 09:15 |  |  | 31-Dec-2017 | 18:37 | NOT WORKING | REPLACED THE NIBP CUFF | Closed | 068E00 4548 | MEDIUM | 33:22:0 | 2
244 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00079 | MONITOR | LAXMI APURKAR | 9423417228 | 19420 | 21-Aug-2018 | 09:00 |  |  | 06-Nov-2018 | 11:07 | NOT WORKING | CHECK AND FOUND SPO2 PROB IS DEFECTIVE .CHANGE THE SAME NOW MONITOR IS WORKING | Closed | 068E00 4548 | MEDIUM | 1850:7:0 | 78
245 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00079 | MONITOR | UJWALA POKALE | 8530322606 | 7902 | 07-Oct-2017 | 09:24 |  |  | 08-Oct-2017 | 20:31 | NOT WORKING | REPLACED THE ADULT SPO2 CABLE | Closed | 068E00 4548 | MEDIUM | 35:7:0 | 2
246 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00079 | MONITOR | SCOTT NURSE | 9764971169 | 36646 | 17-Dec-2019 | 08:26 |  |  | 18-Dec-2019 | 14:30 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE ISSUE. FIXED THE DEFECTIVE POWER CORD WITH THE NEW ONE. FUSE IS BEING FIXED. SMPS IS BEING CROSS-CHECKED. POWER SUPPLY FAILURE ISSUE RESOLVED. MACHINE IS WORKING. | Closed | 068E00 4548 | MEDIUM | 30:4:0 | 2
247 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00080 | SYRINGE PUMP | UJWALA POKLEY | 8530322606 | 5186 | 10-Jul-2017 | 10:05 | 13-Jul-2017 | 10:57 | 01-Aug-2017 | 09:16 | NOT WORKING | RESET THE SETTING AND GUIDED THE TECHNICIANS | Closed | 090216B4.5RY02Z | MEDIUM | 527:11:0 | 23
248 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00081 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | MANJUSHA NEVARE | 9423646267 | 14955 | 16-Apr-2018 | 09:15 |  |  | 18-Apr-2018 | 10:13 | NOT WORKING | REPLACED NEW O2 CONNECTOR AND ALSO CALIBRATION IS DONE ,MACHINE IS WORKING CONDITION. | Closed | ASED0101 | HIGH | 48:58:0 | 3
249 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00081 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | MANJUSHA NEVARE | 9423646267 | 12410 | 01-Feb-2018 | 09:25 |  |  | 05-Feb-2018 | 14:04 | NOT WORKING | RECTIFIED THE COMPRESSOR AUTOMATIC SHUT DOWN PROBLEM AND NOW CHECKED COMPRESS MOTOR IS RUNNING CONTINOUSLY.MACHINE IS WORKING CONDITION. | Closed | ASED0101 | HIGH | 100:39:0 | 5
250 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00081 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | UJWALA POKALE | 8530322606 | 4901 | 29-Jun-2017 | 09:45 | 30-Jun-2017 | 22:03 | 20-Sep-2017 | 15:00 | NOT WORKING | MACHINE HANDOVER IN GOOD WORKING CONDIITON(WORK ORDER WAS ALREADY COMPLTED ON 5TH AUGUST) | Closed | ASED0101 | HIGH | 1997:15:0 | 84
251 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00081 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | Mrs D. V Scott | 9764971169 | 42679 | 28-May-2020 | 09:00 |  |  | 28-May-2020 | 10:45 | NOT WORKING | CHECK AND FOUND CALIBRATION OF OXYGEN FAILED AS OXYGEN SENSOR IS FAULTY. FIXED THE NEW OXYGEN SENSOR AND CALIBRATION OF OXYGEN IS SUCCESSFULLY COMPLETED. NOW, INFOW OF OXYGEN IS MEASURED AND SAME IS DISPLAYED IN FIO2 COLUMN IN VARIOUS MODES OF VENTILATION. VENTILATOR IS NOW WORKING. | Closed | ASED0101 | HIGH | 1:45:0 | 1
252 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00082 | SUCTION PUMP PORTABLE, ELECTRONIC | MANJUSHA NURSE | 9423646267 | 13153 | 23-Feb-2018 | 09:05 |  |  | 27-Feb-2018 | 14:21 | NOT WORKING | CHECKED AND FOUND VACUUM PRESSURE IS NOT ENOUGH,SO WE REPLACED NEW BRASS INTO AN MOTOR BECAUSE OF RUSTED SHAFT AFTER REPLACING IRON IT WON'T WORK PROPERLY.SO WE RECOMMENDED FOR CONDEMNATION. | Closed | RM1000165 | MEDIUM | 101:16:0 | 5
253 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00082 | SUCTION PUMP PORTABLE, ELECTRONIC | DOLLY SCOTT | 9764971169 | 40820 | 12-Mar-2020 | 09:40 |  |  | 12-Mar-2020 | 17:09 | NOT WORKING | CHECK AND FOUND POWER ON/OFF SWITCH IS FAULTY. FIXED THE NEW POWER 'ON/OFF' SWITCH. HYDRAULIC OIL IS BEING FILLED AS SUCTION PRESSURE IS LOW. POWER SUPPLY FAILURE ISSUE RESOLVED AND DESIRED SUCTION PRESSURE IS GENERATED. SUCTION MACHINE IS NOW WORKING. | Closed | RM1000165 | MEDIUM | 7:29:0 | 1
254 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00082 | SUCTION PUMP PORTABLE, ELECTRONIC | MANJUSHA NEVARE | 9423646267 | 14958 | 16-Apr-2018 | 09:15 |  |  | 19-Apr-2018 | 10:16 | NOT WORKING | RECOMMENDED FOR CONDEMNATION PROCESS | Closed | RM1000165 | MEDIUM | 73:1:0 | 4
255 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00083 | MONITOR | UJAWALA GHOKHALE | 8530322606 | 11150 | 30-Dec-2017 | 09:15 |  |  | 31-Dec-2017 | 18:33 | NOT WORKING | REPLACED THE NIBP CUFF | Closed | 068E00 4567 | MEDIUM | 33:18:0 | 2
256 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00083 | MONITOR | SCOTT NURSE | 9764971169 | 37858 | 13-Jan-2020 | 08:05 |  |  | 14-Jan-2020 | 12:45 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS DEFECTIVE AND NIBP CUFF IS HAVING LEAKAGE. FIXED THE SPO2 PROBE AND FIXED THE NEW NIBP CUFF. MONITOR IS NOW WORKING. | Closed | 068E00 4567 | MEDIUM | 28:40:0 | 2
257 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00083 | MONITOR | UJWALA POKALE | 8530322606 | 5202 | 11-Jul-2017 | 09:05 | 14-Jul-2017 | 10:59 | 17-Jul-2017 | 18:42 | NOT WORKING | REPLACE SPO2 CABLE | Closed | 068E00 4567 | MEDIUM | 153:37:0 | 7
258 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00083 | MONITOR | LAXMI APURKAR | 9423417228 | 19421 | 21-Aug-2018 | 09:00 |  |  | 06-Nov-2018 | 11:12 | NOT WORKING | CHECK AND FOUND SPO2 PROB IS DEFECTIVE. CHANGE THE SAME SPO2 NOW MACHINE IS WORKING | Closed | 068E00 4567 | MEDIUM | 1850:12:0 | 78
259 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00083 | MONITOR | DOLLY SCOTT | 9764971169 | 35456 | 20-Nov-2019 | 08:57 |  |  | 20-Nov-2019 | 14:36 | NOT WORKING | CHECK AND FOUND NIBP HOSE PIPE IS BROKEN. CHANGED THE OLD DEFECTIVE HOSE PIPE WITH THE NEW ONE. NOW B.P IS MONITORED WITH NO ERROR. MONITOR IS NOW WORKING. | Closed | 068E00 4567 | MEDIUM | 5:39:0 | 1
260 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00083 | MONITOR | PRADIP WASNIK | 9423605761 | 1568 | 09-Feb-2017 | 10:01 |  |  | 10-Feb-2017 | 14:30 | NOT WORKING | WAITING FOR SPARES PARTS | Closed | 068E00 4567 | MEDIUM | 28:29:0 | 2
261 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00083 | MONITOR | SCOTT NURSE | 9764971169 | 32976 | 24-Sep-2019 | 09:12 |  |  | 25-Sep-2019 | 15:13 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS HAVING LEAKAGE AND DUE TO WHICH BLOOD PRESSURE COULD NOT BE MEASURED. CHANGED THE OLD DEFECTIVE CUFF WITH THE NEW ONE. B. P IS NOW ACCURATELY MEASURED. MONITOR IS NOW WORKING. | Closed | 068E00 4567 | MEDIUM | 30:1:0 | 2
262 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00085 | SYRINGE PUMP | SCOTT NURSE | 9764971169 | 32399 | 13-Sep-2019 | 08:05 |  |  | 14-Sep-2019 | 12:19 | NOT WORKING | CHECK AND FOUND POWER CORD IS DEFECTIVE. CHANGED THE DEFECTIVE POWER CORD WITH THE NEW ONE. SYRINGE PUMP IS NOW WORKING. | Closed | 1202672 | MEDIUM | 28:14:0 | 2
263 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00085 | SYRINGE PUMP | SCOTT NURSE | 9764971169 | 41512 | 07-Apr-2020 | 08:10 |  |  | 09-Apr-2020 | 16:10 | NOT WORKING | CHECK AND FOUND DISPLAY SCREEN IS BROKEN DUE TO MISHANDLING. EQUIPMENT IS APPROXIMATELY 10 YEARS OLD AND IS EXCEEDING RECOMMENDED LIFE SPAN. IT'S SPARES IS NOT AVAILABLE. BODY IS ALSO DAMAGED. SO, SYRINGE PUMP IS PROPOSED FOR CONDEMNATION. | Closed | 1202672 | MEDIUM | 56:0:0 | 3
264 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00085 | SYRINGE PUMP | SCOTT NURSE | 9764971169 | 36645 | 17-Dec-2019 | 08:26 |  |  | 18-Dec-2019 | 10:53 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN DISPLAY IS DAMAGE. RAISED MR FOR DISPLAY WHEN IT COME THEN WE WILL CHANGE THE NEW DISPLAY. MACHINE IS NOT WORKING CONDITION. | Closed | 1202672 | MEDIUM | 26:27:0 | 2
265 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00087 | Mobile X-ray machine (HF) | WALMIK MEHER | 8275298328 | 7480 | 25-Sep-2017 | 08:22 |  |  | 26-Sep-2017 | 11:20 | NOT WORKING | RECTIFIED THE ERROR | Closed | 31097 | MEDIUM | 26:58:0 | 2
266 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00087 | Mobile X-ray machine (HF) | DONWADE | 9403360423 | 18537 | 27-Jul-2018 | 09:30 |  |  | 03-Aug-2018 | 17:30 | NOT WORKING | CHECK AND FOUND THAT IGBT BOARD FAULTY. REPLACED THE IGBT BOARD AS SAME NOW MACHINE IS WORKING SATISFACTORY. HANDOVER TO USER. | Closed | 31097 | MEDIUM | 176:0:0 | 8
267 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00087 | Mobile X-ray machine (HF) | MEHAR | 9834283591 | 39665 | 18-Feb-2020 | 09:06 |  |  | 18-Feb-2020 | 11:30 | NOT WORKING | CHECKED AND FOUND MAINS WIRE WAS FAULTY. REPLACED WITH THE NEW ONE. MACHINE TESTED AND HAND-OVER TO USER IN PROPER WORKING ORDER. | Closed | 31097 | MEDIUM | 2:24:0 | 1
268 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00087 | Mobile X-ray machine (HF) | Ravindra | 8055273031 | 42517 | 22-May-2020 | 08:32 |  |  | 26-May-2020 | 11:14 | NOT WORKING | CHECK ANF FOUND EXPOSURE HAND SWITCH IS FAULTY. REPAIRED THE HAND SWITCH. NOW, MACHINE WORKING OK. NOTE: THERE IS LARGE FLUCTUATION IN VOLTAGE, SO DIRECTING THE STAFF TO INSTALL SINGLE PHASE 3KV STABLIZER. | Closed | 31097 | MEDIUM | 98:42:0 | 5
269 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00087 | Mobile X-ray machine (HF) | MANJUSHA NEVARE | 9423646267 | 15672 | 09-May-2018 | 08:45 |  |  | 12-May-2018 | 10:43 | NOT WORKING | CHECK AND FOUND SHOOTING WAS NOT PROPER AND THE MOVING ARM IS LOOSE. RECTIFIED BOTH THE PROBLEM. THE MACHINE IS IN WORKING CONDITION | Closed | 31097 | MEDIUM | 73:58:0 | 4
270 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00087 | Mobile X-ray machine (HF) | RAVINDRA FATING | 8055273031 | 19657 | 27-Aug-2018 | 09:19 |  |  | 30-Aug-2018 | 12:32 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE.MOV IS BURNT OUT.CHANGED THE NEW MOV.POWET SUPPLY FAILURE IS BEING RECTIFIED.X-RAY MACHINE IS WORKING NOW. | Closed | 31097 | MEDIUM | 75:13:0 | 4
271 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00087 | Mobile X-ray machine (HF) | MANJUSHA NEVARE | 9423646267 | 18092 | 18-Jul-2018 | 10:20 |  |  | 19-Jul-2018 | 13:30 | NOT WORKING | MACHINE CHECK AND FOUND ERROR 33 AND ERROR 16 PROBLEM IN MASTER BOARD AND INVERTOR BOARD. RECTIFIED THE SAME. NOW MACHINE IS IN WORKING PROPERLY. | Closed | 31097 | MEDIUM | 27:10:0 | 2
272 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00087 | Mobile X-ray machine (HF) | Ravindra | 8055273031 | 41628 | 13-Apr-2020 | 09:35 |  |  | 16-Apr-2020 | 15:43 | NOT WORKING | D 195 MOTHER BOARD FAULTY GIVING ALARM OF E14 WHILE SHOOTING. FIXED THE NEW D 195 MOTHER BOARD. E14 ERROR PROBLEM IS RECTIFIED. X-RAY IS BEING TAKEN AND FILM OBTAINED IS SATISFACTORY. X-RAY MACHINE IS NOW WORKING. | Closed | 31097 | MEDIUM | 78:8:0 | 4
273 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00087 | Mobile X-ray machine (HF) | SCOT NURSE | 9764971169 | 25365 | 25-Feb-2019 | 09:11 |  |  | 28-Feb-2019 | 14:22 | NOT WORKING | CHECK AND FOUND DISPLAY BLINKING AND ERROR 13 PROBLEM. INTERNAL WIRES WAS BROKEN DUE TO RAT BITE.DONE THE PROPER WIRING AND MAINS BOARD D9 15 REPAIRED.MACHINE TESTED AND HANDOVER TO TECHNICIAN IN PROPER WORKING ORDER. | Closed | 31097 | MEDIUM | 77:11:0 | 4
274 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00088 | SYRINGE PUMP | UJWALA POKALE | 8530322606 | 5841 | 03-Aug-2017 | 09:15 |  |  | 05-Aug-2017 | 20:07 | NOT WORKING | RESET THE SETTING AND GUIDED THE TECHNICIANS | Closed | ST06 | MEDIUM | 58:52:0 | 3
275 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00088 | SYRINGE PUMP | DOIIY SCOTT | 9464971169 | 34105 | 19-Oct-2019 | 08:45 |  |  | 20-Oct-2019 | 14:53 | NOT WORKING | CHECK AND FOUND PLUNGER IS BROKEN. FIXED THE NEW PLUNGER.CHECKED THE MACHINE AT DIFFERENT RATE OF FLOW BY INSERTING SYRINGE. MACHINE IS WORKING. | Closed | ST06 | MEDIUM | 30:8:0 | 2
276 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00089 | PULSE OXIMETER LINE POWERED | SCOTT NURSE | 9764971169 | 30419 | 25-Jul-2019 | 08:03 |  |  | 25-Jul-2019 | 17:39 | NOT WORKING | CHECK AND FOUND POWER SUPPLY ISSUES IS THERE AS FUSE IS BURNT. CHANGE THE NEW FUSE OF RATING 5A. POWER SUPPLY FAILURE ISSUES IS BEING SORTED OUT. PULSE OXIMETRE IS NOW WORKING. | Closed | 187816 | MEDIUM | 9:36:0 | 1
277 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00089 | PULSE OXIMETER LINE POWERED | D. V Scott | 9764971169 | 43870 | 06-Jul-2020 | 08:22 |  |  | 06-Jul-2020 | 09:45 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS DEFECTIVE. FIXED THE NEW SPO2 PROBE AND PULSE RATE & SPO2 IS MEASURED. READING OBTAINED IS ACCEPTABLE. PULSE OXIMETER LINE POWERED IS NOW WORKING. | Closed | 187816 | MEDIUM | 1:23:0 | 1
278 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00092 | BLOOD GAS ANALYSER (HANDHELD) | MANJUSHA NEVARE | 9423646267 | 12411 | 01-Feb-2018 | 09:25 |  |  | 05-Feb-2018 | 14:15 | NOT WORKING | CONSUMABLE PART ELECTRODES ARE COMPLETED, SO WE NEED TO CHANGE ELECTRODES AND RECOMMENDED TO TECHNICIAN ALSO | Closed | NO INFO | MEDIUM | 100:50:0 | 5
279 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00092 | BLOOD GAS ANALYSER (HANDHELD) | MANJUSHA NEWARE | 9423646267 | 5998 | 10-Aug-2017 | 09:25 |  |  | 14-Aug-2017 | 11:12 | NOT WORKING | REARRANGED POWER CORD SECTION PROPERLY,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 97:47:0 | 5
280 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00092 | BLOOD GAS ANALYSER (HANDHELD) | MAJUSHA NEVARE | 9423646267 | 6829 | 08-Sep-2017 | 08:59 |  |  | 09-Sep-2017 | 13:30 | NOT WORKING | MACHINE IS WORKING CONDITION BUT SOLUTION GET EXPERIED,SO WE REQUEST TO PURCHASE THE SOLUTION FOR THIS MACHINE. | Closed | NO INFO | MEDIUM | 28:31:0 | 2
281 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00092 | BLOOD GAS ANALYSER (HANDHELD) | MANJUSHA NEVARE | 9423646267 | 13402 | 03-Mar-2018 | 10:28 |  |  | 05-Mar-2018 | 14:24 | NOT WORKING | CHECKED AND FOUND ELECTRODES HAS AN PROBLEM, BUT IT IS CONSUMABLE PART SO WE INSTRUCT USER TO BUY NEW ELECTRODES | Closed | NO INFO | MEDIUM | 51:56:0 | 3
282 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00092 | BLOOD GAS ANALYSER (HANDHELD) | PRATIK | 9423605761 | 634 | 31-Dec-2016 | 11:32 |  |  | 05-Jan-2017 | 16:15 | NOT WORKING | CONNECTOR RECTIFIED | Closed | NO INFO | MEDIUM | 124:43:0 | 6
283 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00093 | SYRINGE PUMP | SENDRA NURSE | 9373476576 | 27711 | 17-May-2019 | 09:00 |  |  | 22-May-2019 | 15:15 | NOT WORKING | CHECK AND FOUND PLUNGER IS BROKEN. FIXED THE NEW PLUNGER. SYRINGE PUMP IS NOW WORKING. | Closed | 150216B4.5AJ0101Z | MEDIUM | 126:15:0 | 6
284 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00093 | SYRINGE PUMP | SCOTT NURSE | 9764971169 | 32398 | 13-Sep-2019 | 08:05 |  |  | 14-Sep-2019 | 16:41 | NOT WORKING | CHECK AND FOUND PLUNGER IS BROKEN AND DUE TO WHICH SYRINGE COULD NOT BE FIXED. FIXED THE NEW PLUNGER. SYRINGE PUMP IS NOW WORKING. | Closed | 150216B4.5AJ0101Z | MEDIUM | 32:36:0 | 2
285 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00093 | SYRINGE PUMP | MAJUSHA NIWARE | 9423646267 | 6396 | 23-Aug-2017 | 08:55 | 25-Aug-2017 | 14:47 | 01-Sep-2017 | 12:54 | NOT WORKING | REPLACED THE SYRIANGE HOLDER AND PLUNGER | Closed | 150216B4.5AJ0101Z | MEDIUM | 219:59:0 | 10
286 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00093 | SYRINGE PUMP | SCOTT NURSE | 9764971169 | 41514 | 07-Apr-2020 | 08:10 |  |  | 09-Apr-2020 | 11:35 | NOT WORKING | CHECK AND FOUND OCCLUSION ALARM OCCURS JUST AFTER STARTING THE DOSE AND THE PUMP STOPS. ISSUE WITH PLUNGER SET AND MOTOR. MR HAD BEEN RAISED FOR NEW PLUNGER SET AND MOTOR. AS SOON AS PARTS WILL BE RECEIVED IT WILL BE FIXED AND MACHINE WILL BE MADE FUNCTIONAL. | Closed | 150216B4.5AJ0101Z | MEDIUM | 51:25:0 | 3
287 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00093 | SYRINGE PUMP | AVIKA YADAV |  | 14396 | 01-Apr-2018 | 08:14 |  |  | 04-Apr-2018 | 11:30 | NOT WORKING | REPLACED NEW SYRINGE PLUNGER HOLDER INTO AN MACHINE , NOW MACHINE IS WORKING CONDITION . | Closed | 150216B4.5AJ0101Z | MEDIUM | 75:16:0 | 4
288 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00094 | PULSE OXIMETER LINE POWERED | Mrs D. V Scott | 9764971169 | 42681 | 28-May-2020 | 09:00 |  |  | 28-May-2020 | 14:45 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS FAULTY. FIXED THE NEW SPO2 PROBE. OXYGEN SATURATION AND PULSE IS BEING MONITORED AND FOUND TO BE ACCEPTABLE. READING IS CROSS-CHECKED WITH OTHER PULSE OXIMETER AND FOUND TO BE SATISFIABLE. PULSE OXIMETER IS NOW WORKING. | Closed | 181711 | MEDIUM | 5:45:0 | 1
289 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00097 | AUTOCLAVE HP HORIZONTAL | HARSHA BHUSARI | 9096022114 | 39137 | 06-Feb-2020 | 09:31 |  |  | 06-Feb-2020 | 17:17 | NOT WORKING | CHECK AND FOUND DOOR ISSUE.JACK HANDLE AND JACK PLATE IS SMOOTHENED ON LATHE MACHINE. NEW BALL BEARING IS FIXED. SAFETY VALVES ARE CLEANED. HIGH PRESSURE EXHAUST STEAM PIPE LINE IS CHECKED. COMPLETE AUTOCLAVE PROCESS IS BEING DONE 3-4 TIMES. USER TRAINING IS BEING GIVEN. AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 7:46:0 | 1
290 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00097 | AUTOCLAVE HP HORIZONTAL | BAVANKAR NURSE | 8020873321 | 25720 | 11-Mar-2019 | 08:24 |  |  | 18-Mar-2019 | 21:46 | NOT WORKING | CHECK AND FOUND DOOR WAS COMPLETELY JAMMED AND ITS NUTS AND SCREWS ARE BROKEN. DUE TO USER MISHANDLING THE DOOR WAS JAMMED AND ITS ACCESSORIES ARE DAMAGED AND BROKEN. IT IS PROPOSED FOR RW AND ITS MAINTENANCE COST WILL BE PAID BY HOSPITAL AUTHORITY IF THE PARTS ARE TO BE BOUGHT. AT PRESENT CALL IS BEING CLOSED NOW AND IF PARTS WOULD BE AVAILABLE THEN THE CALL WILL BE RAISED AGAIN AND MACHINE WILL BE MADE FUNCTIONAL. | Closed | NO INFO | MEDIUM | 181:22:0 | 8
291 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00097 | AUTOCLAVE HP HORIZONTAL | SHILA BAVANKAR | 8208723521 | 26818 | 15-Apr-2019 | 08:50 |  |  | 24-Apr-2019 | 09:58 | DOOR ISSUE | CHECK AND FOUND MACHINE PROPOSE FOR CONDEMNATION DUE TO USER MISHANDLING . | Closed | NO INFO | MEDIUM | 217:8:0 | 10
292 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00097 | AUTOCLAVE HP HORIZONTAL | BAVANKAR NURSE | 8020873321 | 27663 | 16-May-2019 | 08:25 |  |  | 20-May-2019 | 16:03 | NOT WORKING | CHECK AND FOUND DOOR OF THE MACHINE IS FAULTY AND NEW BALL BEARING SET IS REQUIRED. MR IS BEING RAISED AND PART IS TO BE REVEIVED YET. AS SOON AS NEW BALL BEARING SET WILL BE RECEIVED,AGAIN COMPLAINT WILL BE RAISED AND MACHINE WILL BE MADE OK. | Closed | NO INFO | MEDIUM | 103:38:0 | 5
293 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00097 | AUTOCLAVE HP HORIZONTAL | BAVANKAR SHEELA | 9422132247 | 7676 | 29-Sep-2017 | 09:29 |  |  | 02-Oct-2017 | 13:22 | NOT WORKING | REPLACED NEW 6K WATTS STRAIGHT HEATING COIL,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 75:53:0 | 4
294 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00097 | AUTOCLAVE HP HORIZONTAL | BAVANKAR | 9422132247 | 16769 | 15-Jun-2018 | 09:40 |  |  | 17-Jun-2018 | 12:30 | NOT WORKING | CHECK AND FOUND LEAKAGE FROM HEATER COIL SIDE. FIX THE LEAKAGE, NOW AUTOCLAVE WORKING SATISFACTORY. | Closed | NO INFO | MEDIUM | 50:50:0 | 3
295 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00097 | AUTOCLAVE HP HORIZONTAL | LAVANKAR NURSE | 9522132247 | 18991 | 10-Aug-2018 | 09:13 |  |  | 16-Aug-2018 | 16:00 | NOT WORKING | CHECK AND FOUND WATER LEAKAGE FROM GLASS TUBE. FIX THE GLASS TUBE PROPERLY. WATER LEAKAGE PROBLEM SOLVED. NOW AUTOCLAVE WORKING SATISFACTORY. | Closed | NO INFO | MEDIUM | 150:47:0 | 7
296 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00098 | AUTOCLAVE HP HORIZONTAL | RAJKUKAR KANOJIA | 9309176931 | 27071 | 24-Apr-2019 | 09:06 |  |  | 25-Apr-2019 | 16:52 | NOT WORKING | CHECK AND FOUND HEATER IS NOT TURNED ON. PROBLEM IS FOUND OF CONTRACTOR SWITCH. REPAIRED THE CONTRACTOR SWITCH. ALL THE VALVES IS BEING CHECKED. AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 31:46:0 | 2
297 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00098 | AUTOCLAVE HP HORIZONTAL | HARSHA BHUSARI | 9096022114 | 36253 | 06-Dec-2019 | 08:50 |  |  | 06-Dec-2019 | 10:17 | NOT WORKING | CHECK AND FOUND MACHINE DISPLAY IS NOT SHOWING DIGITS SO REPAIRED THE DISPLAY BOARD AND FIT ON OPRATING PANNEL AND DIGITS SHOWING ON DISPLAY AND MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 1:27:0 | 1
298 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00099 | AUTOCLAVE HP HORIZONTAL | POTPODE | 9422474142 | 23005 | 14-Dec-2018 | 09:00 |  |  | 18-Dec-2018 | 11:05 | NOT WORKING | CHECK AND FOUND CONTACTOR SWITCH FAULTY AND DUE TO WHICH VIBRATIONAL SOUND IS BEING HEARD. CHANGED THE NEW CONTRACTOR SWITCH. MACHINE IS BEING CHECK HORIZONTAL AUTOCLAVE IS NOW WORKING | Closed | NO INFO | MEDIUM | 98:5:0 | 5
299 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00099 | AUTOCLAVE HP HORIZONTAL | BAVANKAR SHEELA | 9422132247 | 7678 | 29-Sep-2017 | 09:33 |  |  | 03-Oct-2017 | 14:25 | NOT WORKING | REPLACED NEW OUTLET PIPE HALF INCH ON/OFF VALVE AND DONE CLEANING PROCESS ON HEATING COIL SECTION. NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 100:52:0 | 5
300 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00100 | AUTOCLAVE HP HORIZONTAL | BAVANKAR SHEELA | 9422132247 | 7677 | 29-Sep-2017 | 09:31 |  |  | 03-Oct-2017 | 14:10 | NOT WORKING | DONE CLEANING PROCESS ON HEATING COIL SECTION AND RECTIFY PRESSURE LEAKAGE PROBLEM,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 100:39:0 | 5
301 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00100 | AUTOCLAVE HP HORIZONTAL | SHILA BAVANKAR | 8208723521 | 26817 | 15-Apr-2019 | 08:50 |  |  | 18-Apr-2019 | 17:35 | LEAKAGE ISSUE | CHECK AND FOUND HEATING COIL IS BURNT AND DUE TO WHICH WATER COULD NOT BE BOILED. CHANGED ALL THE DEFECTIVE COIL WITH THE NEW ONE OF RATING 9KW.AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 80:45:0 | 4
302 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00100 | AUTOCLAVE HP HORIZONTAL | BAVANKAR NURSE | 9422132247 | 19376 | 20-Aug-2018 | 09:55 |  |  | 22-Aug-2018 | 18:30 | NOT WORKING | CHECK AND FOUND HEATER COIL IS FAULTY , REPLACED HEATER COIL , NOW AUTOCLAVE WORKING SATISFACTORY, WITHOUT ANY INTRUPT. | Closed | NO INFO | MEDIUM | 56:35:0 | 3
303 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00100 | AUTOCLAVE HP HORIZONTAL | BAVANKAR NURSE | 8208733531 | 39956 | 24-Feb-2020 | 09:11 |  |  | 26-Feb-2020 | 15:08 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BLOCKED AND DUE TO WHICH WATER LEVEL COULD NOT BE DETECTED. CLEANED THE GLASS TUBE AND ITS VALVES. NOW, WATER LEVEL IS CLEARLY DETECTED. AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 53:57:0 | 3
304 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00100 | AUTOCLAVE HP HORIZONTAL | BAVANKAR NURSE | 9422132247 | 17071 | 23-Jun-2018 | 09:08 |  |  | 27-Jun-2018 | 16:50 | NOT WORKING | HECK AND FOUND SOME LEAKAGE FROM DOOR SIDE. FIX THE LEAKAGE NOW AUTOCLAVE WORKING SATISFACTORY. | Closed | NO INFO | MEDIUM | 103:42:0 | 5
305 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00100 | AUTOCLAVE HP HORIZONTAL | SHILA BAWANKAR | 9422132247 | 25756 | 12-Mar-2019 | 08:17 |  |  | 14-Mar-2019 | 15:12 | LEAKAG ISSUE | CHEAK AND FOUND COIL CONNECTING WIRE IS OPEND AND DUE TO WHICH WATER COULD NOT BE BOILED. RE-FIXED THE WIRE ACROSS COIL AND CONNECTOR SWITCH IS CHECKED . AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 54:55:0 | 3
306 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00100 | AUTOCLAVE HP HORIZONTAL | SHEELA BAVANKAR | 94221322247 | 19858 | 31-Aug-2018 | 09:05 |  |  | 06-Sep-2018 | 12:00 | NOT WORKING | CHECK AND FOUND HORIZONTAL AUTOCLAVE STEAM VALVE IS NOT WORKING PROPERLY. CHANGED THE SAME ABOVE MENTION VALUE. NOW MACHINE IS WORKING GOOD. | Closed | NO INFO | MEDIUM | 146:55:0 | 7
307 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00102 | AUTOCLAVE HP VERTICAL (2 BIN) DIGITAL | BAWANKAR | 8208733521 | 27155 | 27-Apr-2019 | 11:20 |  |  | 01-May-2019 | 17:49 | NOT WORKING | CHECK AND FOUND COIL IS BURNT AND DUE TO WHICH WATER COULD NOT BE BOILED. CHANGED THE HEATING COIL OF 6KW.AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 102:29:0 | 5
308 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00103 | AUTOCLAVE HP VERTICAL (2 BIN) DIGITAL | HARSHA BHUSARI | 9096022114 | 36252 | 06-Dec-2019 | 08:50 |  |  | 06-Dec-2019 | 10:58 | NOT WORKING | CHECK AND FOUND SOME CHOKUP FROM DRAINIS IN THE GLASS TUBE AND NOT SHOWING WATER LEVEL, SO NEED AND CLEAN GLASS TUBE AND AND MACHINE IS WORKING PROPERLY AND SHOWING WATER LEVEL. | Closed | NO INFO | MEDIUM | 2:8:0 | 1
309 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00106 | ANESTHESIA MACHINE (BOYLES APPARATUS) | Sheela Bawankar | 8208733521 | 44284 | 16-Jul-2020 | 08:42 |  |  |  |  | NOT WORKING |  | Open | NO INFO | HIGH
310 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00106 | ANESTHESIA MACHINE (BOYLES APPARATUS) | DR TARACHAND | 9049155994 | 5217 | 11-Jul-2017 | 09:35 | 12-Jul-2017 | 10:02 | 15-Jul-2017 | 11:55 | NOT WORKING | ARRESTED THE OXYGEN LEAKAGE | Closed | NO INFO | HIGH | 98:20:0 | 5
311 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00106 | ANESTHESIA MACHINE (BOYLES APPARATUS) | SULABHA MESHRAM | 9158285535 | 31714 | 23-Aug-2019 | 09:09 |  |  | 24-Aug-2019 | 15:30 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE OF OXYGEN THROUGH PIPES. CHANGED THE DEFECTIVE PIPES WITH THE NEW ONE. LEAKAGE PROBLEM IS RESOLVED. ANESTHESIA MACHINE IS NOW WORKING. | Closed | NO INFO | HIGH | 30:21:0 | 2
312 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00107 | OPERATION TABLE HYDRAULIC (MAJOR) | SHEELA BAVANKAR | 8208733521 | 29937 | 15-Jul-2019 | 08:09 | 15-Jul-2019 | 11:03 | 19-Jul-2019 | 16:34 | NOT WORKING | CHECK AND FOUND HYDRAULIC UP AND DOWN MOVEMENT IS NOT WORKING AS IT MAKE RUBBING METTALIC SOUND WITH VERY SLOW MOVEMENT. REFILLED THE HYDRAULIC OIL IN OIL CHAMBER. NOW THERE IS SMOOTH AND FAST UP AND DOWN MOVEMENT. OPERATION TABLE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 104:25:0 | 5
313 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00107 | OPERATION TABLE HYDRAULIC (MAJOR) | HARSHA | 9096022114 | 36718 | 18-Dec-2019 | 08:15 |  |  | 18-Dec-2019 | 10:06 | NOT WORKING | CHECK AND FOUND HYDRAULIC UP AND DOWN MOVEMENT IS NOT WORKING AS IT MAKE RUBBING METTALIC SOUND WITH VERY SLOW MOVEMENT. REFILLED THE HYDRAULIC OIL IN OIL CHAMBER. NOW THERE IS SMOOTH AND FAST UP AND DOWN MOVEMENT. OPERATION TABLE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 1:51:0 | 1
314 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00107 | OPERATION TABLE HYDRAULIC (MAJOR) | HARSHA | 9096022114 | 40493 | 05-Mar-2020 | 08:26 |  |  | 09-Mar-2020 | 11:38 | NOT WORKING | THIS TABLE IS ALREADY PROPOSED FOR BER(BEYOND ECONOMIC REPAIRS) AS QUOTATION RECEIVED FROM STAAN BIOMEDICAL IS MORE THAN 75% OF THE COST OF A NEW OT TABLE. DUE TO THIS AS PER GUIDELINES, THIS TABLE NEEDS TO BE CATEGORISED AS BER AS OUR MANAGEMENT WILL NOT SANCTION FUNDS FOR ITS REPAIRS. | Closed | NO INFO | MEDIUM | 99:12:0 | 5
315 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00107 | OPERATION TABLE HYDRAULIC (MAJOR) | HARSHA BHUSARI | 9096022114 | 38038 | 17-Jan-2020 | 08:15 |  |  | 11-Feb-2020 | 15:08 | NOT WORKING | ENGINEER FROM STAAN COMPANY HAS VISITED FOR THIS BREAKDOWN AND FOUND THERE IS LEAKAGE OF OIL FROM UP AND DOWN COLUMN INNER CYLINDER AND ALSO HAND REMOTE AND PANEL REMOTE IS FAULTY AND HE SUGGESTED TO CHANGE THE ABOVE PARTS. RECEIVED QUOTATION FROM STAAN AND ON CALCULATION FOUND THAT REPAIR COST IS MORE THAN 75 PERCENT OF THE ALLOTTED MAINTENANCE COST. SO IT IS PROPOSED FOR BER(BEYOND ECONOMICAL REPAIR). | Closed | NO INFO | MEDIUM | 606:53:0 | 26
316 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00108 | OPERATION TABLE HYDRAULIC (MAJOR) | SAVANKAR | 9422132247 | 16604 | 12-Jun-2018 | 08:15 |  |  | 13-Jun-2018 | 12:30 | NOT WORKING | UP AND DOWN MOVEMENT NOT WORKING. FILLED OT TABLE OUL | Closed | NO INFO | MEDIUM | 28:15:0 | 2
317 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00109 | C ARM SYSTEM (HF) | DR JIBHKATE | 9766918077 | 6698 | 06-Sep-2017 | 11:00 |  |  | 12-Sep-2017 | 10:04 | NOT WORKING | RECONNECT THE ALL DETECTOR CONNECTORS | Closed | V12129132 | MEDIUM | 143:4:0 | 7
318 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00110 | TOURNIQUET MACHINE | HARSHA BHUSARI | 9096022114 | 37920 | 13-Jan-2020 | 09:07 |  |  | 13-Jan-2020 | 16:32 | NOT WORKING | CHECK AND FOUND ARM AND THIGH CUFF IS FAULTY AND WORN OUT. REPLACED THE OLD CUFF WITH THE NEW ONE OF MEASUREMENT 46 CM AND 64 CM. TOURNIQUET MACHINE IS NOW WORKING. | Closed | 1098 | LOW | 7:25:0 | 1
319 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00111 | MANMAN ORTHO DRILL | SMITA KULKARNI | 9637606658 | 14924 | 14-Apr-2018 | 09:03 |  |  | 14-Apr-2018 | 15:20 | NOT WORKING | DONE WELDING ON BROKEN HOLDER, MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 6:17:0 | 1
320 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00111 | MANMAN ORTHO DRILL | REHANA KHAN | 9158548938 | 3940 | 22-May-2017 | 10:25 |  |  | 26-May-2017 | 12:12 | NOT WOKING | REPLACE BEARING | Closed | NO INFO | LOW | 97:47:0 | 5
321 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00112 | SHADOWLESS LAMP STAND TYPE | SHEELA BAVANKAR | 8208733521 | 29938 | 15-Jul-2019 | 08:09 | 15-Jul-2019 | 15:57 | 17-Jul-2019 | 14:29 | NOT WORKING | CHECK AND FOUND THE CALL IS WRONGLY LOGGED. ACTUALLY USER WAS TRYING TO LOG BREAKDOWN COMPLAINTS OF TOURNIQUET MACHINE. AT PRESENT SHADOWLESS LAMP OF ABOVE MENTIONED ASSET NUMBER IS WORKING. | Closed | NO INFO | MEDIUM | 54:20:0 | 3
322 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00113 | SUCTION PUMP PORTABLE, ELECTRONIC | HARSHA BHUSARI | 9096022114 | 36254 | 06-Dec-2019 | 08:50 |  |  | 06-Dec-2019 | 10:01 | NOT WORKING | CHECK AND FOUND INTERNAL TUBING IS OPENED DUE TO WHICH PRESSURE IS NOT GENERATED.FIXED THE INTERNAL TUBING.SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 1:11:0 | 1
323 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00113 | SUCTION PUMP PORTABLE, ELECTRONIC | BAVANKAR | 9422132247 | 20657 | 25-Sep-2018 | 08:05 |  |  | 25-Sep-2018 | 15:00 | NOT WORKING | CHECK AND FOUND INTERNAL TUBING IS OPENED DUE TO WHICH PRESSURE IS NOT GENERATED.FIXED THE INTERNAL TUBING.SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 6:55:0 | 1
324 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00113 | SUCTION PUMP PORTABLE, ELECTRONIC | SHILA BAVANKAR | 8208723521 | 26816 | 15-Apr-2019 | 08:50 |  |  | 16-Apr-2019 | 17:52 | NOT WORKING | CHECK AND FOUND SUCTION PRESSURE IS NOT GENERATED. PIPING IS BLOCKED. CLEANED THE PIPING. NOW, SUCTION MACHINE IS WORKING WITH DESIRED PRESSURE. | Closed | NO INFO | MEDIUM | 33:2:0 | 2
325 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00113 | SUCTION PUMP PORTABLE, ELECTRONIC | NEVADITA NURSE | 9011161387 | 23028 | 15-Dec-2018 | 08:50 |  |  | 17-Dec-2018 | 13:07 | NOT WORKING | CHECK AND FOUND LOW PRESSURE IS GENERATED. REFILLED THE OIL. NOW, DESIRED PRESSURE IS GENERATED. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 52:17:0 | 3
326 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00113 | SUCTION PUMP PORTABLE, ELECTRONIC | SHILA BAWANKAR | 8208733521 | 27089 | 25-Apr-2019 | 08:47 |  |  | 27-Apr-2019 | 16:09 | NOT WORKING | CHECK AND FOUND PRESSURE GENERATED IS NILL.PROBLEM FOUND IS OF INTERNAL TUBING. FIXED THE INTERNAL TUBINGS AND NOW DESIRED SUCTION PRESSURE IS GENERATED. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 55:22:0 | 3
327 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00113 | SUCTION PUMP PORTABLE, ELECTRONIC | BAVANKAR | 9422132247 | 20682 | 25-Sep-2018 | 09:30 |  |  | 01-Oct-2018 | 02:35 | NOT WORKING | CHECK AND FOUND SUCTION PRESSURE IS NOT GENRATED OPNED THE SUCTION MOTOR FIXED THE OTER CUVETTE OF SUCTION AND AIR FRAME . REOLING IS DONE.DESIRED FULL SUCTION PRESUURE IS NOW GENRATED SUCTION MACHINE IS NOW WORKING | Closed | NO INFO | MEDIUM | 137:5:0 | 7
328 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00113 | SUCTION PUMP PORTABLE, ELECTRONIC | BAVANKAR SHEELA | 9422132247 | 7675 | 29-Sep-2017 | 09:27 |  |  | 02-Oct-2017 | 11:56 | NOT WORKING | REPLACED NEW ON/OFF SWITCH AND NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 74:29:0 | 4
329 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00113 | SUCTION PUMP PORTABLE, ELECTRONIC | BAVANKAR SISTER | 820873352 | 25847 | 14-Mar-2019 | 08:17 |  |  | 14-Mar-2019 | 16:04 | NOT WORKING | CHECK AND FOUND ON/OFF TOGGLE SWITCH IS DEFECTIVE. CHANGED THE DEFECTIVE WITH THE NEW ON/OFF TOGGLE SWITCH. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 7:47:0 | 1
330 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00114 | SHADOWLESS LAMP STAND TYPE | SHEELA BAVANKAR | 8208733521 | 29461 | 04-Jul-2019 | 09:19 |  |  | 06-Jul-2019 | 10:31 | NOT WORKING | CALL IS RAISED OF DIFFERENT MACHINE. AT, PRESENT THIS PRESENT MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 49:12:0 | 3
331 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00116 | BP APPARATUS ANEROID | REHANA KHAN | 9158548938 | 3937 | 22-May-2017 | 10:25 |  |  | 25-May-2017 | 19:01 | NOT WORKING | REPLACE TUBE CONNECTOR ,CUFF,BLADDER ,BULB | Closed | RK1000552 | LOW | 80:36:0 | 4
332 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00116 | BP APPARATUS ANEROID | HARSHA BHUSARI | 9096022114 | 34465 | 30-Oct-2019 | 08:35 |  |  | 30-Oct-2019 | 17:48 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF, VALVE AND BULB IS MISSING. FIXED THE NEW BLADDER, CUFF, VALVE AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | RK1000552 | LOW | 9:13:0 | 1
333 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00117 | BP APPARATUS ANEROID | ANITA | 9249217699 | 27441 | 09-May-2019 | 08:35 |  |  | 13-May-2019 | 15:18 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS LEAK. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | R11000565 | LOW | 102:43:0 | 5
334 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00117 | BP APPARATUS ANEROID | Anita Limje | 9049217699 | 42248 | 14-May-2020 | 09:20 |  |  | 14-May-2020 | 10:15 | NOT WORKING | CHECK AND FOUND VALVE AND BULB IS HAVING LEAKAGE. FIXED THE NEW VALVE AND BULB. BP APPARATUS IS NOW WORKING. | Closed | R11000565 | LOW | 0:55:0 | 1
335 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00117 | BP APPARATUS ANEROID | JAYVANT | 8928453705 | 2753 | 24-Mar-2017 | 09:45 |  |  | 29-Mar-2017 | 11:10 | NOT WORKING | REPLACED THE BULB, CUFF, BLADDER, BULB CONNECTOR TUBING AND TUBING CONNECTOR | Closed | R11000565 | LOW | 121:25:0 | 6
336 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00117 | BP APPARATUS ANEROID | ANITA NIMJE | 9049217699 | 21598 | 24-Oct-2018 | 08:40 |  |  | 25-Oct-2018 | 15:35 | NOT WORKING | CHECK AND FOUND CUFF, BLADDER, BULB IS MISSING. FIXED ALL ABOVE NEW SPARES. NOW MACHINE IS WORKING. | Closed | R11000565 | LOW | 30:55:0 | 2
337 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00117 | BP APPARATUS ANEROID | ANITA NURSE | 9049217699 | 26614 | 09-Apr-2019 | 08:15 |  |  | 11-Apr-2019 | 11:06 | BULB ISSUE | CHECK AND FOUND BULB AND VALVE IS LEAK. CONNECTOR IS ALSO BROKEN. FIXED THE NEW BULB, VALVE AND CONNECTOR. B.P APPARATUS IS NOW WORKING. | Closed | R11000565 | LOW | 50:51:0 | 3
338 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00121 | SHADOWLESS LAMP STAND TYPE | SHITAL BAWANKAR | 8208733521 | 32214 | 07-Sep-2019 | 09:12 |  |  | 10-Sep-2019 | 15:35 | NOT WORKING | CHECK ANF FOUND 1 DOME OF SHADOWLESS LAMP IS NOT WORKING. BOTH THE HALOGEN BULB IS FUSED. CHANGED THE OLD DEFECTIVE HALOGEN BULB WITH THE NEW ONE. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 78:23:0 | 4
339 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00124 | SUCTION PUMP PORTABLE, ELECTRONIC | SHEELA | 8208733521 | 32235 | 09-Sep-2019 | 08:20 |  |  | 11-Sep-2019 | 12:03 | NOT WORKING | CHECK AND FOUND MOTOR WIRING IS BURNT AND HYDRAULIC VACUUM PUMP IS COMPLETELY JAMMED. GAVE THE MOTOR TO LOCAL VENDOR FOR REPAIR. AT PRESENT THIS BREAKDOWN CALL IS CLOSED. AS SOON AS MOTOR WILL BE RECEIVED FROM VENDOR, AGAIN THE CALL WOULD BE RAISED AND MACHINE WILL BE MADE FUNCTIONAL. | Closed | NO INFO | MEDIUM | 51:43:0 | 3
340 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00124 | SUCTION PUMP PORTABLE, ELECTRONIC | BAWANKAR SISTER | 8208733521 | 15889 | 18-May-2018 | 08:03 |  |  | 18-May-2018 | 10:40 | NOT WOKING | THE SUCTION TUBE IS BLOCKED WITH FLUIDS, CLEANED THE TUBING.NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO | MEDIUM | 2:37:0 | 1
341 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00124 | SUCTION PUMP PORTABLE, ELECTRONIC | BAWANKAR SISTER | 8208733521 | 15883 | 18-May-2018 | 08:30 |  |  | 18-May-2018 | 10:40 | NOT WORKING | THE SUCTION TUBE IS BLOCKED WITH FLUIDS,CLEANED THE TUBING.NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO | MEDIUM | 2:10:0 | 1
342 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00124 | SUCTION PUMP PORTABLE, ELECTRONIC | BAVANKAR NURSE | 8020873321 | 27664 | 16-May-2019 | 08:25 |  |  | 30-May-2019 | 01:30 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM MACHINE DOSE NOT TAKEN PRESSURE. REPAIR THE MOTOR WINDING NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 329:5:0 | 15
343 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00124 | SUCTION PUMP PORTABLE, ELECTRONIC | HARSHA BHUSARI | 9096022114 | 35772 | 26-Nov-2019 | 08:50 |  |  | 27-Nov-2019 | 14:51 | NOT WORKIG | CHECK AND FOUND WINDING OF THE MOTOR IS BURNT AND VACUUM PUMP(HYDRAULIC) IS JAMMED. SUCTION MOTOR IS REPAIRED THROUGH VENDOR. REPAIRED MOTOR IS BEING FIXED. SUCTION MOTOR IS WORKING WITH DESIRED SUCTION PRESSURE. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 30:1:0 | 2
344 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00125 | BP APPARATUS ANEROID | SHEELA BAWANKAR | 9422132247 | 25400 | 26-Feb-2019 | 09:10 |  |  | 26-Feb-2019 | 11:30 | NOT WORKING | BP APPRATUS IS NOT PROPERLY WORK, BECAUSE VALVE AND BULB IS UNUSEABLE. SO NEW BULB AND VALVE IS APPLY AND MASHINE IS WELL AND PROPERLY WORK AT PRESENT. | Closed | NO INFO | LOW | 2:20:0 | 1
345 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00126 | ANESTHESIA MACHINE (BOYLES APPARATUS) | REHANA KHAN | 9158548938 | 32627 | 17-Sep-2019 | 09:35 |  |  | 19-Sep-2019 | 14:40 | NOT WORKING | CHECK AND FOUND INTERNAL TUBINGS IS HAVING LEAKAGE AND DUE TO WHICH THERE IS LEAKAGE OF OXYGEN. CHANGED THE OLD DEFECTIVE DEGRADED TUBINGS WITH THE NEW ONE. LEAKAGE ISSUE IS BEING RESOLVED. ANESTHESIA MACHINE IS NOW WORKING. | Closed | NO INFO | HIGH | 53:5:0 | 3
346 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00126 | ANESTHESIA MACHINE (BOYLES APPARATUS) | BAVANKAR SHILA | 8208733521 | 28605 | 15-Jun-2019 | 08:32 |  |  | 19-Jun-2019 | 16:28 | NOT WORKING | CHECK AND FOUND INTERNAL TUBINGS IS BROKEN AND DUE TO WHICH LEAKAGE PROBLEM IS DETECTED. FIXED THE NEW LEGRIS TUBING(2.5 * 4) POLYURETHANE CALIBRATED TUBE. LEAKAGE PROBLEM IS SORTED OUT. ANESTHESIA MACHINE IS NOW WORKING. | Closed | NO INFO | HIGH | 103:56:0 | 5
347 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00126 | ANESTHESIA MACHINE (BOYLES APPARATUS) | HARSHA BHUSARI | 9096022114 | 32496 | 14-Sep-2019 | 09:25 |  |  | 14-Sep-2019 | 17:01 | NOT WORKING | CHECK AND FOUND INTERNAL TUBINGS IS LEAK AND DEGRADED. CHANGED THE OLD DEFECTIVE TUBINGS WITH THE NEW ONE. LEAKAGE PROBLEM IS RESOLVED. ANESTHESIA MACHINE IS NOW WORKING. | Closed | NO INFO | HIGH | 7:36:0 | 1
348 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00126 | ANESTHESIA MACHINE (BOYLES APPARATUS) | SHEELA BAVANKAR | 9422132247 | 28096 | 01-Jun-2019 | 09:00 |  |  | 03-Jun-2019 | 03:15 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND O2 TUBE HAS BROKEN. CLEAN LOCK CONNECTER AND RECONECT O2 TUBE NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | HIGH | 42:15:0 | 3
349 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00126 | ANESTHESIA MACHINE (BOYLES APPARATUS) | BAVANKAR | 9422132247 | 19582 | 25-Aug-2018 | 08:12 |  |  | 01-Sep-2018 | 11:35 | NOT WORKING | CHECK AND FOUND INTERNAL CONNECTING TUBING IS BROKEN. NEW TUBING IS BEING FIXED AND PASSAGE OF N2O AND O2 THROUGH THE TUBING IS BEING CHECKED. NO LEAKAGE OF O2 AND N2O IS PRESENT. ANESTHESIA MACHINE IS NOW WORKING. | Closed | NO INFO | HIGH | 171:23:0 | 8
350 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00128 | BINOCULAR MICROSCOPE | SHILA | 8208733521 | 38495 | 25-Jan-2020 | 08:21 |  |  | 25-Jan-2020 | 11:11 | NOT WORKING | CHECK AND FOUND IT IS ENT OPTICAL MICROSCOPE. MACHINE IS FULLY FUNCTIONAL WITH NO ISSUE. HANDS ON TRAINING IS GIVEN TO SISTER INCHARGE. CALL WAS LOGGED JUST TO CHECK THE MACHINE AS USER IS NOT ABLE TO TURN 'ON' THE MACHINE. | Closed | 09-F-7086 | MEDIUM | 2:50:0 | 1
351 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00132 | SUCTION PUMP PORTABLE, ELECTRONIC | BAVANKAR | 9422132247 | 20656 | 25-Sep-2018 | 08:05 |  |  | 25-Sep-2018 | 13:00 | NOT WORKING | CHECK AND FOUND LOW PRESSURE IS GENERATED.OILING OF MOTOR IS DONE.NOW DESIRED PRESSURE IS GENERATED SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 4:55:0 | 1
352 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00132 | SUCTION PUMP PORTABLE, ELECTRONIC | USHA SANGODE | 9421735065 | 5434 | 18-Jul-2017 | 10:20 |  |  | 19-Jul-2017 | 13:18 | NOT WORKING | REARRANGE THE VACUUM PRESSURE | Closed | NO INFO | MEDIUM | 26:58:0 | 2
353 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00132 | SUCTION PUMP PORTABLE, ELECTRONIC | BAVANKAR | 9422132247 | 20681 | 25-Sep-2018 | 09:30 |  |  | 31-Oct-2018 | 12:00 | NOT WORKING | CHECK AND FOUND MACHINE SUCTION IS LOW PRESSURE NOW MACHINE IS IN UNDER OBSERVATION,NOW MACHINE IS WORKING | Closed | NO INFO | MEDIUM | 866:30:0 | 37
354 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00132 | SUCTION PUMP PORTABLE, ELECTRONIC | REHANA KHAN | 9158548938 | 3634 | 11-May-2017 | 09:00 |  |  | 15-May-2017 | 13:21 | NOT WORKING | REPLACED THE MOTOR BEARING | Closed | NO INFO | MEDIUM | 100:21:0 | 5
355 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00132 | SUCTION PUMP PORTABLE, ELECTRONIC | SHILPA BAVANKAR` | 9422132247 | 25474 | 28-Feb-2019 | 09:04 |  |  | 01-Mar-2019 | 15:25 | NOT WORKING | CHECK ANF FOUND LOW SUCTION PRESSURE IS GENERATED. OILING IS BEING DONE. NOW PROPER SUCTION PRESSURE IS GENERATED.SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 30:21:0 | 2
356 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00132 | SUCTION PUMP PORTABLE, ELECTRONIC | SHEELA BAVANKAR | 9322132247 | 7695 | 02-Oct-2017 | 08:20 |  |  | 03-Oct-2017 | 16:45 | NOT WORKING | REPLACED NEW MOTOR BEARING, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 32:25:0 | 2
357 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00132 | SUCTION PUMP PORTABLE, ELECTRONIC | NEVADITA NURSE | 9011161387 | 23029 | 15-Dec-2018 | 08:50 |  |  | 17-Dec-2018 | 11:09 | NOT WORKING | CHECK AND FOUND POWER CORD IS FAULTY. CHANGED THE NEW POWER CORD. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 50:19:0 | 3
358 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00133 | OPERATION TABLE HYDRAULIC (MINOR) | BHAVANKAR | 9422132247 | 16283 | 31-May-2018 | 08:00 |  |  | 05-Jun-2018 | 01:41 | NOT WORKING | DOWN MOVEMENT JAMMED . REFILLED OT OIL | Closed | NO INFO | MEDIUM | 113:41:0 | 6
359 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00134 | SUCTION PUMP PORTABLE, ELECTRONIC | BAVANKAR SHEELA | 9422132247 | 7674 | 29-Sep-2017 | 09:25 |  |  | 02-Oct-2017 | 10:49 | NOT WORKING | ROTAR ASSEMBLY CLEANED AND FILLED OIL INTO AN MOTOR SECTION,NOW SUCTION DONE PROPERLY MACHINE HANDOVER TO USER AS WORKING CONDITION | Closed | NO INFO | MEDIUM | 73:24:0 | 4
360 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00134 | SUCTION PUMP PORTABLE, ELECTRONIC | HARSHA BHUSARI | 9096022114 | 36255 | 06-Dec-2019 | 08:50 |  |  | 06-Dec-2019 | 10:11 | NOT WORKING | CHECK AND FOUND INTERNAL TUBING IS OPEND DUE TO WHICH PRESSURE IS NOT GENRATED. FIXD THE INTERNAL TUBING SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 1:21:0 | 1
361 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00135 | SUCTION PUMP PORTABLE, ELECTRONIC | BAVANKAR SISTER | 9422132247 | 20609 | 22-Sep-2018 | 09:30 |  |  | 22-Sep-2018 | 18:30 | NOT WORKING | CHECK AND FOUND POWER ON/OFF SWITCH FAULTY.CHANGED THE NEW ON/OFF SWITCH.SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 9:0:0 | 1
362 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00137 | Anesthesia Work Station | SHILA NURSE | 9422132247 | 10577 | 15-Dec-2017 | 09:06 |  |  | 16-Dec-2017 | 15:03 | NOT WORKING | REPLACED THE NIBP CUFF | Closed | EA28000692 | HIGH | 29:57:0 | 2
363 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00137 | Anesthesia Work Station | HARSHA NURSE | 9096022114 | 37813 | 11-Jan-2020 | 08:10 |  |  | 12-Jan-2020 | 17:23 | NOT WORKING | CHECK AND FOUND CYLINDER LOCKING KNOB IS FAULTY. OXYGEN SENSOR IS ALSO FAULTY AS DURING CALIBRATION OF OXYGEN, TEST FAILS. OXYGEN MEASURED IS INACCURATE. FIXED THE LOCKING KNOB THROUGH WELDING. NEW OXYGEN SENSOR IS FIXED AND CALIBRATION IS DONE. ANESTHESIA MACHINE IS NOW WORKING. | Closed | EA28000692 | HIGH | 33:13:0 | 2
364 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00137 | Anesthesia Work Station | NIVIDHITA JOHN | 9011167387 | 35298 | 18-Nov-2019 | 08:33 |  |  | 19-Nov-2019 | 14:57 | NOT WORKING | CHECK AND FOUND KNOB ATTACHED WITH CYLINDER YOKES IS LOOSE AND OXYGEN CYLINDER COULD NOT BE TIGHTENED. FIXED THE KNOB OF CYLINDER YOKES. NOW OXYGEN CYLINDER CAN BE FIXED AND TIGHTENED. ANESTHESIA WORK STATION IS NOW WORKING. | Closed | EA28000692 | HIGH | 30:24:0 | 2
365 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00137 | Anesthesia Work Station | SHEELA BAWANKAR | 9422132247 | 25401 | 26-Feb-2019 | 09:10 |  |  | 03-Mar-2019 | 12:56 | O2 SENSOR ISSUE | CHECK AND FOUND OXYGEN SENSOR IS DEFECTIVE. CHANGED THE NEW OXYGEN SENSOR. CALIBRATION HAS BEEN DONE. ANESTHESIA MACHINE IS NOW WORKING. | Closed | EA28000692 | HIGH | 123:46:0 | 6
366 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00138 | ANESTHESIA MACHINE (BOYLES APPARATUS) | URMILA JAMBHULKAR | 9371876261 | 760 | 06-Jan-2017 | 12:38 |  |  | 10-Jan-2017 | 18:30 | NOT WORKING | WASHER REPLACED AND RECTIFIED | Closed | HT12L1324 | HIGH | 101:52:0 | 5
367 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00138 | ANESTHESIA MACHINE (BOYLES APPARATUS) | NIVIDHITA JOHN | 9011167387 | 35299 | 18-Nov-2019 | 08:35 |  |  | 19-Nov-2019 | 16:01 | NOT WORKING | CHECK AND FOUND KNOB ATTACHED WITH CYLINDER YOKES IS BROKEN. LEAKAGE FROM AUXILLARY OXYGEN FLOWMETER ATTACHED WITH ANESTHESIA MACHINE. FIXED THE NEW KNOB OF CYLINDER YOKES. LEAKAGE PROBLEM FROM AUXILLARY OXYGEN FLOWMETER IS RECTIFIED. ANESTHESIA MACHINE IS WORKING. | Closed | HT12L1324 | HIGH | 31:26:0 | 2
368 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00139 | PULSE OXIMETER LINE POWERED | SMITA DHALE | 9637606658 | 14534 | 04-Apr-2018 | 09:13 |  |  | 05-Apr-2018 | 13:09 | NOT WORKING | ACTUAL MACHINE IS PULSE OXIMETER NOT A RADIANT WARMER SO WE CHECKED AND FOUND ADULT FINGER PROBE HAS AN PROBLEM THEN WE REPLACED WITH NEW PROBE, MACHINE IS WORKING CONDITION. | Closed | 1211 | MEDIUM | 27:56:0 | 2
369 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00140 | MONITOR | SHILA BAWANKAR | 9422132247 | 13095 | 21-Feb-2018 | 09:28 |  |  | 22-Feb-2018 | 15:53 | NOT WORKING | REPLACED THE SPO2 PROBE | Closed | NO INFO | MEDIUM | 30:25:0 | 2
370 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00140 | MONITOR | SHEELA BAWANKAR | 8208733521 | 34019 | 17-Oct-2019 | 09:26 |  |  | 17-Oct-2019 | 15:57 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS HAVING LEAKAGE. FIXED THE NEW NIBP CUFF. NOW, BLOOD PRESSURE IS MEASURED WITHOUT ANY ERROR. MONITOR IS NOW WORKING. | Closed | NO INFO | MEDIUM | 6:31:0 | 1
371 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00140 | MONITOR | HARSHA | 9496022114 | 39786 | 20-Feb-2020 | 09:13 |  |  | 20-Feb-2020 | 10:00 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS LEAKAGE.CHANGE THE NIBP CUFF NEW ONE. NOW MONITOR IS WORKING WITH NIBP CUFF. | Closed | NO INFO | MEDIUM | 0:47:0 | 1
372 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00140 | MONITOR | NEVADITA NURSE | 9011161387 | 23027 | 15-Dec-2018 | 08:50 |  |  | 22-Dec-2018 | 11:38 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS DEFECTIVE. CHANGE THE NIBP CUFF L, NOW MACHINE IS WORKING | Closed | NO INFO | MEDIUM | 170:48:0 | 8
373 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00141 | DIATHERMY MACHINE (ELECTRIC CAUTERY) | SHILA NURSE | 9422132247 | 23952 | 15-Jan-2019 | 08:06 |  |  | 18-Jan-2019 | 16:05 | NOT WORKING | CHECK AND FOUND DISPLAY CABLE CONNECTOR IS LOOSE. FIXED THE CABLE CONNECTOR, RECHECKED THE ENTIRE BOARD AND ITS CONNECTION. NOW, DISPLAY IS PROPER WITHOUT ANY DISRUPTION. CAUTERY MACHINE IS NOW WORKING. | Closed | ASG-M1303682 | MEDIUM | 79:59:0 | 4
374 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00147 | ANESTHESIA MACHINE (BOYLES APPARATUS) | REHANA KHAN | 9158548938 | 32625 | 17-Sep-2019 | 09:35 |  |  | 19-Sep-2019 | 11:34 | O2 LEAKAGE | CHECK AND FOUND THERE IS LEAKAGE OF OXYGEN THROUGH INTERNAL TUBINGS. INTERNAL TUBING IS DEGRADED AND THERE IS LEAKAGE. CHANGED THE INTERNAL POLYURETHANE CALIBRATED TUBE. LEAKAGE PROBLEM IS RESOLVED. ANESTHESIA MACHINE IS NOW WORKING. | Closed | NO INFO | HIGH | 49:59:0 | 3
375 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00147 | ANESTHESIA MACHINE (BOYLES APPARATUS) | URMILA JAMBHULKAR | 9371876261 | 761 | 06-Jan-2017 | 12:39 |  |  | 10-Jan-2017 | 18:30 | NOT WORKING | WASHER REPLACED AND RESOLVED | Closed | NO INFO | HIGH | 101:51:0 | 5
376 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00150 | BP APPARATUS ANEROID | SHEELA BAVANKAR | 9422132247 | 7377 | 21-Sep-2017 | 09:43 |  |  | 25-Sep-2017 | 11:39 | NOT WORKING | REPLACED THE BULB | Closed | RM1000152 | LOW | 97:56:0 | 5
377 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00150 | BP APPARATUS ANEROID | HARSHA BHUSARI | 9096022114 | 38695 | 29-Jan-2020 | 08:32 |  |  | 29-Jan-2020 | 10:03 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND BULB. FIXED THE NEW BLADDER AND BULB. BP APPARATUS IS NOW WORKING. | Closed | RM1000152 | LOW | 1:31:0 | 1
378 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00150 | BP APPARATUS ANEROID | SHEELA BAWANKAR | 9422132247 | 25399 | 26-Feb-2019 | 09:10 |  |  | 27-Feb-2019 | 11:08 | NOT WORKING | BP CUFF BLADDER IS LEAK. SO NEW BP CUFF BLADDER IS APPLY AND WE'LL AND PROPERLY WORK AT PRESENT. | Closed | RM1000152 | LOW | 25:58:0 | 2
379 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00153 | BP APPARATUS ANEROID | SHEELA BAVANKAR | 9422132247 | 7378 | 21-Sep-2017 | 09:43 |  |  | 25-Sep-2017 | 11:09 | NOT WORKING | REPLACED THE BLADDER AND TUBING CONNECTOR | Closed | 19194 | LOW | 97:26:0 | 5
380 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00154 | SUCTION PUMP PORTABLE, ELECTRONIC | SHILA BAVANKAR | 9422132247 | 7321 | 20-Sep-2017 | 08:25 |  |  | 26-Sep-2017 | 14:50 | NOT WORKING | REPLACED NEW ON/OFF SWITCH AND POUR SOME OIL INTO AN MACHINE NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 150:25:0 | 7
381 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00154 | SUCTION PUMP PORTABLE, ELECTRONIC | SHILA NURSE | 9422132247 | 10576 | 15-Dec-2017 | 09:06 |  |  | 16-Dec-2017 | 15:04 | NOT WORKING | REWIND THE MOTOR | Closed | NO INFO | MEDIUM | 29:58:0 | 2
382 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00154 | SUCTION PUMP PORTABLE, ELECTRONIC | USHA SANGODE | 9421735065 | 5435 | 18-Jul-2017 | 10:25 |  |  | 20-Jul-2017 | 11:08 | NOT WORKING | REPLACE 6A ON/OFF SWITCH | Closed | NO INFO | MEDIUM | 48:43:0 | 3
383 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00155 | MONITOR | APARNA JANJAL | 9423308188 | 23283 | 22-Dec-2018 | 08:05 |  |  | 26-Dec-2018 | 16:35 | PROBE ISSUE | CHECKED AND FOUND SPO2 PROBE FAULTY.CHANGED THE NEW SPO2 PROBE(6PIN,DUAL LOCK). SPO2 READING IS MONITORED ALONG WITH PULSE RATE. READING OBTAINED IS SATISFACTORY. MONITOR IS NOW WORKING. | Closed | NO INFO | MEDIUM | 104:30:0 | 5
384 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00155 | MONITOR | DR YELNE | 9049155994 | 4503 | 15-Jun-2017 | 09:12 | 16-Jun-2017 | 16:39 | 20-Jun-2017 | 10:22 | NOT WORKING | REPLACE NEW AUTOMATIC SINGLE TUBE NIBP CUFF | Closed | NO INFO | MEDIUM | 121:10:0 | 6
385 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00155 | MONITOR | BAVANKAR NURSE | 9422132247 | 19434 | 22-Aug-2018 | 08:02 |  |  | 27-Aug-2018 | 11:57 | NOT WORKING | CHECK AND FOUND B.P MODULE IS DETACHED AND WENT INSIDE THE MACHINE.FIXED THE B.P MODULE B.P IS MEASURED .MULTIPARAMONITOR IS WORKING | Closed | NO INFO | MEDIUM | 123:55:0 | 6
386 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00155 | MONITOR | SHILA BAWANKAR | 9422132247 | 13096 | 21-Feb-2018 | 09:28 |  |  | 24-Feb-2018 | 15:31 | NOT WORKING | REPLACED THE SPO2 PROBE | Closed | NO INFO | MEDIUM | 78:3:0 | 4
387 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00156 | BP APPARATUS ANEROID | SHILA BAVANKAR | 9422132247 | 10954 | 26-Dec-2017 | 09:06 |  |  | 27-Dec-2017 | 11:25 | NOT WORKING | REPLACED THE BLADDER | Closed | RM100546 | LOW | 26:19:0 | 2
388 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00156 | BP APPARATUS ANEROID | SHILA BAVAN | 8208733521 | 26866 | 16-Apr-2019 | 09:05 |  |  | 16-Apr-2019 | 17:50 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF, VALVE AND BULB IS MISSING. FIXED THE NEW BLADDER, CUFF, VALVE AN BULB. B. P APPARATUS IS NOW WORKING. | Closed | RM100546 | LOW | 8:45:0 | 1
389 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00156 | BP APPARATUS ANEROID | REHANA KHAN | 9158548938 | 3938 | 22-May-2017 | 10:25 |  |  | 25-May-2017 | 02:05 | NOT WORKING | REPLACE CUFF,BLADDER ,BULB,BULB CONNECTOR | Closed | RM100546 | LOW | 63:40:0 | 4
390 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00157 | BP APPARATUS ANEROID | SHEELA BAVANKAR | 9422132247 | 7376 | 21-Sep-2017 | 09:43 |  |  | 25-Sep-2017 | 11:14 | NOT WORKING | REPLACED THE BULB AND TUBING | Closed | RK1000563 | LOW | 97:31:0 | 5
391 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00157 | BP APPARATUS ANEROID | BHAVNA SISTER | 9405949739 | 20399 | 15-Sep-2018 | 08:32 |  |  | 18-Sep-2018 | 11:07 | NOT WORKING | CHECK AND FOUND VALVE AND CONNECTING PIPE IS DEFECTIVE .CHANGED THE NEW VALVE AND CONNECTING PIPES B.P APPRATUS IS NOW WORKING | Closed | RK1000563 | LOW | 74:35:0 | 4
392 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00157 | BP APPARATUS ANEROID | BAVANKAR | 9422132247 | 19583 | 25-Aug-2018 | 08:12 |  |  | 28-Aug-2018 | 12:05 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT BLADDER AND VALVE IS LEAK BULB IS ALSO WORN OUT.CHANGED CUFF BLADDER VALVE AND BULB.BP APPARTUS IS NOW WORKING | Closed | RK1000563 | LOW | 75:53:0 | 4
393 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00157 | BP APPARATUS ANEROID | HARSHA BHUSARI | 9096022114 | 38254 | 22-Jan-2020 | 09:10 |  |  | 22-Jan-2020 | 10:51 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS LEAK. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | RK1000563 | LOW | 1:41:0 | 1
394 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00158 | ANESTHESIA MACHINE (BOYLES APPARATUS) | REHANA KHAN | 9158548938 | 32626 | 17-Sep-2019 | 09:35 |  |  | 19-Sep-2019 | 12:37 | O2 LEAKAGE | CHECK AND FOUND THERE IS LEAKAGE IN INTERNAL TUBINGS AND DUE TO WHICH LEAKAGE OF OXYGEN IS TAKING PLACE. CHANGED THE OLD DEGRADED TUBINGS WITH THE NEW ONE. LEAKAGE PROBLEM IS RESOLVED. ANESTHESIA MACHINE IS NOW WORKING. | Closed | O-03-0123 | HIGH | 51:2:0 | 3
395 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00158 | ANESTHESIA MACHINE (BOYLES APPARATUS) | JOHN NURSE | 9011161387 | 31792 | 26-Aug-2019 | 08:05 |  |  | 30-Aug-2019 | 14:16 | NOT WORKING | CHECK AND FOUND ALL INTERNAL TUBING IS DEGRADED AND BROKEN AND THERE IS ENTIRE PROBLEM OF LEAKAGE. MR IS BEING RAISED AND TUBING IS BEING ORDERED. THERE IS SLIGHT DELAY IN RECEIVING THE SPARES AS TUBING IS NOT EASILY AVAILABLE. AT PRESENT THIS PRESENT BREAKDOWN IS BEING CLOSED. AS SOON AS THE SPARES WOULD BE AVAILABLE, AGAIN THE COMPLAINT WOULD BE RAISED AND MACHINE WOULD BE MADE FUNCTIONAL. | Closed | O-03-0123 | HIGH | 102:11:0 | 5
396 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00160 | CARDIOTOCOGRAPHY MACHINE | ANITA | 9049721760 | 43940 | 08-Jul-2020 | 08:10 |  |  | 09-Jul-2020 | 10:30 | NOT WORKING | PROBLEM REPORTED THAT CHARGING COULDN'T BE DONE. CHECK AND FOUND MALE CONNECTOR OF POWER ADAPTER IS FAULTY. REPAIRED THE SAME. NOW, POWER ADAPTER IS WORKING AND ALSO THE CARDIOTOCOGRAPHY MACHINE. | Closed | JA3C10028 | MEDIUM | 26:20:0 | 2
397 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00160 | CARDIOTOCOGRAPHY MACHINE | ANITA LINGE | 9049217699 | 10124 | 06-Dec-2017 | 10:20 |  |  | 07-Dec-2017 | 11:54 | NOT WORKING | CHANGE FETAL DOPPLER PROBE | Closed | JA3C10028 | MEDIUM | 25:34:0 | 2
398 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00162 | BP APPARATUS ANEROID | RAHATE | 9822694627 | 34741 | 06-Nov-2019 | 09:18 |  |  | 08-Nov-2019 | 13:42 | NOT WORKING | CHECK AND FOUND BLADDER, BULB, CUFF AND VALVE IS MISSING.FIXED THE NEW BLADDER, BULB, CUFF AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 52:24:0 | 3
399 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00164 | SUCTION PUMP PORTABLE, ELECTRONIC | ANITA LINGE | 9049217699 | 10128 | 06-Dec-2017 | 10:20 |  |  | 11-Dec-2017 | 12:59 | NOT WORKING | DONE REWINDING ON MOTOR AND ALSO REPLACED NEW BEARINGS. | Closed | NO INFO | MEDIUM | 122:39:0 | 6
400 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00169 | BP APPARATUS ANEROID | ARCHANA VASNIK | 8378035077 | 12026 | 20-Jan-2018 | 09:22 |  |  | 20-Jan-2018 | 18:03 | NOT WORKING | REPLACED TUBING AND CONNECTOR | Closed | NO NFO | LOW | 8:41:0 | 1
401 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00171 | ADULT WEIGHING MACHINE | SUNITA DARWATE | 9604346238 | 11430 | 09-Jan-2018 | 09:00 |  |  | 10-Jan-2018 | 23:52 | NOT WORKING | RECONNECT THE FSR CABLE CONNECTION | Closed | NO INFO | LOW | 38:52:0 | 2
402 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00171 | ADULT WEIGHING MACHINE | T S UKEY | 9145309697 | 7287 | 19-Sep-2017 | 11:27 |  |  | 20-Sep-2017 | 18:10 | NOT WORKING | REPLACE THE ON/OFF SWITCH | Closed | NO INFO | LOW | 30:43:0 | 2
403 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00171 | ADULT WEIGHING MACHINE | PRAJAVATI K | 9145309697 | 17215 | 27-Jun-2018 | 08:30 |  |  | 03-Jul-2018 | 13:30 | NOT WORKING | CHECK AND FOUND SOME INTERNAL WIRING ISSUES. FIX THE ISSUES AND DID CALIBRATION, NOW WEIGHING MACHINE WORKING CONDITION | Closed | NO INFO | LOW | 149:0:0 | 7
404 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00173 | BP APPARATUS ANEROID | SUNITA DARWATE | 9604346238 | 10102 | 06-Dec-2017 | 09:05 |  |  | 07-Dec-2017 | 11:03 | NOT WORKING | REPLACED THE BULB | Closed | NO NFO | LOW | 25:58:0 | 2
405 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00174 | NEBULIZER ELECTRIC | PRIYANKA MISHRA | 9458203936 | 11541 | 10-Jan-2018 | 09:50 |  |  | 11-Jan-2018 | 14:39 | NOT WORKING | SOLDERING THE WIRE CONNECTION | Closed | NO INFO | MEDIUM | 28:49:0 | 2
406 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00178 | SUCTION PUMP PORTABLE, ELECTRONIC | PRAJAVATI NURSE | 9145309697 | 18253 | 21-Jul-2018 | 08:45 |  |  | 24-Jul-2018 | 14:30 | NOT WORKING | CHECK AND FOUND SUCTION PRESSURE NOT GENERATING. CHECK ALL TUBING FOUND JAMMED AND BLOCK WITH DUST. CLEAN THE TUBINGS, NOW SUCTION PRESSURE IS GENERATED, MACHINE IS WORKING SATISFACTORY. | Closed | NO INFO | MEDIUM | 77:45:0 | 4
407 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00179 | BP APPARATUS ANEROID | PATIL | 9420866915 | 11445 | 09-Jan-2018 | 09:20 |  |  | 10-Jan-2018 | 15:12 | NOT WORKING | REPLACED THE BLADDER, METAL VALVE, BULB, CUFF , CONNECTOR | Closed | NO NFO | LOW | 29:52:0 | 2
408 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00179 | BP APPARATUS ANEROID | MAYA WASNIK | 8149503946 | 7275 | 19-Sep-2017 | 09:15 |  |  | 20-Sep-2017 | 11:33 | NOT WORKING | REPLACE GLASS TUBE AND METAL CONNECTOR | Closed | NO NFO | LOW | 26:18:0 | 2
409 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00179 | BP APPARATUS ANEROID | GHARME NURSE | 9511971047 | 10591 | 15-Dec-2017 | 09:18 |  |  | 16-Dec-2017 | 23:35 | NOT WORKING | REFILLED A MERCURY | Closed | NO NFO | LOW | 38:17:0 | 2
410 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00179 | BP APPARATUS ANEROID | MAYA | 8149503946 | 36673 | 17-Dec-2019 | 09:10 |  |  | 17-Dec-2019 | 15:48 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS LEAK. FIXED THE NEW BLADDER AND BULB. B.P APPARATUS IS NOW WORKING. | Closed | NO NFO | LOW | 6:38:0 | 1
411 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00179 | BP APPARATUS ANEROID | PRADIP VACHNIR | 9433605761 | 3322 | 25-Apr-2017 | 09:02 |  |  | 01-May-2017 | 10:17 | NOT WORKING | REPLACED NEW RUBBER BLADER | Closed | NO NFO | LOW | 145:15:0 | 7
412 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00180 | SUCTION PUMP PORTABLE, ELECTRONIC | GHARME NURSE | 9511971047 | 10575 | 15-Dec-2017 | 09:04 |  |  | 15-Dec-2017 | 23:32 | NOT WORKING | RECONNECT THE POWER CARD PROPERLY | Closed | NO INFO | MEDIUM | 14:28:0 | 1
413 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00180 | SUCTION PUMP PORTABLE, ELECTRONIC | PRADIP WASNIK | 9423605761 | 856 | 10-Jan-2017 | 10:05 |  |  | 12-Jan-2017 | 09:30 | NOT WORKING | REINSERT THE SUCTION TUBE PROPERLY | Closed | NO INFO | MEDIUM | 47:25:0 | 3
414 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00181 | PULSE OXIMETER LINE POWERED | GUMARE NURSE | 9511971047 | 24809 | 08-Feb-2019 | 09:10 |  |  | 08-Feb-2019 | 15:07 | NOT WORKING | CHECK AND FOUND MACHINE IS NOT TURNING ON. POWER CABLE IS DEFECTIVE. CHANGED THE OLD AND DEFECTIVE ONE WITH NEW POWER CABLE. PULSE OXIMETER IS NOW WORKING | Closed | 4131864 | MEDIUM | 5:57:0 | 1
415 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00181 | PULSE OXIMETER LINE POWERED | AMARPALI MATGHARE | 9765334027 | 11428 | 09-Jan-2018 | 09:00 |  |  | 10-Jan-2018 | 23:48 | NOT WORKING | REPLACED THE SPO2 PROBE | Closed | 4131864 | MEDIUM | 38:48:0 | 2
416 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00181 | PULSE OXIMETER LINE POWERED | DHUMBARE NURSE | 9511971147 | 26902 | 18-Apr-2019 | 09:05 |  |  | 21-Apr-2019 | 15:13 | READING ISSUE | CHECK AND FOUND SPO2 PROBE FAULTY. CHANGED THE OLD WITH NEW D-TYPE, 9 PIN SPO2 PROBE. PULSE OXIMETER IS NOW WORKING. | Closed | 4131864 | MEDIUM | 78:8:0 | 4
417 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00181 | PULSE OXIMETER LINE POWERED | MAYA WASNIK | 8149503946 | 18002 | 17-Jul-2018 | 09:06 |  |  | 19-Jul-2018 | 10:00 | NOT WORKING | NO PULSE BEING OBTAINED AND COMPLAINT HAS BEEN LOGGED SAME. CHECK AND FOUND SPO2 PROBE NOT PROPERLY CONNECTED. CONNECTED SAME SPO2 SPO2 READING IS BEING MONITORED. MACHINE IS IN WORKING CONDITION. | Closed | 4131864 | MEDIUM | 48:54:0 | 3
418 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00181 | PULSE OXIMETER LINE POWERED | MAYA VASNIK | 8149503946 | 16054 | 23-May-2018 | 09:15 |  |  | 25-May-2018 | 02:42 | NOT WORKING | REPLACED NEW SP O2 PROBE . NOW THE MACHINE IS IN WORKING CONDITION | Closed | 4131864 | MEDIUM | 41:27:0 | 3
419 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00181 | PULSE OXIMETER LINE POWERED | VASMIK | 8149503946 | 32260 | 09-Sep-2019 | 09:23 |  |  | 11-Sep-2019 | 14:16 | NOT WORKING | CHECK AND FOUND A/D CONVERTER PCB IS DEFECTIVE AND DUE TO WHICH SPO2 PROBE IS NOT WORKING. SENT THE PULSE OXIMETER TO MUMBAI PCB LAB FOR REPAIR. AT PRESENT THIS BREAKDOWN IS CLOSED. AS SOON AS REPAIRED MACHINE WILL BE RECEIVED FROM MUMBAI, AGAIN THE CALL WILL BE RAISED AND MACHINE WILL BE HANDED TO CONCERNED DEPARTMENT. | Closed | 4131864 | MEDIUM | 52:53:0 | 3
420 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00182 | NEBULIZER ELECTRIC | MAYA WASNIK | 8149503946 | 7276 | 19-Sep-2017 | 09:10 |  |  | 20-Sep-2017 | 11:20 | NOT WORKING | RECONNECT THE POWER CABLE | Closed | NO INFO | MEDIUM | 26:10:0 | 2
421 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00184 | SYRINGE NEEDLE DESTROYER | MAYA WASNIK | 8149503946 | 35030 | 12-Nov-2019 | 08:49 |  |  | 12-Nov-2019 | 15:23 | NOT WORKING | CHECK AND FOUND POWER CORD IS DEFECTIVE. CHANGED THE DEFECTIVE POWER CORD WITH THE NEW ONE. SYRINGE NEEDLE DESTROYER IS NOW WORKING. | Closed | NO INFO | LOW | 6:34:0 | 1
422 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00185 | BP APPARATUS ANEROID | HEMLATA | 7875861287 | 4922 | 29-Jun-2017 | 08:30 |  |  | 01-Jul-2017 | 09:25 | NOT WORKING | REPLACED THE BULB, CONNECTOR, BLADDER AND CUFF | Closed | NO NFO | LOW | 48:55:0 | 3
423 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00185 | BP APPARATUS ANEROID | PRADIP WASNIK | 9423605761 | 1099 | 19-Jan-2017 | 10:01 |  |  | 21-Jan-2017 | 11:17 | NOT WORKING | BLADDER AND RUBBER WASHER CHANGED | Closed | NO NFO | LOW | 49:16:0 | 3
424 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00189 | BP APPARATUS ANEROID | SUNANDA | 8007445066 | 28517 | 13-Jun-2019 | 08:40 |  |  | 13-Jun-2019 | 15:43 | NOT WORKING | CHECK AND FOUND BLADDER, BULB, VALVE IS LEAK AND CUFF IS WORN OUT. FIXED THE NEW BLADDER, BULB, VALVE AND CUFF. B. P APPARATUS IS NOW WORKING. | Closed | 254018 | LOW | 7:3:0 | 1
425 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00189 | BP APPARATUS ANEROID | PRADIP WASNIK | 9423605761 | 2559 | 17-Mar-2017 | 09:48 |  |  | 21-Mar-2017 | 11:06 | NOT WORKING | REPLACED THE BULB, CUFF AND BLADDER | Closed | 254018 | LOW | 97:18:0 | 5
426 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | SUNANDA | 8007445066 | 28500 | 13-Jun-2019 | 08:05 |  |  | 18-Jun-2019 | 13:29 | NOT WORKING | CHECK AND FOUND DOPPLER PROBE IS FAULTY. CHANGED THE NEW DOPPLER PROBE. CARDIOTOCOGRAPHY MACHINE IS NOW WORKING. | Closed | JA3C100219 | MEDIUM | 125:24:0 | 6
427 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | BHAGWAT NURSE | 9309169328 | 17740 | 11-Jul-2018 | 09:00 |  |  | 13-Jul-2018 | 15:26 | NOT WORKING | CHECK AND FOUND PROBE IS NOT FITTED PROPERLY. FIT IT PROPERLY, NOW MACHINE IN WORKING CONDITION. | Closed | JA3C100219 | MEDIUM | 54:26:0 | 3
428 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | MONALI KATAWALE | 8600834483 | 21091 | 05-Oct-2018 | 08:18 |  |  | 09-Oct-2018 | 15:10 | NOT WORKING | CHECK AND FOUND ON/OFF SWITCH IS NOT WORKING. REPAIRED THE BUTTON TYPE ON/OFF SWITCH. CARDIOTOCOGRAPHY MACHINE IS NOW WORKING. | Closed | JA3C100219 | MEDIUM | 102:52:0 | 5
429 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | MISHRAM | 8007729587 | 40392 | 03-Mar-2020 | 09:25 |  |  | 06-Mar-2020 | 11:39 | NOT WORKING | CHECK AND FOUND POWER ADAPTER IS FAULTY. FIXED THE NEW POWER ADAPTER OF RATING 15V,4A. CARDIOTOCOGRAPHY MACHINE IS NOW WORKING. | Closed | JA3C100219 | MEDIUM | 74:14:0 | 4
430 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | DAHEKAR NURSE | 9423672785 | 41113 | 19-Mar-2020 | 08:55 |  |  | 19-Mar-2020 | 11:00 | NOT WORKING | CHECK AND FOUND POWER CABLE IS SHORT, CHANGE THE POWER CABLE FIX THE ADAPTER AND CONNECT TO MACHINE, NOW MACHINE IS WORKING PROPERLY | Closed | JA3C100219 | MEDIUM | 2:5:0 | 1
431 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | UMBERKAR | 8007445066 | 38597 | 28-Jan-2020 | 08:02 |  |  | 28-Jan-2020 | 16:50 | NOT WORKING | CHECK AND FOUND POWER CABLE IS LOOSE TO ADAPTER AND DUE TO WHICH MACHINE IS NOT TURNING ON. FIXED THE LOOSE CONNECTION OF POWER CABLE AND ADAPTER. CARDIOTOCOGRAPHY MACHINE IS NOW WORKING. | Closed | JA3C100219 | MEDIUM | 8:48:0 | 1
432 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | MINAKSHI BAVANKUDE | 7507957370 | 18992 | 10-Aug-2018 | 09:14 |  |  | 13-Aug-2018 | 15:14 | NOT WORKING | CHECK AND FOUND POWER ON/OFF IS NOT WORKING RESOLDERING THE PUB OF POWER ON/OFF CARDIOTOCOGRAPHY MACHINE IS NOW WORKING | Closed | JA3C100219 | MEDIUM | 78:0:0 | 4
433 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | UMBARKAR | 8007445066 | 27151 | 27-Apr-2019 | 10:30 |  |  | 01-May-2019 | 18:00 | NOT WORKING | CHECK AND FOUND ON/OFF POWER SWITCH IS FAULTY. REPAIRED THE POWER ON/OFF SWITCH. CARDIOTOCOGRAPHY MACHINE IS NOW WORKING. | Closed | JA3C100219 | MEDIUM | 103:30:0 | 5
434 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | UBARKAR SISTER | 8007445066 | 16521 | 07-Jun-2018 | 08:00 |  |  | 08-Jun-2018 | 17:52 | NOT WORKING | PROBE PINS ARE BENT . STRAIGHTEN THE PROBE PINS. NOW THE MACHINE IS IN WORKING CONDITION | Closed | JA3C100219 | MEDIUM | 33:52:0 | 2
435 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | Mishra | 8007729587 | 41589 | 11-Apr-2020 | 08:30 |  |  | 11-Apr-2020 | 11:42 | NOT WORKING | CHECK AND FOUND POWER CABLE IS FAULTY. FIXED THE NEW POWER CABLE. POWER SUPPLY ISSUE IS RESOLVED. CARDIOTOCOGRAPHY MACHINE IS NOW WORKING. | Closed | JA3C100219 | MEDIUM | 3:12:0 | 1
436 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | BHAGWAT KOKILA | 9921463322 | 20418 | 16-Sep-2018 | 09:01 |  |  | 18-Sep-2018 | 12:06 | NOT WORKING | CHECK AND FOUND ON OFF SWITCH IS NOT WORKING.SCREW PROBLEM IS ALSO RECTIFIED CARDIOTOCOGRAPHY IS NOW WORKING | Closed | JA3C100219 | MEDIUM | 51:5:0 | 3
437 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | SUNANDA UMBARKER | 8007445006 | 22622 | 05-Dec-2018 | 08:00 |  |  | 10-Dec-2018 | 11:05 | NOT WORKING | CHECK AND FOUND ON/OFF BUTTON IS NOT WORKING BUTTON IS LOOSE. FIXED THE BUTTON PROPERLY NOW MACHINE IS WORKING GOOD | Closed | JA3C100219 | MEDIUM | 123:5:0 | 6
438 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00191 | CARDIOTOCOGRAPHY MACHINE | PRIYA NURSE | 9604098679 | 11185 | 01-Jan-2018 | 09:40 |  |  | 04-Jan-2018 | 11:33 | NOT WORKING | RECONNECTED PROBE CONNECTION | Closed | JA3C100219 | MEDIUM | 73:53:0 | 4
439 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00192 | DIGITAL BP APPARATUS | MS GUNJAN | 8421137019 | 4641 | 20-Jun-2017 | 09:35 |  |  | 24-Jun-2017 | 10:49 | NOT WORKING | REPLACED THE BATTERY | Closed | 15066801984 | LOW | 97:14:0 | 5
440 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00192 | DIGITAL BP APPARATUS | REHANA KHAN | 9158548938 | 3939 | 22-May-2017 | 10:25 |  |  | 25-May-2017 | 01:09 | NOT WORKING | REPLACE BULB | Closed | 15066801984 | LOW | 62:44:0 | 4
441 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00192 | DIGITAL BP APPARATUS | NAGDEVE SISTER | 9403379340 | 14341 | 29-Mar-2018 | 09:10 |  |  | 30-Mar-2018 | 12:56 | NOT WORKING | REPLACED NEW GLASS TUBE AND MERCURY TANK WITH MERCURY, MACHINE IS WORKING CONDITION. | Closed | 15066801984 | LOW | 27:46:0 | 2
442 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00196 | SUCTION PUMP PORTABLE, ELECTRONIC | SONALI BADGE | 7038409193 | 9504 | 22-Nov-2017 | 09:07 |  |  | 22-Nov-2017 | 16:04 | NOT WORKING | CLEAN THE DUST PARTICLES | Closed | NO INFO | MEDIUM | 6:57:0 | 1
443 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00197 | PULSE OXIMETER LINE POWERED | GUNJAN PALI | 9226201417 | 14797 | 11-Apr-2018 | 08:24 |  |  | 12-Apr-2018 | 09:55 | NOT WORKING | RECONNECTED ALL DISCONNECTING WIRES. NOW MACHINE IS WORKING CONDITION. | Closed | 4131860 | MEDIUM | 25:31:0 | 2
444 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00197 | PULSE OXIMETER LINE POWERED | SANGITA TOTALWAR | 7276117184 | 21466 | 19-Oct-2018 | 08:12 |  |  | 21-Nov-2018 | 11:18 | NOT WORKING | CHECK AND FOUND MACHINE SHOWING EPROM ISSUE. REPAIRED THE CONTROLLER BOARD. NOW,MACHINE IS WORKING NORMALLY WITH PROPER READING OF SPO2 AND PULSE RATE. | Closed | 4131860 | MEDIUM | 795:6:0 | 34
445 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00197 | PULSE OXIMETER LINE POWERED | SONALI | 8657733630 | 43076 | 12-Jun-2020 | 08:45 |  |  | 13-Jun-2020 | 10:45 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS FAULTY. REPAIRED & FIXED THE SPO2 PROBE. CHECKED THE SPO2 (OXYGEN SATURATION) AND PULSE RATE AND READING OBTAINED IS ACCEPTABLE. PULSE OXIMTER IS NOW WORKING. | Closed | 4131860 | MEDIUM | 26:0:0 | 2
446 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00197 | PULSE OXIMETER LINE POWERED | MS GUNJAN | 8421137019 | 4640 | 20-Jun-2017 | 09:35 |  |  | 24-Jun-2017 | 11:16 | NOT WORKING | REPLACE NEW ADULT SPO2 CABLE | Closed | 4131860 | MEDIUM | 97:41:0 | 5
447 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00197 | PULSE OXIMETER LINE POWERED | TOTALWAR NURSE | 7276114184 | 26931 | 20-Apr-2019 | 09:04 |  |  | 26-Apr-2019 | 10:01 | NOT WORKING | CHECK AND FOUND MACHINE POWER SUPPLY IS CUT OUT . SOLDER THE POWER SUPPLY LINE. NOW MACHINE IS WORKING GOOD | Closed | 4131860 | MEDIUM | 144:57:0 | 7
448 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00197 | PULSE OXIMETER LINE POWERED | SONALI | 8208297987 | 43801 | 03-Jul-2020 | 08:44 |  |  | 03-Jul-2020 | 14:30 | NOT WORKING | CHECK AND FOUND MACHINE AUTOMATICALLY SHUTS DOWN AS SOON AS IT TURNED ON. FIXED THE LOOSE CONNECTION BETWEEN INTERFACING BOARD, DISPLAY CONTROLLER BOARD AND POWER SUPPLY BOARD. PROBLEM IS RESOLVED AND PULSE OXIMETER LINE POWERED IS NOW WORKING. | Closed | 4131860 | MEDIUM | 5:46:0 | 1
449 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00197 | PULSE OXIMETER LINE POWERED | NAGDEVE | 9403379340 | 29448 | 04-Jul-2019 | 09:11 |  |  | 06-Jul-2019 | 10:23 | NOT WORKING | CHECK AND FOUND MACHINE MAKES VIBRATIONAL SOUND WHILE TURNING 'ON' AND DISPLAY COULD NOT BE SEEN. EPROM ISSUE IS BEING FOUND OUT. FIXED THE NEW EPROM. NOW, MACHINE IS WORKING WITHOUT ANY ERROR. | Closed | 4131860 | MEDIUM | 49:12:0 | 3
450 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00197 | PULSE OXIMETER LINE POWERED | TOTALWAR | 7276114184 | 31813 | 26-Aug-2019 | 08:45 |  |  | 27-Aug-2019 | 11:45 | NOT WORKING | CHECK AND FOUND SPO2 READING IS NOT COMING. PINS OF THE SPO2 PROBE IS BEND AND NOT PROPERLY INSERTED. SPO2 PROBE PINS GOT FIXED AND REPAIRED AND PROBE IS INSERTED PROPERLY. NOW, READING IS OBTAINED. PULSE OXIMETER IS NOW WORKING. | Closed | 4131860 | MEDIUM | 27:0:0 | 2
451 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00197 | PULSE OXIMETER LINE POWERED | GUNJAN PALI | 9226201417 | 11429 | 09-Jan-2018 | 09:00 |  |  | 10-Jan-2018 | 23:55 | NOT WORKING | RECONNECT THE ALL CONNECTION | Closed | 4131860 | MEDIUM | 38:55:0 | 2
452 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00198 | NEBULIZER ELECTRIC | GUNJAN PALI | 9226201417 | 12835 | 13-Feb-2018 | 10:20 |  |  | 14-Feb-2018 | 21:19 | NOT WORKING | RESOLDER THE POWER CARD | Closed | NO INFO | MEDIUM | 34:59:0 | 2
453 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00198 | NEBULIZER ELECTRIC | SAMNGITA SISTER | 7276114184 | 17617 | 07-Jul-2018 | 08:35 |  |  | 07-Jul-2018 | 17:30 | NOT WORKING | CHECK AND FOUND MAIN SUPPLY WIRE REMOVE DUE TO STRETCH. RECONNECT DONE. NOW NEBULIZER WORKING SATISFACTORY. | Closed | NO INFO | MEDIUM | 8:55:0 | 1
454 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00198 | NEBULIZER ELECTRIC | SANGITA NURSE | 7276114184 | 30875 | 06-Aug-2019 | 08:02 |  |  | 06-Aug-2019 | 11:10 | DUE TO SHORT CIRCIT MACIN IS NOT WORKING | CHECK AND FOUND WIRE ACROSS ON/OFF SWITCH IS BURNT DUE TO SHORT CIRCUIT. FIXED THE WIRE AND REPAIRED IT. NEBULIZER MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 3:8:0 | 1
455 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00198 | NEBULIZER ELECTRIC | NAGDEVE NURSE | 9403379340 | 12018 | 20-Jan-2018 | 09:16 |  |  | 20-Jan-2018 | 18:11 | NOT WORKING | SOLDERING THE WIRE CONNECTION | Closed | NO INFO | MEDIUM | 8:55:0 | 1
456 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00199 | ECG UNIT 3 CHANNEL PORTABLE | SANDIP AWSARMOL | 9021099179 | 4458 | 13-Jun-2017 | 09:22 |  |  | 15-Jun-2017 | 14:22 | NOT WORKING | REPLACED 10 LEAD ECG CABLE | Closed | DETE6E3718 | MEDIUM | 53:0:0 | 3
457 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00199 | ECG UNIT 3 CHANNEL PORTABLE | GUNJAN PATIL | 9226201417 | 5730 | 29-Jul-2017 | 08:15 | 31-Jul-2017 | 11:20 | 21-Sep-2017 | 17:38 | NOT WORKING | REPLACED THE PRINTER CAP AND ROLLER | Closed | DETE6E3718 | MEDIUM | 1305:23:0 | 55
458 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00199 | ECG UNIT 3 CHANNEL PORTABLE | AJIT KUZEKAR | 8830242857 | 12984 | 16-Feb-2018 | 10:00 |  |  | 16-Feb-2018 | 21:47 | NOT WORKING | RESET THE SETTINGS | Closed | DETE6E3718 | MEDIUM | 11:47:0 | 1
459 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00199 | ECG UNIT 3 CHANNEL PORTABLE | SONALI | 8657733630 | 42642 | 26-May-2020 | 09:10 |  |  | 26-May-2020 | 15:30 | NOT WORKING | PROBLEM REPORTED OF CHEST LEAD 5 (V5) GRAPH OBTAINED IS BLANK. REPAIRED THE FEMALE ECG CABLE CONNECTOR PCB BOARD. NOW, ALL THE LIMB LEADS GRAPH IS OBTAINED. MACHINE IS NOW WORKING. | Closed | DETE6E3718 | MEDIUM | 6:20:0 | 1
460 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00199 | ECG UNIT 3 CHANNEL PORTABLE | SANGITA TOTALWAR | 7276114184 | 10392 | 12-Dec-2017 | 09:50 |  |  | 12-Dec-2017 | 18:25 | NOT WORKING | REPLACED THE 12 LEAD CABLE | Closed | DETE6E3718 | MEDIUM | 8:35:0 | 1
461 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00199 | ECG UNIT 3 CHANNEL PORTABLE | SANGITA | 7276114184 | 39396 | 12-Feb-2020 | 08:50 |  |  | 12-Feb-2020 | 10:00 | NOT WORKING | CHECK AND FOUND FEMALE ECG CABLE CONNECTOR IS FAULTY AND DUE TO WHICH GRAPH OBTAINED IS SIMPLY DASH LINE. FIXED THE NEW FEMALE ECG CABLE CONNECTOR. NOW, SIGNALS IS BEING OBTAINED THROUGH ECG LEAD CABLE. GRAGH OBTAINED IS SATISFACTORY. ECG MACHINE IS NOW WORKING. | Closed | DETE6E3718 | MEDIUM | 1:10:0 | 1
462 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00199 | ECG UNIT 3 CHANNEL PORTABLE | SANGITA DEULKAR | 72761741182 | 16050 | 23-May-2018 | 09:16 |  |  | 25-May-2018 | 16:33 | NOT WORKING | CLEANED THE ELECTRODES AND REPLACED PRINTER PAD. NOW THE MACHINE IS IN WORKING CONDITION | Closed | DETE6E3718 | MEDIUM | 55:17:0 | 3
463 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00199 | ECG UNIT 3 CHANNEL PORTABLE | AJITH KUZEKAR | 8830242875 | 13062 | 20-Feb-2018 | 08:40 |  |  | 20-Feb-2018 | 15:35 | NOT WORKING | RESOLDER THE LEAD CABLE | Closed | DETE6E3718 | MEDIUM | 6:55:0 | 1
464 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00199 | ECG UNIT 3 CHANNEL PORTABLE | SANGITA SISTER | 7276114184 | 9229 | 16-Nov-2017 | 11:15 |  |  | 17-Nov-2017 | 19:53 | NOT WORKING | PASTED THE BROKEN PART THEN WORKING PROPERLY | Closed | DETE6E3718 | MEDIUM | 32:38:0 | 2
465 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00200 | SUCTION PUMP PORTABLE, ELECTRONIC | SEEMA NURSE | 8055416152 | 29449 | 04-Jul-2019 | 09:12 |  |  | 06-Jul-2019 | 14:07 | FILTER ISSUE | CHECK AND FOUND INLET FILTER IS BROKEN.CHANGED THE OLD DEFECTIVE INLET FILTER WITH THE NEW ONE. SUCTION PUMP IS NOW WORKING. | Closed | DETE6J4401 | MEDIUM | 52:55:0 | 3
466 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00200 | SUCTION PUMP PORTABLE, ELECTRONIC | SHENDE NURSE | 9823149772 | 12698 | 08-Feb-2018 | 09:23 |  |  | 12-Feb-2018 | 09:25 | NOT WORKING | REWIND THE MOTOR AND DROPPED THE 2T OIL ALSO | Closed | DETE6J4401 | MEDIUM | 96:2:0 | 5
467 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00204 | BP APPARATUS ANEROID | SEEMA KURAI | 8055416152 | 30912 | 06-Aug-2019 | 08:39 |  |  | 06-Aug-2019 | 10:17 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS LEAK. CHANGED THE NEW BLADDER,VALVE AND BULB. B. P APPRATUS IS NOW WORKING. | Closed | RH010494 | LOW | 1:38:0 | 1
468 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00204 | BP APPARATUS ANEROID | SEEMA NURSE | 8055416552 | 26976 | 22-Apr-2019 | 09:02 |  |  | 27-Apr-2019 | 11:11 | NOT WORKING | CHECK AND FOUND BLADDER AND VALVE IS LEAK. CHANGED THE OLD DEFECTIVE PART WITH NEW. B. P APPARATUS IS NOW WORKING. | Closed | RH010494 | LOW | 122:9:0 | 6
469 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00205 | LOW COST GLUCOMETER | GAYATRI PATALE | 7350412543 | 3202 | 19-Apr-2017 | 10:49 |  |  | 20-Apr-2017 | 13:41 | NOT WORKING | REPLACED THE 3V LITHIUM CELL | Closed | NO INFO | MEDIUM | 26:52:0 | 2
470 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00206 | ECG UNIT 3 CHANNEL PORTABLE | Seema Kurai | 8055416152 | 43369 | 23-Jun-2020 | 09:05 |  |  | 24-Jun-2020 | 10:30 | NOT WORKING | PROBLEM REPORTED THAT GRAPH OF ALL LEADS IS NOT COMING. SETTING OF MODE IS BEING DONE. GRAPH PRINTING SLOT IS CLEANED AND REFIXED. NOW, ECG GRAPH IS OBTAINED OF ALL LEADS SYSTEM. HANDS-ON-TRAINING IS GIVEN. ECG MACHINE IS NOW WORKING. | Closed | DETE6J4401 | MEDIUM | 25:25:0 | 2
471 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00209 | ECG UNIT 3 CHANNEL PORTABLE | D G SATHAWANE | 9403243102 | 38264 | 22-Jan-2020 | 09:18 |  |  | 22-Jan-2020 | 15:42 | NOT WORKING | CHECK AND FOUND FEMALE ECG CABLE CONNECTOR PCB BOARD IS NOT WORKING AND DUE TO WHICH SIGNAL COULD NOT BE OBTAINED IN THE FORM OF GRAPH. GRAPH OBTAINED IS SIMPLY STRAIGHT LINE. FIXED THE NEW FEMALE ECG CABLE CONNECTOR PCB BOARD,NEW ECG LEAD CABLE AND ECG CHEST BULB. ECG IS BEING OBTAINED OF PATIENT AND GRAPH OBTAINED IS SATISFACTORY. ECG MACHINE IS NOW WORKING. | Closed | DETE6H4438 | MEDIUM | 6:24:0 | 1
472 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00209 | ECG UNIT 3 CHANNEL PORTABLE | PRIYA SAHANE | 9325073518 | 29863 | 12-Jul-2019 | 08:50 |  |  | 15-Jul-2019 | 15:02 | NOT WORKING | CHECK AND FOUND ECG CABLE OF RIGHT ARM IS BROKEN. FIXED AND REPAIRED THE CABLE. TOOK THE ECG AND FOUND READING IS SATISFACTORY OK. E.C.G MACHINE IS NOW WORKING. | Closed | DETE6H4438 | MEDIUM | 78:12:0 | 4
473 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00209 | ECG UNIT 3 CHANNEL PORTABLE | T B SHILLEDAR | 8208684604 | 33330 | 03-Oct-2019 | 09:02 |  |  | 05-Oct-2019 | 14:44 | NOT WORKING | CHECK AND FOUND ECG GRAPH OBTAINED IS SIMPLY DASH LINE. CROSS-CHECKED THE ECG LEAD CABLE ON OTHER BPL ECG MACHINE AND FOUND ECG CABLE IS DEFECTIVE. CHANGED THE DEFECTIVE ECG CABLE WITH THE NEW ONE. ECG IS BEING TAKEN OF PATIENT AND GRAPH OBTAINED IS SATISFACTORY. ECG MACHINE IS NOW WORKING. | Closed | DETE6H4438 | MEDIUM | 53:42:0 | 3
474 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00212 | SUCTION PUMP PORTABLE, ELECTRONIC | SEEMA MATE | 7769811631 | 35779 | 26-Nov-2019 | 09:12 |  |  | 26-Nov-2019 | 14:57 | NOT WORKING | CHECK AND FOUND SUCTION PRESSURE COULD NOT BE GENERATED AS HYDRAULIC VACUUM PUMP IS FAULTY. ENTIRE NEW SHEETS OF VACUUM PUMP IS FIXED THROUGH VENDOR. NOW, DESIRED SUCTION PRESSURE IS GENERATED. SUCTION PRESSURE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 5:45:0 | 1
475 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00212 | SUCTION PUMP PORTABLE, ELECTRONIC | SCOTT NURSE | 9403243102 | 33766 | 12-Oct-2019 | 09:32 |  |  | 14-Oct-2019 | 15:07 | NOT WORKING | CHECK AND FOUND SUCTION PRESSURE IS NOT GENERATED. PROBLEM FOUND IS OF HYDRAULIC VACUUM PUMP. ENTIRE SHEETS OF VACUUM PUMP IS DEFECTIVE AND NEW SHEETS IS TO BE MADE IN LATHE MAHCHINE.GAVE THE MOTOR TO LOCAL VENDOR FOR REPAIR. AT PRESENT THIS BREAKDOWN CALL IS CLOSED.AS SOON AS REPAIRED MOTOR WILL BE RECEIVED, AGAIN THE CALL WILL BE RAISED AND MACHINE WILL BE MADE TOTALLY FUNCTIONAL. | Closed | NO INFO | MEDIUM | 53:35:0 | 3
476 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00212 | SUCTION PUMP PORTABLE, ELECTRONIC | TRISHNA POTPODE | 9422474142 | 19183 | 14-Aug-2018 | 09:36 |  |  | 18-Aug-2018 | 15:05 | NOT WORKING | CHECK AND FOUND SHAFT OF THE MOTOR IS JAMMED MOTOR IS REPAIRED.OILING IS SUCTION MACHINE IS NOW WORKING | Closed | NO INFO | MEDIUM | 101:29:0 | 5
477 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00216 | ECG UNIT 3 CHANNEL PORTABLE | SHASHIKALA KAMBALE | 9922114086 | 33331 | 03-Oct-2019 | 09:03 |  |  | 05-Oct-2019 | 15:46 | NOT WORKING | CHECK AND FOUND RUBBER BULB AND METAL ELECTRODES OF CHEST LEADS ARE DEGRADED AND WORN OUT. CHANGED THE OLD RUBBER BULB & METAL ELECTRODES WITH THE NEW ONE. E.C.G MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 54:43:0 | 3
478 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00218 | BP APPARATUS ANEROID | PUSHPA SAVARKAR | 8657782281 | 23324 | 24-Dec-2018 | 08:24 |  |  | 28-Dec-2018 | 15:23 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK. CHANGED THE NEW BLADDER. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 102:59:0 | 5
479 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00218 | BP APPARATUS ANEROID | PATALE NURSE | 8007576014 | 33199 | 28-Sep-2019 | 09:20 |  |  | 28-Sep-2019 | 11:00 | NOT WORKING | CHECK AND FOUND BLADER, CUFF AND EXTENSION TUBE IS LEAK. CHANGE THE DEFECTIVE PARTS WITH THE NEW BLADDER ,CUFF AND EXTENSION TUBE, SO B. P. APPRATUS IS NOW WORKING AND HANDOVER TO USER. | Closed | NO INFO | LOW | 1:40:0 | 1
480 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | WALMIK D MEHAL | 9834283591 | 38244 | 22-Jan-2020 | 09:03 |  |  | 24-Jan-2020 | 13:00 | NOT WORKING | CHECKED AND FOUND BUCKY SUPPLY WIRE MALFUNCTIONING. RECTIFIED THE SAME AND NOW MACHINE WORKING PROPERLY. KEPT UNDER OBSERVATION FOR 3 DAYS. | Closed | 12C016 | MEDIUM | 51:57:0 | 3
481 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | KOMAL SINGH RATHOD | 9822565510 | 18459 | 26-Jul-2018 | 08:35 |  |  | 27-Aug-2018 | 17:30 | NOT WORKING | CHECK AND FOUND HT CANTROL BOARD IS DEFECTIVE CHANGED THE NEW HT CONTROL BOARD.X-RAY MACHINE IS NOW WORKING. | Closed | 12C016 | MEDIUM | 776:55:0 | 33
482 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | RAVINDRA FATING | 8055273031 | 15137 | 21-Apr-2018 | 10:02 |  |  | 25-Apr-2018 | 15:00 | NOT WORKING | REPLACED X-RAY HEAD TUBE INTO AN MACHINE,NOW MACHINE IS WORKING CONDITION. | Closed | 12C016 | MEDIUM | 100:58:0 | 5
483 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | RAVINDRA | 8055273031 | 34152 | 20-Oct-2019 | 08:18 |  |  | 21-Oct-2019 | 16:18 | NOT WORKING | CHECK AND FOUND POWER CABLE IS FAULTY AS FUMES AND SPARKS IS GENERATED FROM IT. CHANGED THE NEW POWER CABLE OF RATING 240V~,16A.CHECKED THE MACHINE BY TAKING BLANK X-RAY'S WITH NO ERROR DISPLACED ON SCREEN. MACHINE IS NOW WORKING. | Closed | 12C016 | MEDIUM | 32:0:0 | 2
484 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | MEHAR | 9834283591 | 34620 | 04-Nov-2019 | 09:26 |  |  | 04-Nov-2019 | 15:03 | NOT WORKING | CHECK AND FOUND IMAGE SHIFTING PROBLEM. DONE PROPER CALIBRATION OF COLLIMETER LIGHT AND RADIATION LIGHT. CHECKED MACHINE NOW FOUND OK. HANDOVER TO USER IN PROPER WORKING ORDER. | Closed | 12C016 | MEDIUM | 5:37:0 | 1
485 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | RAVINDRA | 8055273031 | 43090 | 13-Jun-2020 | 08:20 |  |  | 15-Jun-2020 | 14:30 | NOT WORKING | CHECK AND FOUND X-RAY CANNOT BE SHOOT ON BUCKY MODE GIVING ERROR ALARM. INTERNAL CONNECTION THROUGH BUCKY IS RAT BITTEN. NEW WIRING THROUGH BUCKY IS DONE. NOW, X-RAY CAN BE SHOOT ON BUCKY MODE. X-RAY MACHINE IS NOW WORKING. NOTE: USER IS ADVISED ON URGENT ATTENTION TO MONITOR RAT AND USE REPELLENT TO KEEP RAT AWAY. | Closed | 12C016 | MEDIUM | 54:10:0 | 3
486 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | Meher | 9834283591 | 43992 | 09-Jul-2020 | 08:37 |  |  | 11-Jul-2020 | 15:30 | NOT WORKING | CHECK AND FOUND BUCKY WIRING IS BURNT DUE TO SHORT CIRCUIT. SHORT CIRCUIT IS CAUSED DUE TO RAT BITE. REWIRING OF BUCKY CONNECTION IS DONE AND BUCKY IS MADE FUNCTIONAL. X-RAY MACHINE IS NOW WORKING. NOTE: ALREADY INSTRUCTED X-RAY TECHNICIANS TO USE RAT REPELLENT TO KEEP RAT AWAY AND AVOID MAJOR PROBLEM BUT STILL ISSUES CAUSED BY RAT IS THERE. URGENT ACTION TO BE TAKEN BY HOSPITAL STAFF. | Closed | 12C016 | MEDIUM | 54:53:0 | 3
487 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | W D MEHAR | 9423114953 | 10795 | 21-Dec-2017 | 08:30 |  |  | 26-Dec-2017 | 23:22 | NOT WORKING | REPLACED THE ON/OFF SWITCH AND REPAIRED POWER SUPPLY BOARD | Closed | 12C016 | MEDIUM | 134:52:0 | 6
488 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | WALMIK MEHAR | 9423114953 | 17692 | 10-Jul-2018 | 09:10 |  |  | 12-Jul-2018 | 16:50 | NOT WORKING | CHECK AND FOUND MACHINE WAS HANG DURING EXPOSURE. ALSO FOUND MOISTURE ON CONTROL UNIT DUE TO HUMIDITY. DRY THE CONTROL UNIT AND P.C.B. MACHINE TESTED AND HANDOVER IN PROPER WORKING ORDER. | Closed | 12C016 | MEDIUM | 55:40:0 | 3
489 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00222 | XRAY ABOVE 500mA (HF) | NIMBADSE | 9423114953 | 6954 | 12-Sep-2017 | 09:14 |  |  | 13-Sep-2017 | 14:14 | NOT WORKING | COLIMETER SETTING WAS CHANGED AND BUCKE CONNECTION MANUALLY ARRANGED | Closed | 12C016 | MEDIUM | 29:0:0 | 2
490 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00224 | DENTAL CHAIR UNIT | TRUPTI | 7498189655 | 40942 | 16-Mar-2020 | 09:14 |  |  | 16-Mar-2020 | 10:00 | NOT WORKING | CHECK AND FOUND TRANFORMER WIRE CONECTION IS SHORT. CONNECT THE WIRE CONECTION AND NOW MACHINE IS WORKING. | Closed | NO INFO | LOW | 0:46:0 | 1
491 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00227 | BLOOD WEIGHT MONITOR | DR SONAVANE | 9763117108 | 23174 | 19-Dec-2018 | 08:25 |  |  | 25-Dec-2018 | 14:56 | NOT WORKING | CHECK AND FOUND SMPS AND CONTROLLER BOARD ISSUE. NEW SMPS OF RATING 15V,3 AMPERE IS INSERTED. CONTROLLER BOARD IS REPAIRED. NOW, BLOOD WEIGHT MONITOR IS WORKING. | Closed | 14024051 | LOW | 150:31:0 | 7
492 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00227 | BLOOD WEIGHT MONITOR | TIJARE SITER | 9970876734 | 32510 | 16-Sep-2019 | 08:02 |  |  | 17-Sep-2019 | 17:32 | NOT WORKING | CHECK AND FOUND POWER SUPPLY ISSUE AS FUSE IS BURNT. FIXED THE NEW FUSE AND POWER SUPPLY ISSUE IS RESOLVED. BLOOD WEIGHT MONITOR IS NOW WORKING. | Closed | 14024051 | LOW | 33:30:0 | 2
493 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00229 | BP APPARATUS ANEROID | KAILASH | 8237933451 | 34843 | 08-Nov-2019 | 08:25 |  |  | 08-Nov-2019 | 15:48 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN BLADDER AND EXTENSION TUBE. CHANGED THE DEFECTIVE PARTS WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | RM100163 | LOW | 7:23:0 | 1
494 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00229 | BP APPARATUS ANEROID | DR APARNA JAKKAL | 8888863689 | 8252 | 24-Oct-2017 | 08:50 |  |  | 24-Oct-2017 | 19:10 | NOT WORKING | REPLACED TUBE AND CLEAN THE FILTER | Closed | RM100163 | LOW | 10:20:0 | 1
495 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00230 | BLOOD BANK REFRIGERATOR 2-8°C | LOKESH | 7184250725 | 32507 | 15-Sep-2019 | 09:00 |  |  | 16-Sep-2019 | 17:26 | NOT WORKING | CHECK AND FOUND PP ERROR IS FLASHED ON SUB-ZERO AND TEMPERATURE IS SHOWN IN MINUS RANGE THOUGH INTERNAL TEMPERATURE OF BBR IS NORMAL. PROBLEM IN TEMPERATURE SENSOR PROBE IS FOUND. FIXED THE BREAKAGE IN WIRE OF SENSOR PROBE. NOW, ERROR IS REMOVED AND TEMPERATURE IS SHOWN AS NORMAL. BBR MACHINE IS NOW WORKING. | Closed | BDIT-1321 | MEDIUM | 32:26:0 | 2
496 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00230 | BLOOD BANK REFRIGERATOR 2-8°C | SAPANA | 94065986554 | 17718 | 10-Jul-2018 | 09:50 |  |  | 17-Jul-2018 | 17:30 | NOT WORKING | CHECK AND FOUND COOLING MOTOR IS FAULTY. REPLACED THE MOTOR WITH FAN . NOW BBR WORKING CONDITION. | Closed | BDIT-1321 | MEDIUM | 175:40:0 | 8
497 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00230 | BLOOD BANK REFRIGERATOR 2-8°C | APARNA JAKKAL | 888863689 | 36004 | 30-Nov-2019 | 09:15 |  |  | 30-Nov-2019 | 19:39 | NOT WORKING | CHECK AND FOUND COMPRESSOR DOESN'T TAKE CUT-OFF AT SET TEMPERATURE. SUB-ZERO CONTROLLER IS FAULTY. FIXED THE NEW SUB-ZERO CONTROLLER. CUT-OFF OF COMPRESSOR TAKES PLACE MAINTAINING THE DESIRED TEMPERATURE BETWEEN 6-8*C. BLOOD BANK REFRIGERATOR IS NOW WORKING. | Closed | BDIT-1321 | MEDIUM | 10:24:0 | 1
498 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00230 | BLOOD BANK REFRIGERATOR 2-8°C | SEEMA TIJARE | 9970876734 | 27584 | 14-May-2019 | 08:25 |  |  | 15-May-2019 | 16:10 | NOT WORKING | CHECK AND FOUND COMPRESSOR DOESN'T TURNS ON AND DUE TO WHICH SYSTEM COULD NOT BE COOLED. COMPRESSOR CONNECTING WIRING IS OPENED DUE TO OVER HEATING. FIXED THE CONNECTING WIRES AND COMPRESSOR TURNED ON. NOW, PROPER COOLING IS MAINTAINED. BBR IS NOW WORKING. | Closed | BDIT-1321 | MEDIUM | 31:45:0 | 2
499 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00230 | BLOOD BANK REFRIGERATOR 2-8°C | DR SONAVANE | 9763117108 | 23175 | 19-Dec-2018 | 08:25 |  |  | 21-Dec-2018 | 11:17 | NOT WORKING | CHECK AND FOUND TEMPERATURE CONTROLLER BOARD IS NOT WORKING. REPAIRED THE INTERNAL WIRING OF TEMPERATURE CONTROLLER BOARD. NOW, IT IS WORKING. BBR IS ALSO WORKING AS TEMPERATURE IS MAINTAINED WITHIN THE DESIRED LIMIT. | Closed | BDIT-1321 | MEDIUM | 50:52:0 | 3
500 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00230 | BLOOD BANK REFRIGERATOR 2-8°C | LAXMIKANT VAIDHYA | 7972507041 | 30377 | 24-Jul-2019 | 08:09 |  |  | 27-Jul-2019 | 16:45 | NOT WORKING | CHECK AND FOUND "COMPRESSOR ERROR" IS FLASHED ON DISPLAY. PROBLEM FOUND IS OF SR RELAY, BATTERY AND 24 V TRANSFORMER WHICH GIVES COMMAND TO COMPRESSOR TO TURN ON/OFF. FIXED THE SUB-ZERO CONTROLLER AND PROBLEM IS RESOLVED. BBR MACHINE IS NOW WORKING. | Closed | BDIT-1321 | MEDIUM | 80:36:0 | 4
501 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00232 | CENTRIFUGE (BENCH TOP) | RAHATE | 9822694627 | 34739 | 06-Nov-2019 | 09:18 |  |  | 07-Nov-2019 | 15:17 | NOT WORKING | CHECK AND FOUND SPEED REGULATOR IS NOT WORKING. CHANGED THE DEFECTIVE SPEED REGULATOR WITH THE NEW ONE. NOW, USER CAN MAINTAIN THE SPEED OF MOTOR AT DESIRED SETTING. CENTRIFUGE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 29:59:0 | 2
502 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00233 | BINOCULAR MICROSCOPE | TIJARE | 9970876734 | 31181 | 13-Aug-2019 | 08:08 |  |  | 14-Aug-2019 | 04:00 | NOT WORKING | BINACULAR MICROSCOPE CHECK AND FOUND HALOGEN BULB IS NOT PROPERLY WORK. BECAUSE BULB IS LOOSE IN SOKET. SO CHANGE THE DEFECTIVE SOKET WITH THE NEW ONE. BINACULAR MICROSCOPE IS NOW WORKING. | Closed | 26233 | MEDIUM | 19:52:0 | 2
503 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00233 | BINOCULAR MICROSCOPE | SEEMA TIJARE | 9970876734 | 33970 | 17-Oct-2019 | 09:00 |  |  | 17-Oct-2019 | 15:59 | NOT WORKING | CHECK AND FOUND HALOGEN LIGHT BULB IS FUSED. FIXED THE NEW HALOGEN BULB OF RATING 6V,20W. BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | 26233 | MEDIUM | 6:59:0 | 1
504 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00233 | BINOCULAR MICROSCOPE | DR APARNA JAKKAL | 8888863689 | 31600 | 21-Aug-2019 | 09:55 |  |  | 22-Aug-2019 | 16:05 | NOT WORKING | CHECK AND FOUND EYEPIECE AND OBJECTIVE LENS(10X/0.25) IS FAULTY AND DUE TO WHICH USER IS FACING PROBLEM WHILE VIEWING THE SLIDES. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | 26233 | MEDIUM | 30:10:0 | 2
505 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00233 | BINOCULAR MICROSCOPE | KAILASH | 8237933451 | 34842 | 08-Nov-2019 | 08:25 |  |  | 08-Nov-2019 | 15:46 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. FIXED THE NEW HALOGEN BULB OF RATING 6V,20W.BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | 26233 | MEDIUM | 7:21:0 | 1
506 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00233 | BINOCULAR MICROSCOPE | SEEMA | 9970876734 | 30249 | 22-Jul-2019 | 08:40 |  |  | 22-Jul-2019 | 15:24 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. CHANGED THE DEFECTIVE HALOGEN BULB WITH THE NEW ONE. BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | 26233 | MEDIUM | 6:44:0 | 1
507 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00235 | SEMI AUTOMATED ELISA READER AND WASHER | PRADIP WASNIK | 9423605761 | 939 | 13-Jan-2017 | 10:02 |  |  | 16-Jan-2017 | 11:18 | NOT WORKING | ADJUST THE LAMP AMPLITUDE AND CLEANING THE LENS | Closed | 2T0620313RBK | MEDIUM | 73:16:0 | 4
508 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00235 | SEMI AUTOMATED ELISA READER AND WASHER | SAPNA | 9405986554 | 38741 | 30-Jan-2020 | 08:42 |  |  | 31-Jan-2020 | 15:38 | NOT WORKING | CHECK AND FOUND ABRUPT CHANGE IN COLOURS ON DISPLAYED SCREEN AND ALPHABET IS UNREADABLE. FIXED THE LOOSE CONNECTION OF CABLES OF PRIMARY BOARD. PROBLEM IS RESOLVED. NOW, DISPLAY IS WORKING AND ALPHABET ALONG WITH READING ARE DISPLAYED CLEARLY. ELISA PLATE ANALYSER IS NOW WORKING. | Closed | 2T0620313RBK | MEDIUM | 30:56:0 | 2
509 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00235 | SEMI AUTOMATED ELISA READER AND WASHER | SAPNA BANSODE | 9405986554 | 19837 | 30-Aug-2018 | 09:38 |  |  | 31-Aug-2018 | 11:50 | NOT WORKING | CHECK AND FOUND PLATE ANALYSER THERMAL PRINTER IS NOT WORKING.FIXED THE PROBLEM.NOW MACHINE WORKING SATISFACTORY AND HANDED OVER TO USER . | Closed | 2T0620313RBK | MEDIUM | 26:12:0 | 2
510 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00238 | SEMI AUTOMATED ELISA READER AND WASHER | SAPNA NURSE | 9405986554 | 40092 | 26-Feb-2020 | 09:05 |  |  | 28-Feb-2020 | 14:43 | NOT WORKING | CHECK AND FOUND INTERNAL CABLE CONNECTOR IS FAULTY AND SUCTION TUBE IS HAVING LEAKAGE. FIXED THE NEW SUCTION TUBE AND INTERNAL CABLE CONNECTOR IS REPAIRED. SEMI AUTOMATED ELISA WASHER IS NOW WORKING. NOTE: MACHINE IS ELISA WASHER, IT IS NOT ELISA READER. | Closed | AW0580313RBK | MEDIUM | 53:38:0 | 3
511 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00238 | SEMI AUTOMATED ELISA READER AND WASHER | SAPNA BANSODE | 9405986554 | 7681 | 29-Sep-2017 | 09:37 |  |  | 30-Sep-2017 | 12:45 | NOT WORKING | CLEAN THE BLOCKAGE TUBES AND CONNECT PROPERLY | Closed | AW0580313RBK | MEDIUM | 27:8:0 | 2
512 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00238 | SEMI AUTOMATED ELISA READER AND WASHER | SAPNA BANSOD | 9405986554 | 5464 | 19-Jul-2017 | 10:05 |  |  | 21-Jul-2017 | 11:35 | NOT WORKING | REMOVE TUBES BLOCKAGE | Closed | AW0580313RBK | MEDIUM | 49:30:0 | 3
513 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00243 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR APARNA JAKKAL | 8888836389 | 34882 | 08-Nov-2019 | 09:27 |  |  | 08-Nov-2019 | 16:52 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT. POWER CORD IS ALSO SHORT. FIXED THE NEW HEATING COIL OF 2KW AND POWER CORD OF RATING 240V~,16A.AUTOCLAVE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 7:25:0 | 1
514 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00245 | Hot Air Oven | LOKESH | 7184250725 | 34705 | 05-Nov-2019 | 09:38 |  |  | 07-Nov-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN POWER CABLE. CHANGE THE NEW POWER CABLE. NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | NO INFO | MEDIUM | 50:22:0 | 3
515 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00246 | BLOOD BANK REFRIGERATOR 2-8°C | TIJARE SITER | 9970876734 | 28331 | 07-Jun-2019 | 08:31 |  |  | 10-Jun-2019 | 16:11 | NOT WORKING | CHECK AND FOUND TEMPERATUTE COULD NOT BE MAINTAINED. PROBLEM FOUND IN FAN MOTOR. REPAIRED THE FAN MOTOR. NOW, FAN MOTOR IS WORKING AND TEMPERATUTE IS MAINTAINED. BLOOD BANK REFRIGERATOR IS NOW WORKING. | Closed | BDIT-1322 | MEDIUM | 79:40:0 | 4
516 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00246 | BLOOD BANK REFRIGERATOR 2-8°C | SAPNA | 9405986554 | 39568 | 15-Feb-2020 | 08:42 |  |  | 19-Feb-2020 | 17:47 | NOT WORKING | CHECK AND FOUND COMPRESSOR IS FAULTY. CONDENSER FAN MOTOR AND EVAPORATOR FAN MOTOR IS ALSO FAULTY. FIXED THE NEW COMPRESSOR AND GAS CHARGING IS BEING DONE. NEW CONDENSER AND EVAPORATOR FAN MOTOR IS FIXED. PROPER COOLING IS BEING DONE. CUT-OFF OF COMPRESSOR IS TAKING PLACE AT SET TEMPERATURE. BBR MACHINE IS NOW WORKING. | Closed | BDIT-1322 | MEDIUM | 105:5:0 | 5
517 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00246 | BLOOD BANK REFRIGERATOR 2-8°C | SEEMA TIZARE | 9970876734 | 21813 | 01-Nov-2018 | 09:09 |  |  | 07-Nov-2018 | 11:28 | NOT WORKING | CHECK AND FOUND COMPRESSOR FAN MOTOR AND TEMPRATURE CONTROLLER IS FAULTY. CHANGE THE NEW COMPRESSOR AND GAS FILLING IS DONE. NEW FAN MOTOR AND TEMPERATURE CONTROLLER BOARD IS CHANGED. NOW TEMPERATURE IS MAINTENED BETWEEN 4-8C.NOW BBR MACHINE IS WORKING | Closed | BDIT-1322 | MEDIUM | 146:19:0 | 7
518 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00246 | BLOOD BANK REFRIGERATOR 2-8°C | DR SONAWANE | 9763117108 | 19609 | 25-Aug-2018 | 08:50 |  |  | 30-Sep-2018 | 11:06 | NOT WORKING | CHECK AND FOUND COMPRESSOR MOTOR CAPACITOR IS FAULTY. CHANGE THE SAME & REPLACE WITH NEW ONE, NOW MACHINE IS WORKING | Closed | BDIT-1322 | MEDIUM | 866:16:0 | 37
519 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00246 | BLOOD BANK REFRIGERATOR 2-8°C | SAPNA BANSOD | 9405986554 | 8120 | 16-Oct-2017 | 09:20 |  |  | 20-Oct-2017 | 23:15 | NOT WORKING | REPLACED TWO POWER SUPPLY RELAYS AND COMPRESSOR CAPACITOR | Closed | BDIT-1322 | MEDIUM | 109:55:0 | 5
520 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00246 | BLOOD BANK REFRIGERATOR 2-8°C | SAPANA NURSE | 9405986554 | 7311 | 20-Sep-2017 | 08:07 |  |  | 21-Sep-2017 | 16:15 | NOT WORKING | DUE_TO HIGH VOLTAGE SOCKET WAS BURNED SO INTERCHANGE THE SOCKET | Closed | BDIT-1322 | MEDIUM | 32:8:0 | 2
521 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00246 | BLOOD BANK REFRIGERATOR 2-8°C | Rahate | 9822694627 | 42889 | 06-Jun-2020 | 09:13 |  |  | 08-Jun-2020 | 15:30 | NOT WORKING | CHECK AND FOUND TEMPERATURE ON SUB-ZERO CONTROLLER BOX IS ABRUPTLY CHANGING FORM LOW TO HIGH AND VICE VERSA. PT 100 SENSOR IS FAULTY. FIXED THE NEW SUB-ZERO CONTROLLER AND THE TEMPERATURE SENSOR. CUT-OFF OF COMPRESSOR IS CHECKED AT SET TEMPERATURE AND ITS OK. PROPER COOLING IS BEING ATTAINED. BLOOD BANK REFRIGERATOR IS NOW WORKING. | Closed | BDIT-1322 | MEDIUM | 54:17:0 | 3
522 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00248 | BLOOD BANK REFRIGERATOR (-40*C) | TIJARE | 9970876734 | 27435 | 09-May-2019 | 08:26 |  |  | 14-May-2019 | 16:43 | NOT WORKING | CHECK AND FOUND FAN MOTOR IS FAULTY. CHANGED THE OLD WITH THE NEW FAN MOTOR. NOW, TEMPERATUTE IS MAINTAINED.BLOOD BANK REFRIGERATOR IS NOW WORKING. | Closed | 7100053 | MEDIUM | 128:17:0 | 6
523 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00248 | BLOOD BANK REFRIGERATOR (-40*C) | DR SONAWANE | 9763117108 | 19804 | 30-Aug-2018 | 09:14 |  |  | 06-Sep-2018 | 11:57 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING | Closed | 7100053 | MEDIUM | 170:43:0 | 8
524 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00253 | ULTRACENTRIFUGE FLOOR MOUNTED | SEEMA SISTER | 9970876734 | 20367 | 14-Sep-2018 | 08:45 |  |  | 19-Sep-2018 | 03:48 | NOT WORKING | UNABLE TO SEE PARAMETERS IN THE DISPLAY BECAUSE OF DEFECTIVE CONTROL PANEL BOARD. WE HAD REPLACE A NEW CONTROL PANEL BOARD AND THEN CHECKED PERFORMENCE OF THE INSTRUMENT.FOUND IT IS WORKING SATISFACTORY | Closed | 336 | MEDIUM | 115:3:0 | 6
525 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00253 | ULTRACENTRIFUGE FLOOR MOUNTED | SEEMA TIJARE | 9970876734 | 25796 | 13-Mar-2019 | 08:05 |  |  | 14-Mar-2019 | 10:10 | NOT WORKING | CHECK AND FOUND LOW VOLTAGE POWER SUPPLY INPUT GIVEN TO THE MACHINE. GIVES THE REQUIRED POWER SUPPLY. NOW MACHINE IS WORKING GOOD | Closed | 336 | MEDIUM | 26:5:0 | 2
526 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00253 | ULTRACENTRIFUGE FLOOR MOUNTED | TIJARE SITER | 9970876734 | 32603 | 17-Sep-2019 | 08:52 |  |  | 19-Sep-2019 | 16:08 | NOT WORKING | CHECK AND FOUND ERROR MESSAGE OF 'OPEN/OEFFNEN' IS DISPLAYED EVEN AFTER DOOR IS CLOSED. PROBLEM IS FOUND IN 4-WAY SWITCH ATTACHED WITH PCB BENEATH THE DOOR. A METALLIC ROLLER ATTACHED WITH DOOR IS RESET AND ALSO THE 4-WAY SWITCH ATTACHED WITH PCB. ERROR IS REMOVED.NOW,ULTRACENTRIFUGE IS WORKING AS NORMAL. | Closed | 336 | MEDIUM | 55:16:0 | 3
527 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00253 | ULTRACENTRIFUGE FLOOR MOUNTED | TIJARE | 9970876734 | 31182 | 13-Aug-2019 | 08:10 |  |  | 14-Aug-2019 | 05:34 | NOT WORKING | CHECK AND FOUND UNTRACENTRIFUGE FLOOR MOUNTED IS NOT WORKING. (MACHINE IS NOT ON). BECAUSE POWER ISSUE. AT PRESENT MACHINE IS WELL AND PROPERLY WORK. | Closed | 336 | MEDIUM | 21:24:0 | 2
528 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00253 | ULTRACENTRIFUGE FLOOR MOUNTED | TIJARE | 9970876734 | 27885 | 24-May-2019 | 08:30 |  |  | 28-May-2019 | 15:13 | NOT WORKING | CHECK ANF FOUND ERROR 50 IS DISPLAYED ON SCREEN. THIS ERROR ARISES BECAUSE ROOM TEMPERATURE IS VERY HIGH.FOR MACHINE TO RUN NORMALLY, ROOM TEMPERATURE SHOULD BE 22*C. INSTRUCTED THE INCHARGE TO MAINTAIN TEMPERATURE BY INSTALLING AC. | Closed | 336 | MEDIUM | 102:43:0 | 5
529 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00253 | ULTRACENTRIFUGE FLOOR MOUNTED | SEEMA | 9970876734 | 30248 | 22-Jul-2019 | 08:40 |  |  | 23-Jul-2019 | 12:18 | NOT WORKING | CHECK AND FOUND "MAINS INTERRUPT ERROR" IS DISPLAYED ON SCREEN. FIXED THE ERROR. ULTRA CENTRIFUGE MACHINE IS NOW WORKING. | Closed | 336 | MEDIUM | 27:38:0 | 2
530 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00253 | ULTRACENTRIFUGE FLOOR MOUNTED | SAPNA NURSE | 9405986554 | 18237 | 20-Jul-2018 | 10:30 |  |  | 26-Jul-2018 | 17:30 | NOT WORKING | CHECK AND FOUND ERROR 50 GIVEN BY MACHINE DUE TO VOLTAGE PROBLEM. DID RESET THE MACHINE, PROBLEM SOLVED. ALSO RECOMMEND TO DO THE SERVICE OF I STEBLIZER. | Closed | 336 | MEDIUM | 151:0:0 | 7
531 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00253 | ULTRACENTRIFUGE FLOOR MOUNTED | SEEMA TIJARE | 9970876734 | 36111 | 04-Dec-2019 | 08:35 |  |  | 05-Dec-2019 | 16:16 | NOT WORKING | CHECK AND FOUND ERROR 50 IS DISPLAYED ON SCREEN JUST 6-7 MINUTE AFTER STARTING THE MACHINE AND MACHINE STOPS ROTATING. UNBALANCE ERROR IS ALSO THERE AND CREATES LOUD SOUND. ERROR 50 PROBLEM RESOLVED. UNBALANCE ERROR IS ALSO RESOLVED.ONE OF BLOOD BAG BUCKET IS BROKEN AND IT COMES UNDER CONSUMABLE ITEMS SO DIRECTED THE PATHOLOGY INCHARGE TO BUY. ALSO DIRECTED THE TECHNICIAN AND INCHARGE NOT TO USE THE MACHINE UNTIL NEW BLOOD BAG BUCKET IS FIXED.CENTRIFUGE MACHINE IS WORKING. | Closed | 336 | MEDIUM | 31:41:0 | 2
532 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MULMULE | 9422144593 | 38919 | 03-Feb-2020 | 08:37 |  |  | 04-Feb-2020 | 11:55 | NOT WORKING | CHECK AND FOUND INPUT RANGE COULD NOT BE ENTERED FOR ANY TEST. THERE IS ISSUE OF PRIMARY BOARD. AS THIS MODEL IS OBSOLETE, SPARE PARTS OF THIS MACHINE IS NOT AVAILABLE NOW. A NEWER MODEL OF THIS MACHINE IS AVAILABLE WITH BETTER FEATURES,EFFICIENCY AND CAPACITY. SO, THIS MACHINE IS PROPOSED FOR CONDEMNATION. | Closed | 108ESOH02520 | MEDIUM | 27:18:0 | 2
533 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MULMULE | 9422144593 | 23893 | 14-Jan-2019 | 08:02 |  |  | 17-Jan-2019 | 02:05 | NOT WORKING | CHECK AND FOUND MACHINE LAMP IS FUSE BLOWN UP. CHANGE THE SAME. NOW MACHINE IS WORKING GOOD | Closed | 108ESOH02520 | MEDIUM | 66:3:0 | 4
534 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MULMULE | 9422144592 | 33959 | 16-Oct-2019 | 09:26 |  |  | 17-Oct-2019 | 16:13 | NOT WORKING | CHECK AND FOUND SAMPLE COULD NOT ASPIRATE AS THERE IS BLOCKAGE IN TUBINGS THROUGH FLOW CELL. CLEANED THE TUBINGS THROUGH FLOW CELL AND GENERAL CLEANING PROCESS IS DONE. NOW, SAMPLES ARE ASPIRATED AND READINGS OBTAINED ARE SATISFACTORY. SEMI AUTOMATED BIOCHEMISTRY ANALYSER IS NOW WORKING. | Closed | 108ESOH02520 | MEDIUM | 30:47:0 | 2
535 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | KHADE | 9561393453 | 27585 | 14-May-2019 | 08:30 |  |  | 18-May-2019 | 15:50 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. CHANGED THE NEW HALOGEN BULB OF RATING 6V,10W.SEMI AUTOMATED BIOCHEMISTRY ANALYSER IS NOW WORKING. | Closed | 108ESOH02520 | MEDIUM | 103:20:0 | 5
536 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MS KHADE | 8983299433 | 2742 | 24-Mar-2017 | 08:15 |  |  | 27-Mar-2017 | 12:48 | NOT WORKING | ADJUST THE LAMP AMPLITUDE , CLEANING AND WASHING | Closed | 108ESOH02520 | MEDIUM | 76:33:0 | 4
537 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | ARUN MULMULE | 9422144593 | 29468 | 05-Jul-2019 | 09:02 |  |  | 06-Jul-2019 | 14:24 | NOT WORKING | CHECK AND FOUND FLOW CELL IS BLOCKED. CLEANED THE FLOW CELL AND MAINTENANCE IS DONE. SEMI-BIOCHEMISTRY ANALYZER IS NOW WORKING. | Closed | 108ESOH02520 | MEDIUM | 29:22:0 | 2
538 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MULMULE | 9422144593 | 37938 | 14-Jan-2020 | 08:32 |  |  | 16-Jan-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS PROGRAM IS NOT SAVED. THEN CALLIBRATE THE MACHINE . NOW MACHINE IS IN WORKING CONDITION. | Closed | 108ESOH02520 | MEDIUM | 51:28:0 | 3
539 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MULMULE | 9422144593 | 25842 | 14-Mar-2019 | 08:05 |  |  | 16-Mar-2019 | 12:05 | NOT WORKING | CHECK AND FOUND IN MACHINE HOLLOGEN LAMP IS BLOWN OUT. NEED TO CHANGE. NOW MACHINE IS WORKING GOOD | Closed | 108ESOH02520 | MEDIUM | 52:0:0 | 3
540 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MULMULE | 9422144593 | 33709 | 11-Oct-2019 | 09:00 |  |  | 11-Oct-2019 | 16:25 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED AND CHECK LIGHT PATH ERROR IS DISPLAYED ON SCREEN. FIXED THE NEW HALOGEN BULB OF RATING 6V,10W.PROPER CLEANING OF MACHINE IS DONE. BIOCHEMISTRY ANALYZER IS NOW WORKING. | Closed | 108ESOH02520 | MEDIUM | 7:25:0 | 1
541 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MULMULE | 9422144593 | 22179 | 17-Nov-2018 | 09:00 |  |  | 23-Nov-2018 | 11:18 | NOT WORKING | CHECK AND FOUND LAMP OFF. CHANGED THE HALOGEN LAMP THEN CHECK THE INSTRUMENT AND RUN THE SAMPLES READING ARE ACCEPTABLE NOW MACHINE IS WORKING FINE | Closed | 108ESOH02520 | MEDIUM | 146:18:0 | 7
542 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MRS KHADE | 8983299435 | 4238 | 03-Jun-2017 | 09:25 |  |  | 09-Jun-2017 | 03:55 | NOT WORKING | REPLACED 6V 10 WATTS BULB AND FEED NEW PROGRAM FOR UREA,NOW ALL ERRORS GET RECTIFIED AND MACHINE IS WORKING CONDITION | Closed | 108ESOH02520 | MEDIUM | 138:30:0 | 7
543 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00262 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | RAHUL SIR | 9021601692 | 10235 | 08-Dec-2017 | 09:50 |  |  | 09-Dec-2017 | 08:30 | NOT WORKING | REPLACED NEW 6V 10 WATTS BULB INTO AN MACHINE,NOW MACHINE IS WORKING CONDITION. | Closed | 108ESOH02520 | MEDIUM | 22:40:0 | 2
544 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00265 | ELECTROPHORESIS MACHINE | RAHUL | 9021601692 | 6691 | 06-Sep-2017 | 10:30 |  |  | 07-Sep-2017 | 15:51 | NOT WORKING | REPLACED NEW 7 AMPS FUSE,NOW MACHINE IS WORKING CONDITION. | Closed | LA690/164/VD0029 | MEDIUM | 29:21:0 | 2
545 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00267 | BINOCULAR MICROSCOPE | LANJEWAR | 7709231577 | 7105 | 14-Sep-2017 | 09:00 |  |  | 16-Sep-2017 | 14:24 | NOT WORKING | REPLACE 6V/20W BULP | Closed | NO INFO | MEDIUM | 53:24:0 | 3
546 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00267 | BINOCULAR MICROSCOPE | ARUN MULMULE | 9422144593 | 40775 | 12-Mar-2020 | 08:25 |  |  | 14-Mar-2020 | 15:05 | NOT WORKING | CHECK AND FOUND SLIDE COULD NOT BE VIEWED CLEARLY AS EYEPIECE IS DUSTY. CLEANED THE EYEPIECE AND LIGHT INTENSITY REGULATOR IS BEING REPAIRED.NOW, SLIDE VIEWED IS VERY CLEAR. BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 54:40:0 | 3
547 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00274 | CENTRIFUGE (BENCH TOP) | LANJEAR | 7709231577 | 34797 | 07-Nov-2019 | 09:03 |  |  | 08-Nov-2019 | 11:37 | NOT WORKING | CHECK AND FOUND CARBON BRUSH IS DEGRADED. FIXED THE NEW CARBON BRUSH. CLEANED THE ARMATURE OF MOTOR. CENTRIFUGE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 26:34:0 | 2
548 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00274 | CENTRIFUGE (BENCH TOP) | RAHUL MATALE | 9021601692 | 16268 | 30-May-2018 | 08:55 |  |  | 31-May-2018 | 02:03 | NOT WORKING | POLISHED THE CARBON BRUSH IN ORDER TO CONTACT WITH ROTOR. NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO | MEDIUM | 17:8:0 | 2
549 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00274 | CENTRIFUGE (BENCH TOP) | RAHUL MATALE | 9021601692 | 22079 | 15-Nov-2018 | 08:08 |  |  | 19-Nov-2018 | 17:18 | NOT WORKING | CHECK AND FOUND CARBON BRUSH IS DEGRADED.CHANGED THE NEW CARBON BRUSH. CENTRIFUGE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 105:10:0 | 5
550 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00274 | CENTRIFUGE (BENCH TOP) | RAHUL MOTALE | 9021601692 | 23024 | 15-Dec-2018 | 08:25 |  |  | 25-Dec-2018 | 15:41 | NOT WORKING | CHECK AND FOUND CARBON BRUSH IS DEGRADED. FIXED THE NEW CARBON BRUSH. NOW, CENTRIFUGE MACHINE IS WORKING. | Closed | NO INFO | MEDIUM | 247:16:0 | 11
551 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00274 | CENTRIFUGE (BENCH TOP) | MULMULE | 9422144593 | 37619 | 07-Jan-2020 | 09:15 |  |  | 07-Jan-2020 | 10:48 | NOT WORKING | CHECK AND FOUND CARBON BRUSH IS DEGRADED. FIXED THE NEW CARBON BRUSH. CLEANED THE ARMATURE OF MOTOR. CENTRIFUGE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 1:33:0 | 1
552 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00274 | CENTRIFUGE (BENCH TOP) | GURUMULA | 9422144593 | 32237 | 09-Sep-2019 | 08:30 |  |  | 10-Sep-2019 | 15:20 | NOT WORKING | CHECK AND FOUND MOTOR IS NOT ROTATING. CARBON BRUSH IS DEGRADED. INSERTED THE NEW CARBON BRUSH. NOW, MOTOR IS ROTATING AT DESIRED RPM.CENTRIFUGE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 30:50:0 | 2
553 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00274 | CENTRIFUGE (BENCH TOP) | RAHUL MATALE | 9021601692 | 16666 | 13-Jun-2018 | 08:48 |  |  | 15-Jun-2018 | 17:52 | NOT WORKING | REPLACED WITH NEW CARBON BRUSH | Closed | NO INFO | MEDIUM | 57:4:0 | 3
554 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00276 | AUTOMATIC TISSUE PROCESSOR | PRADEEP WASNIK | 9423605761 | 836 | 09-Jan-2017 | 12:52 |  |  | 11-Jan-2017 | 11:30 | NOT WORKING | OPERATING METHODOLOGY IS NOT KNOWN BY TECHNICIANS | Closed | 1262 | MEDIUM | 46:38:0 | 3
555 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 29251 | 29-Jun-2019 | 08:15 |  |  | 03-Jul-2019 | 16:46 | NOT WORKING | CHECK AND FOUND HAEMOGLOBIN(HB) READING OBTAINED IS VERY LOW. CHAMBER MAINTENANCE DONE AND CLEANING OF TUBES IS DONE. NOW, READING OBTAINED IS SATISFACTORY. 3-PART HAEMATOLOGY ANALYZER IS NOW WORKING. | Closed | 107ESOH02370 | MEDIUM | 104:31:0 | 5
556 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | ARUN | 9422144593 | 20934 | 01-Oct-2018 | 08:00 |  |  | 06-Oct-2018 | 11:27 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING NOW | Closed | 107ESOH02370 | MEDIUM | 123:27:0 | 6
557 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | KHADE | 8983299433 | 29615 | 08-Jul-2019 | 08:40 |  |  | 09-Jul-2019 | 12:35 | NOT WORKING | CHECK AND FOUND READING OBTAINED IS NOT ACCEPTABLE AS REAGENT IS NOT ORIGINAL. DIRECTED THE TECHNICIAN TO ORDER ORIGINAL HORIBA MANUFACTURED REAGENTS. REST, MACHINE IS WORKING NORMALLY. | Closed | 107ESOH02370 | MEDIUM | 27:55:0 | 2
558 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 27273 | 03-May-2019 | 09:41 |  |  | 06-May-2019 | 08:16 | NOT WORKING | CHECK AND FOUND MACHINE CLAENING IS NOT DONE. DO THE CLEANING AND NOW MACHINE IS WORKING GOOD | Closed | 107ESOH02370 | MEDIUM | 70:35:0 | 4
559 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | RAHUL MATALE | 9021601692 | 20450 | 17-Sep-2018 | 09:11 |  |  | 19-Sep-2018 | 12:01 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING NOW | Closed | 107ESOH02370 | MEDIUM | 50:50:0 | 3
560 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 23649 | 05-Jan-2019 | 09:15 |  |  | 06-Jan-2019 | 14:30 | NOT WORKING | CHECK AND FOUND RBC READING MEASURED IS NOT ACCURATE. CLEANED THE CHAMBER. CALIBRATION IS BEING DONE. NOW, READING MEASURED IS ACCURATE. 3-PART HAEMATOLOGY ANALYZER IS WORKING NOW. ~ | Closed | 107ESOH02370 | MEDIUM | 29:15:0 | 2
561 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 26321 | 28-Mar-2019 | 08:07 |  |  | 29-Mar-2019 | 07:35 | NOT WORKING | CHEAK AND FOUND HB REDING IS NOT CONING. CLEANED THE CHAMBERS BACK FLASH IS GIVEN. NOW READING OBTAINED IS SATISFACTORY. | Closed | 107ESOH02370 | MEDIUM | 23:28:0 | 2
562 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 41000 | 17-Mar-2020 | 09:00 |  |  | 19-Mar-2020 | 16:25 | NOT WORKING | CHECK AND FOUND INTERNAL WIRING SYSTEM & HBC BOARD IS RAT BITTEN. INTERNAL TUBING IS DAMAGED AND CARRIAGE HORIZONTAL BELT IS BROKEN. SO, IT IS PROPOSED FOR RW. NEW HBC BOARD, HB CHAMBER, HORIZONTAL BELT CARRIAGE IS TO BE NEWLY FIXED. AS SOON AS ABOVE PARTS WILL BE RECEIVED, MACHINE WILL BE MADE FUNCTIONAL. | Closed | 107ESOH02370 | MEDIUM | 55:25:0 | 3
563 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | MULMUL | 9422144593 | 20104 | 06-Sep-2018 | 08:07 |  |  | 11-Sep-2018 | 12:16 | NOT WORKING | CHECK AND FOUND PROBLEM IN VALVE NUMBER 5 REPLACE VALVE NO. 5 3-PART HEMOTOLOGY ANALYSER IS WORKING | Closed | 107ESOH02370 | MEDIUM | 124:9:0 | 6
564 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 21602 | 24-Oct-2018 | 08:46 |  |  | 26-Oct-2018 | 12:25 | NOT WORKING | CHECKED AND FOUND SYRINGE MAINTENANCE CLEANING NOT DONE. SYRINGE MAINTENANCE CLEANING DONE. NOW MACHINE IS WORKING FINE. | Closed | 107ESOH02370 | MEDIUM | 51:39:0 | 3
565 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | RAHUL MATALE | 9021601692 | 20588 | 21-Sep-2018 | 08:37 |  |  | 24-Sep-2018 | 11:48 | NOT WORKING | CHECK AND FOUND MACHINE IS NOW WORKING GOOD | Closed | 107ESOH02370 | MEDIUM | 75:11:0 | 4
566 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00280 | 3 PART HAEMATOLOGY ANALYZER | PRADIP WASNIK | 9423605761 | 664 | 03-Jan-2017 | 15:00 | 04-Jan-2017 | 10:55 | 06-Jan-2017 | 13:16 | NOT WORKING | WAITING FOR SPARES PARTS | Closed | 107ESOH02370 | MEDIUM | 70:16:0 | 4
567 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | SEEMA | 9970876734 | 39495 | 14-Feb-2020 | 08:30 |  |  | 16-Feb-2020 | 10:00 | NOT WORKING | CHECK AND FOUND PROBLEM IN VALVE NUMBER 3. REPLACE THE OLD DEFECTIVE VALVE WITH THE NEW VALVE NUMBER 3 NOW,MACHINE IS IN WORKING CONDITION. . | Closed | 801516026E | MEDIUM | 49:30:0 | 3
568 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 42861 | 05-Jun-2020 | 08:20 |  |  | 06-Jun-2020 | 10:45 | NOT WORKING | CHECK AND FOUND WBC, RBC AND HAEMOGLOBIN COUNT MEASURED IS HIGH OF EVERY BLOOD SAMPLE. VALVE NUMBER 1 IS FAULTY AND SAME HAS BEEN CHANGED. EARTHING CONNECTING PORT IS CLEANED AND PROPERLY FIXED. CALIBRATION OF PARAMETERS IS BEING DONE. NOW, EVERY PARAMETERS COUNT MEASURED OF BLOOD SAMPLE ARE ACCURATE AS COMPARED WITH PARAMETERS OBTAINED FROM OTHER NEW MACHINE. 3-PART HAEMATOLOGY ANALYZER IS NOW WORKING. | Closed | 801516026E | MEDIUM | 26:25:0 | 2
569 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | DR JAGGAL | 8888863689 | 1799 | 21-Feb-2017 | 11:01 |  |  | 27-Feb-2017 | 11:54 | NOT WORKING | SOFTWARE CHANGED BY A TRIVITRON COMPANY PERSON | Closed | 801516026E | MEDIUM | 144:53:0 | 7
570 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | SEEMA TIJARE | 9970876734 | 21574 | 23-Oct-2018 | 08:10 |  |  | 26-Oct-2018 | 16:28 | NOT WORKING | CHECK AND FOUND RBC CHAMBER CLEANING IS NOT DONE. REPLACE THE NEW REAGENT AND CLEANED THE CHAMBER. MACHINE IS WORKING GOOD. | Closed | 801516026E | MEDIUM | 80:18:0 | 4
571 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | DR APARNA JAKKAL | 8888863689 | 31599 | 21-Aug-2019 | 09:55 |  |  | 22-Aug-2019 | 15:38 | NOT WORKING | CHECK AND FOUND PROBLEM IN VALVE NUMBER 10.REPLACE THE OLD DEFECTIVE VALVE WITH THE NEW VALVE NUMBER 10.NOW,MACHINE IS IN WORKING CONDITION. NOTE:THERE IS PROBLEM OF EARTHING IN BLOOD BANK DEPARTMENT. REQUESTED THE INCHARGE TO TAKE URGENT ACTION AS SOON AS POSSIBLE. | Closed | 801516026E | MEDIUM | 29:43:0 | 2
572 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | RAHUL | 9021601692 | 6692 | 06-Sep-2017 | 10:45 |  |  | 09-Sep-2017 | 13:29 | NOT WORKING | REPLACED NEW WBC CHAMBER AND 3 WAY VALVE,NOW MACHINE IS WORKING CONDITION | Closed | 801516026E | MEDIUM | 74:44:0 | 4
573 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | MR RAHUL | 9021601692 | 1361 | 02-Feb-2017 | 09:45 |  |  | 04-Feb-2017 | 02:30 | NOT WORKING | SOFTWARE CHANGED | Closed | 801516026E | MEDIUM | 40:45:0 | 3
574 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | RAHUL SIR | 9021601692 | 10236 | 08-Dec-2017 | 09:50 |  |  | 13-Dec-2017 | 11:58 | NOT WORKING | REPLACED THE VALVES AND APERATURE | Closed | 801516026E | MEDIUM | 122:8:0 | 6
575 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | ARUN | 9422144593 | 30506 | 27-Jul-2019 | 11:15 |  |  | 27-Jul-2019 | 17:18 | NOT WORKING | READING OBTAINED IS NOT ACCEPTABLE. NEW SOFTWARE IS INSTALLED AND MAINTENANCE IS BEING DONE. 3-PART HAEMATOLOGY ANALYZER IS NOW WORKING. | Closed | 801516026E | MEDIUM | 6:3:0 | 1
576 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | ARUN MULMULE | 9422144593 | 29469 | 05-Jul-2019 | 09:02 |  |  | 09-Jul-2019 | 13:35 | NOT WORKING | CHECK AND FOUND THERE IS ISSUE OF VALVE NUMBER 1.CHANGED THE OLD WITH NEW VALVE. READING OBTAINED IS SATISFACTORY. 3-PART HAEMATOLOGY ANALYZER IS NOW WORKING. | Closed | 801516026E | MEDIUM | 100:33:0 | 5
577 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | RAHUL MATALE | 9021601692 | 16318 | 01-Jun-2018 | 08:00 |  |  | 07-Jun-2018 | 12:19 | NOT WORKING | DONE THE ROUTINE MAINTENANCE.NOW MACHINE WORKING | Closed | 801516026E | MEDIUM | 148:19:0 | 7
578 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | KHADE SISTER | 8983299433 | 18258 | 21-Jul-2018 | 09:07 |  |  | 26-Jul-2018 | 11:24 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING | Closed | 801516026E | MEDIUM | 122:17:0 | 6
579 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | RAHATEA | 9822694627 | 31222 | 13-Aug-2019 | 09:12 |  |  | 16-Aug-2019 | 02:00 | NOT WORKING | CHECK AND FOUND THERE IS ISSUE OF VALVE NUMBER 10.CHANGED THE OLD WITH NEW VALVE. READING OBTAINED IS SATISFACTORY. 3-PART HAEMATOLOGY ANALYZER IS NOW WORKING. | Closed | 801516026E | MEDIUM | 64:48:0 | 4
580 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | MULMULE A L | 9422144593 | 35026 | 12-Nov-2019 | 08:35 |  |  | 12-Nov-2019 | 16:32 | NOT WORKING | CHECK AND FOUND RT HEMAT SOFTWARE IS CORRUPTED. INSTALLED NEW RT HEMAT SOFTWARE AND CHANGE VALVE NUMBER 3.CHECK THE INSTRUMENT AND MACHINE WORKING NOW. | Closed | 801516026E | MEDIUM | 7:57:0 | 1
581 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 34176 | 22-Oct-2019 | 08:15 |  |  | 23-Oct-2019 | 16:09 | NOT WORKING | CHECK AND FOUND SOFTWARE IS CORRUPT. RE-INSTALLED THE SOFTWARE, BASIC CLEANING PROCESS IS DONE. CHECK THE READING BY TAKING SAMPLE AND THE READING OBTAINED IS SATISFACTORY. 3-PART HAEMATOLOGY ANALYZER IS NOW WORKING. | Closed | 801516026E | MEDIUM | 31:54:0 | 2
582 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | SAPNA | 9405986554 | 38554 | 27-Jan-2020 | 08:55 |  |  | 27-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN VALVE V1 IS NOT WORKING. THEN CHANGE THE NEW VALVE V1 AND DONE CALIBRATION . NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | 801516026E | MEDIUM | 8:5:0 | 1
583 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | DR APARNA JAKKAL | 8888863689 | 37033 | 24-Dec-2019 | 08:22 |  |  | 25-Dec-2019 | 10:15 | NOT WORKING | AFTER CHECKED A MACHINE RBC CHAMBER GET OVER FLOW FOR THAT WE CLEANED 2 VALVE, NOW MACHINE IS WORKING AND READING ALSO SHOWING PROPERLY . | Closed | 801516026E | MEDIUM | 25:53:0 | 2
584 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | KHADE SISTER | 8983299433 | 11782 | 16-Jan-2018 | 08:54 |  |  | 17-Jan-2018 | 11:31 | NOT WORKING | CLEANING PROCESS IS DONE | Closed | 801516026E | MEDIUM | 26:37:0 | 2
585 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | DR JAKKAL | 8888863689 | 5651 | 26-Jul-2017 | 09:10 |  |  | 01-Aug-2017 | 13:18 | NOT WORKING | CALIBRATION IS DONE | Closed | 801516026E | MEDIUM | 148:8:0 | 7
586 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | LANGEWAR NURSE | 7709231577 | 17047 | 22-Jun-2018 | 09:50 |  |  | 22-Jun-2018 | 11:32 | NOT WORKING | REINSERTED THE POWER CHORD PROPERLY | Closed | 801516026E | MEDIUM | 1:42:0 | 1
587 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00281 | 3 PART HAEMATOLOGY ANALYZER | DR JAKKAL | 8888863689 | 5009 | 04-Jul-2017 | 09:35 | 06-Jul-2017 | 10:00 | 08-Jul-2017 | 03:30 | NOT WORKING | AFTER CHECKED A MACHINE RBC CHAMBER GET OVERFLOW FOR THAT WE CLEANED 2 VALVES,NOW MACHINE IS WORKING AND READING ALSO SHOWING PROPERLY. | Closed | 801516026E | MEDIUM | 89:55:0 | 5
588 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00283 | ELECTROLYTE ANALYSER | MR RAHUL | 9021601692 | 1362 | 02-Feb-2017 | 09:45 |  |  | 04-Feb-2017 | 11:43 | NOT WORKING | THIS MACHINE HAS WARRANTY PERIOD | Closed | NO INFO | MEDIUM | 49:58:0 | 3
589 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00283 | ELECTROLYTE ANALYSER | RAHUL MATALE | 9021601692 | 5862 | 03-Aug-2017 | 11:15 |  |  | 07-Aug-2017 | 15:30 | NOT WORKING | REARRANGED ELECTRODES ASSEMBLY PROPERLY AND DONE CLEANING PROCESS.NOW ERROR PROBLEM GET RECTIFIED, MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 100:15:0 | 5
590 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00285 | WATER BATH | MULMULE | 9422144593 | 43787 | 03-Jul-2020 | 08:15 |  |  | 05-Jul-2020 | 10:30 | NOT WORKING | CHECK AND FOUND THERMOSTAT IS BURNT DUE TO WHICH HEATING ELEMENT IS NOT WORKING. FIXED THE NEW THERMOSTAT. NOW, HEATING ELEMENT IS WORKING. WATER IS BOILED AT DESIRED SET TEMPERATURE. WATER BATH IS NOW WORKING. | Closed | NO INFO | LOW | 50:15:0 | 3
591 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00285 | WATER BATH | MULMUL | 9422144593 | 20103 | 06-Sep-2018 | 08:07 |  |  | 11-Sep-2018 | 12:05 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE REPAIRED THE POWER SUPPLY FAILURE ISSUE WATER BATH IS NOW WORKING | Closed | NO INFO | LOW | 123:58:0 | 6
592 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00285 | WATER BATH | MULMULE | 9422144593 | 43333 | 22-Jun-2020 | 08:18 |  |  | 23-Jun-2020 | 10:45 | NOT WORKING | CHECK AND FOUND CONNECTING WIRE ACROSS THERMOSTAT IS BURNT. REWIRING ACROSS THERMOSTAT IS DONE. HEATING COIL IS NOW WORKING. WATER BATH IS NOW WORKING. | Closed | NO INFO | LOW | 26:27:0 | 2
593 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00285 | WATER BATH | KHADE NURSE | 8983299433 | 29917 | 13-Jul-2019 | 08:49 |  |  | 15-Jul-2019 | 12:23 | HEATING ISSUE | CHECK AND FOUND POWER CORD IS DEFECTIVE. REPAIRED THE POWER CORD. POWER SUPPLY PROBLEM IS RESOLVED. WATER BATH IS NOW WORKING. | Closed | NO INFO | LOW | 51:34:0 | 3
594 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00285 | WATER BATH | Mulmule | 9422144593 | 43028 | 11-Jun-2020 | 08:25 |  |  | 11-Jun-2020 | 11:45 | NOT WORKING | CHECK AND FOUND POWER CABLE IS BURNT AND DUE TO WHICH THE MACHINE IS NOT TURNING 'ON'. REPAIRED THE POWER CABLE. POWER SUPPLY FAILURE ISSUE IS RESOLVED. WATER BATH IS NOW WORKING. | Closed | NO INFO | LOW | 3:20:0 | 1
595 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00285 | WATER BATH | GURUMULA | 9422144593 | 32238 | 09-Sep-2019 | 08:32 |  |  | 10-Sep-2019 | 16:23 | NOT WORKING | CHECK AND FOUND INTERNAL CONNECTING WIRE OF THERMOSTAT IS BURNT. FIXED THE NEW WIRE ACROSS THERMOSTAT. NOW, WATER IS BOILED AND HEATER IS WORKING AT DESIRED TEMPERATURE. WATER BATH IS NOW WORKING. | Closed | NO INFO | LOW | 31:51:0 | 2
596 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00285 | WATER BATH | Arun Mulmule | 9422144593 | 42382 | 19-May-2020 | 08:00 |  |  | 20-May-2020 | 11:30 | NOT WORKING | CHECK AND FOUND THERMOSTAT CONNECTING WIRE IS BURNT. REWIRING ACROSS THERMOSTAT IS BEING DONE. POWER SUPPLY FAILURE ACROSS COIL IS RESOLVED. WATER BATH IS NOW WORKING. | Closed | NO INFO | LOW | 27:30:0 | 2
597 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00286 | FULLY AUTOMATED BIOCHEMISTRY ANALYSER | RAHUL MATALE | 9021601692 | 22080 | 15-Nov-2018 | 08:08 |  |  | 19-Nov-2018 | 16:13 | NOT WORKING | NO REAGENT IS AVAILABLE AND ALSO CALIBRATION KIT IS ALSO NOT AVAILABLE. MACHINE IS UNDER CMC. SO, AT PRESENT IT IS NOT IN SCOPE. | Closed | B11203 | MEDIUM | 104:5:0 | 5
598 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00286 | FULLY AUTOMATED BIOCHEMISTRY ANALYSER | PRADEEP WASNIK | 9423605761 | 837 | 09-Jan-2017 | 12:53 |  |  | 10-Jan-2017 | 12:17 | NOT WORKING | RESETTING THE BITTING OPTIONS | Closed | B11203 | MEDIUM | 23:24:0 | 2
599 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00286 | FULLY AUTOMATED BIOCHEMISTRY ANALYSER | RAHUL MATALE | 9021601692 | 16665 | 13-Jun-2018 | 08:46 |  |  | 15-Jun-2018 | 17:11 | NOT WORKING | MACHINE IS NOT FULLY INSTALLED AND REAGENTS ARE NOT AVAILABLE | Closed | B11203 | MEDIUM | 56:25:0 | 3
600 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00286 | FULLY AUTOMATED BIOCHEMISTRY ANALYSER | DR APARNA | 8888863689 | 5241 | 12-Jul-2017 | 10:45 |  |  | 01-Sep-2017 | 14:47 | NOT WORKING | REPLACE THE P.M KIT | Closed | B11203 | MEDIUM | 1228:2:0 | 52
601 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00286 | FULLY AUTOMATED BIOCHEMISTRY ANALYSER | MULMULE | 9422144593 | 43788 | 03-Jul-2020 | 08:15 |  |  |  |  | NOT WORKING |  | WIP | B11203 | MEDIUM
602 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00292 | COLOUR DOPPLER FLOW ULTRASOUND | DR KUPADE | 7350365224 | 6267 | 19-Aug-2017 | 09:27 | 20-Aug-2017 | 11:09 | 22-Aug-2017 | 11:52 | NOT WORKING | RESOLDER AND CLEAN PROPERLY THE CONVEX PROBE | Closed | USS-SAR7E3U/WR | MEDIUM | 74:25:0 | 4
603 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00292 | COLOUR DOPPLER FLOW ULTRASOUND | MR SUBODH | 9423114953 | 8351 | 26-Oct-2017 | 09:17 |  |  | 31-Oct-2017 | 15:24 | NOT WOKING | REINSTALLED THE SOFTWARE AND RESET THE SETTINGS | Closed | USS-SAR7E3U/WR | MEDIUM | 126:7:0 | 6
604 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00294 | BP APPARATUS ANEROID | NANDA DUPARE | 9423113691 | 13005 | 17-Feb-2018 | 08:47 |  |  | 17-Feb-2018 | 15:15 | NOT WORKING | REPLACED THE BULB CONNECTOR | Closed | NO INFO | LOW | 6:28:0 | 1
605 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00294 | BP APPARATUS ANEROID | PRADIP WASNIK | 9423605761 | 2558 | 17-Mar-2017 | 09:46 |  |  | 21-Mar-2017 | 11:01 | NOT WORKING | REFILLING THE MERCURY | Closed | NO INFO | LOW | 97:15:0 | 5
606 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00294 | BP APPARATUS ANEROID | B.V.AYALWAR | 9405949739 | 33498 | 07-Oct-2019 | 08:18 |  |  | 07-Oct-2019 | 13:57 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS HAVING LEAKAGE. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE SPARES WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 5:39:0 | 1
607 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00295 | SUCTION PUMP PORTABLE, ELECTRONIC | DUPARE SISTER | 9423113691 | 8432 | 28-Oct-2017 | 08:20 |  |  | 30-Oct-2017 | 15:54 | NOT WORKING | INFORMED THE TECHNICIAN ASSEMBLE HIGH AMPERE TRIPPING SWITCH | Closed | NO INFO | MEDIUM | 55:34:0 | 3
608 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00295 | SUCTION PUMP PORTABLE, ELECTRONIC | MANDA DUPARE | 9423113691 | 6625 | 04-Sep-2017 | 09:35 | 05-Sep-2017 | 12:29 | 09-Sep-2017 | 14:57 | NOT WORKING | REWINDED THE MOTOR | Closed | NO INFO | MEDIUM | 125:22:0 | 6
609 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00295 | SUCTION PUMP PORTABLE, ELECTRONIC | MANDA DUPARE | 9423113691 | 40708 | 11-Mar-2020 | 08:20 |  |  | 12-Mar-2020 | 11:54 | NOT WORKING | CHECK AND FOUND POWER CABLE IS FAULTY. REPAIRED THE POWER CABLE. POWER SUPPLY FAILURE IS RESOLVED. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 27:34:0 | 2
610 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00296 | SUCTION PUMP PORTABLE, ELECTRONIC | MANDA DUPARE | 9423113691Q | 40709 | 11-Mar-2020 | 08:20 |  |  | 12-Mar-2020 | 11:58 | NOT WORKING | CHECK AND FOUND POWER PLUCK IS BROKEN. FIXED THE NEW POWER PLUCK. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 27:38:0 | 2
611 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00298 | PULSE OXIMETER LINE POWERED | DUPARI | 9421113691 | 32309 | 11-Sep-2019 | 08:13 |  |  | 11-Sep-2019 | 15:09 | NOT WORKING | CHECK AND FOUND POWER SUPPLY ISSUE AS FUSE IS BURNT AND MACHINE IS NOT TURNING 'ON'. FIXED THE NEW FUSE AND POWER SUPPLY ISSUE IS BEING RESOLVED. PULSE OXIMETER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 6:56:0 | 1
612 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00298 | PULSE OXIMETER LINE POWERED | BHAVNA SISTER | 9405949739 | 20362 | 14-Sep-2018 | 08:35 |  |  | 15-Sep-2018 | 11:35 | NOT WORKING | CHECK AND FOUND PULSE OXIMETER POWER SUPPLY CABLE IS NOT FIXED. FIXED THE POWER SUPPLY CABLE.PULSE OXIMETER LINE POWER IS NOW WORKING | Closed | NO INFO | MEDIUM | 27:0:0 | 2
613 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00298 | PULSE OXIMETER LINE POWERED | PORNIMA POTDUKE | 9881409188 | 27989 | 28-May-2019 | 09:25 |  |  | 30-May-2019 | 11:10 | NOT WORKING | CHECK AND FOUND POWER CORD IS FAULTY AND DUE TO WHICH MACHINE IS NOT TURNING ON. CHANGED THE OLD DEFECTIVE CORD WITH THE NEW ONE. PULSE OXIMETER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 49:45:0 | 3
614 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00300 | CARDIOTOCOGRAPHY MACHINE | DUPARE SISTER | 9423113691 | 8431 | 28-Oct-2017 | 08:20 |  |  | 01-Nov-2017 | 09:40 | NOT WORKING | REPLACED NEW ECG 10 LEADS CABLE. NOW MACHINE IS WORKING CONDITION. | Closed | JA3C100360 | MEDIUM | 97:20:0 | 5
615 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00300 | CARDIOTOCOGRAPHY MACHINE | MANDA DHUPRE | 9423113691 | 4829 | 27-Jun-2017 | 10:00 |  |  | 29-Jun-2017 | 14:06 | NOT WORKING | REPLACE THE 10 LEAD ECG CABLE | Closed | JA3C100360 | MEDIUM | 52:6:0 | 3
616 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00303 | BP APPARATUS ANEROID | PORNIMA POTDUKHE | 9881409188 | 14936 | 16-Apr-2018 | 08:25 |  |  | 17-Apr-2018 | 04:10 | NOT WORKING | REPLACED NEW BULB | Closed | NO INFO | LOW | 19:45:0 | 2
617 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00305 | SUCTION PUMP PORTABLE, ELECTRONIC | Mrs. D V Scott | 9764971169 | 42879 | 06-Jun-2020 | 09:02 |  |  | 06-Jun-2020 | 10:45 | NOT WORKING | CHECK AND FOUND WINDING OF THE MOTOR IS BURNT AND VACUUM PUMP IS TOTALLY JAMMED. MOTOR AND VACUUM PUMP IS BEING REPAIRED BY VENDOR. REPAIRED MOTOR FROM VENDOR IS BEING COLLECTED AND FIXED TO MACHINE. CHECKED THE MACHINE AND DESIRED SUCTION PRESSURE IS BEING OBTAINED. SUCTION MACHINE IS NOW WORKING. | Closed | 94216 |  | 1:43:0 | 1
618 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00305 | SUCTION PUMP PORTABLE, ELECTRONIC | SCOTT NURSE | 9764971169 | 36644 | 17-Dec-2019 | 08:26 |  |  | 17-Dec-2019 | 16:19 | NOT WORKING | CHECK AND FOUND MOTOR IS NOT WORKING AS WINDING OF THE MOTOR IS SHORT AND BURNT. ALSO ENTIRE SHEETS OF THE VACUUM PUMP IS DEGRADED AND NEW SHEET IS TO BE MADE IN LATHE MACHINE. GAVE THE MOTOR TO LOCAL VENDOR FOR REPAIR. AT PRESENT, THIS BREAKDOWN COMPLAINT IS BEING CLOSED. AS SOON AS REPAIRED MOTOR WILL BE RECEIVED BY VENDOR, AGAIN THE COMPLAINT WILL BE RAISED AND MOTOR WILL BE FIXED & MACHINE WILL BE MADE FUNCTIONAL. | Closed | 94216 |  | 7:53:0 | 1
619 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00307 | BP APPARATUS ANEROID | DR WADHORE | 9421715677 | 12836 | 13-Feb-2018 | 10:25 |  |  | 14-Feb-2018 | 21:26 | NOT WORKING | REPLACED THE METAL VALVE | Closed | DETE6H4534 |  | 35:1:0 | 2
620 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00307 | BP APPARATUS ANEROID | B.V.AYALWAR | 9405949739 | 33499 | 07-Oct-2019 | 08:19 |  |  | 07-Oct-2019 | 10:54 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN BLADDER, BULB AND VALVE. CUFF IS ALSO WORN OUT. CHANGED THE DEFECTIVE PARTS WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | DETE6H4534 |  | 2:35:0 | 1
621 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00307 | BP APPARATUS ANEROID | MANDA DHUPARE | 9423173691 | 22520 | 30-Nov-2018 | 09:14 |  |  | 01-Dec-2018 | 10:09 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT AND BLADDER BULB AND VALVE LEAKAGE.NO MERCURY IS ALSO PRESENT. CHANGED THE NEW CUFF BLADDER VALVE AND BULB. MERCURY IS ALSO REFILLED B. B. P APPARATUS IS NOW WORKING | Closed | DETE6H4534 |  | 24:55:0 | 2
622 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00307 | BP APPARATUS ANEROID | PORNIMA POTDUKHE | 9881409188 | 14935 | 16-Apr-2018 | 08:25 |  |  | 17-Apr-2018 | 09:27 | NOT WORKING | REPLACED NEW BALDDER | Closed | DETE6H4534 |  | 25:2:0 | 2
623 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00307 | BP APPARATUS ANEROID | NANDA DIPANE | 9423113691 | 17386 | 02-Jul-2018 | 09:35 |  |  | 03-Jul-2018 | 11:30 | NOT WORKIING | FOUND BLADDER IS FAULTY. REPLACED BLADDER, NOW WORKING SATISFACTORY. | Closed | DETE6H4534 |  | 25:55:0 | 2
624 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00308 | BP APPARATUS ANEROID | PORNIMA POTDUKE | 9881409188 | 33395 | 04-Oct-2019 | 09:09 |  |  | 05-Oct-2019 | 15:50 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN BLADDER, BULB AND VALVE. CHANGED THE DEFECTIVE PARTS WITH THE NEW ONE. B.P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 30:41:0 | 2
625 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00308 | BP APPARATUS ANEROID | NANDA DUPARE | 9423113691 | 13006 | 17-Feb-2018 | 08:47 |  |  | 17-Feb-2018 | 15:18 | NOT WORKING | REPLACED THE BULB , BLADDER , TUBING , BULB CONNECTOR | Closed | NO INFO |  | 6:31:0 | 1
626 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00308 | BP APPARATUS ANEROID | MANDA DHUPARE | 9423173691 | 22518 | 30-Nov-2018 | 09:14 |  |  | 01-Dec-2018 | 11:02 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER BULB AND VALVE.CUFF IS WORN OUT. MERCURY IS ALSO NOT PRESENT. CHANGED THE NEW CUFF BLADDER VALVE AND BULB.MERCURY IS ALSO REFILLED B. P APPRATUS IS NOW WORKING | Closed | NO INFO |  | 25:48:0 | 2
627 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00311 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | ARUN MULMULE | 9422144593 | 40777 | 12-Mar-2020 | 08:25 |  |  | 14-Mar-2020 | 15:52 | NOT WORKING | CHECK AND FOUND THE MACHINE IS COAGULATION ANALYZER. IT IS NOT SEMI AUTOMATED BIOCHEMISTRY ANALYZER. FOUND PRINTER ASSEMBLY FAULTY. SENT THE PRINTER ASSEMBLY TO NAGPUR OFFICE FOR REPAIR. AS SOON AS PRINTER ASSEMBLY WILL BE REPAIRED, IT WILL BE FIXED. AT PRESENT THIS BREAKDOWN COMPLAINT IS CLOSED. | Closed | NO INFO |  | 55:27:0 | 3
628 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00312 | ECG UNIT 3 CHANNEL PORTABLE | KOHALE | 9421795126 | 26054 | 19-Mar-2019 | 09:12 |  |  | 22-Mar-2019 | 10:07 | NOT WORKING | CHECK AND FOUND ECG CABLE IS NOT WORKING NEED TO CHANGE THE CABLE . NOW CHANGE THE CABLE WITH NEW ONE NOW MACHINE IS WORKING | Closed | NO INFO |  | 72:55:0 | 4
629 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00313 | AUTO REFRACTOMETER | JITENDRA PADOLE | 8149913123 | 18430 | 25-Jul-2018 | 08:39 |  |  | 01-Aug-2018 | 10:00 | NOT WORKING | CHECK AND FOUND MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 169:21:0 | 8
630 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00317 | OPERATING MICROSCOPE | L S PATIL | 9420866915 | 7359 | 21-Sep-2017 | 09:11 |  |  | 22-Sep-2017 | 13:47 | NOT WORKING | CLEANED OBJECTIVE LENS AND EYE PIECE AND NOW IMAGE IS SHOWING PROPERLY MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 28:36:0 | 2
631 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00320 | ECG UNIT 3 CHANNEL PORTABLE | LAXMI | 8149503046 | 41255 | 24-Mar-2020 | 08:38 |  |  | 24-Mar-2020 | 11:45 | NOT WORKING | CHECK AND FOUND ROLLING OF GRAPH PAPER IS NOT CONTINUOUS AND ALSO GRAPH OBTAINED IS BLANK. CLEANED THE GRAPH PRINTER SLOT. REFIXED THE PRINTER CONNECTION. NOW, GRAPH OBTAINED IS CONTINOUS WITH NO HINDRANCE AND READING ON GRAGH IS ALSO OBTAINED WHICH IS SATISFACTORY AS OBSERVED ON DIFFERENT PATIENT. | Closed | ASE00124 |  | 3:7:0 | 1
632 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00320 | ECG UNIT 3 CHANNEL PORTABLE | SHARDHA VANIKAR | 9623392861 | 33456 | 05-Oct-2019 | 09:16 |  |  | 07-Oct-2019 | 13:23 | NOT WORKING | CHECK ANF FOUND ECG PLASTIC CLAMP IS BROKEN AND RUBBER BULB & METAL ELECTRODE OF CHEST LEAD IS DEGRADED. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. ECG MACHINE IS NOW WORKING. | Closed | ASE00124 |  | 52:7:0 | 3
633 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00320 | ECG UNIT 3 CHANNEL PORTABLE | MAYA | 8149503946 | 26948 | 20-Apr-2019 | 09:16 |  |  | 22-Apr-2019 | 12:13 | NOT WORKING | CHECK AND FOUND CHEST LEAD RUBBER BULB IS DEGRADED. CHANGED THE OLD WITH NEW CHEST LEAD BULB. E.C.G MACHINE IS NOW WORKING. | Closed | ASE00124 |  | 50:57:0 | 3
634 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00320 | ECG UNIT 3 CHANNEL PORTABLE | SONALI BAWANKAR | 9049257037 | 38385 | 23-Jan-2020 | 09:40 |  |  | 23-Jan-2020 | 16:01 | NOT WORKING | CHECK AND FOUND FEMALE ECG CABLE CONNECTOR IS FAULTY AND DUE TO WHICH GRAPH OBTAINED IS SIMPLY DASH LINE. FIXED THE NEW FEMALE ECG CABLE CONNECTOR. NOW, SIGNALS IS BEING OBTAINED THROUGH ECG LEAD CABLE. GRAGH OBTAINED IS SATISFACTORY. ECG MACHINE IS NOW WORKING. | Closed | ASE00124 |  | 6:21:0 | 1
635 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00321 | CT SCAN | NIMBARKHE | 8423114953 | 20949 | 01-Oct-2018 | 08:30 |  |  | 30-Oct-2018 | 11:29 | NOT WORKING | CHECK AND FOUND COMPANY SERVICE ENGINEER REPAIRED THE MACHINE AND HANDOVER TO THE USER IN WORKING CONDITION NOW MACHINE IS WORKING | Closed | 3KD1612677 |  | 698:59:0 | 30
636 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00325 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | MANJUSHA NEVARE | 9423646267 | 14954 | 16-Apr-2018 | 09:15 |  |  | 18-Apr-2018 | 10:11 | NOT WORKING | REPLACED NEW O2 CONNECTOR, NOW MACHINE IS WORKING CONDITION. | Closed | ASED0124 |  | 48:56:0 | 3
637 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00325 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | Mrs D. V Scott | 9764971169 | 42680 | 28-May-2020 | 09:00 |  |  | 28-May-2020 | 11:45 | NOT WORKING | CHECK AND FOUND INFLOW OF OXYGEN COULD NOT BE MEASURED AND CALIBRATION OF OXYGEN FAILS. OXYGEN SENSOR IS FOUND TO BE FAULTY AND ITS DIAPHRAGM PCB IS BURNT. FIXED THE NEW OXYGEN SENSOR AND ITS DIAPHRAGM. NOW, CALIBRATION OF OXYGEN IS SUCCESSFULLY COMPLETED AND INFLOW OF OXYGEN IS MEASURED AND DISPLAYED IN FIO2 COLUMN IN VARIOUS MODES OF VENTILATION. VENTILATOR IS NOW WORKING. | Closed | ASED0124 |  | 2:45:0 | 1
638 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00325 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | MANJUSHA NEVARE | 9423646267 | 13400 | 03-Mar-2018 | 10:28 |  |  | 08-Mar-2018 | 10:49 | NOT WORKING | REPLACED WITH NEW STANDBY COMPRESSOR INTO A MACHINE NOW MACHINE IS WORKING CONDITION. | Closed | ASED0124 |  | 120:21:0 | 6
639 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00325 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | SCOTT NURSE | 9764971169 | 41431 | 03-Apr-2020 | 08:00 |  |  | 30-Apr-2020 | 17:42 | NOT WORKING | CHECK AND FOUND ITS INBUILT COMPRESSOR UNIT IS FAULTY AND WAS TAKEN BY OUR STAFF LONG BACK FOR REPAIR. INBUILT COMPRESSOR UNIT IS BEYOND ECONOMICAL REPAIR. SO, NEW EXTERNAL COMPRESSOR UNIT IS INSTALLED. CHECKED THE VENTILATOR ON DIFFERENT MODES. CALIBRATION OF FLOW AND OXYGEN IS BEING DONE. MACHINE IS WORKING NORMALLY WITHOUT ANY ERROR. HANDED OVER THE VENTILATOR IS GOOD WORKING CONDITION. | Closed | ASED0124 |  | 657:42:0 | 28
640 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00325 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | MANJUSHA NEVARE | 9423646267 | 12857 | 14-Feb-2018 | 08:50 |  |  | 17-Feb-2018 | 15:45 | NOT WORKING | REPLACED THE NEW POWER CABLE AND OXYGEN FLOW CONNECTOR AND RECONNECT THE CIRCUIT | Closed | ASED0124 |  | 78:55:0 | 4
641 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00325 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | SCOTT NURSE | 9764971169 | 29885 | 13-Jul-2019 | 08:06 |  |  | 18-Jul-2019 | 15:02 | NOT WORKING | CHECK AND FOUND ALL PARAMETERS AND MODES ARE WORKING EFFICIENTLY. CHECK AND CLEANED THE OXYGEN SENSOR AND FLOW SENSOR. TIDAL VOLUME AND PEEP PARAMETERS IS CHECKED AND FOUND TO BE OK. INTERNAL AND EXTERNAL BATTERY IS OK. INTENSIVE CARE VENTILATOR IS WORKING EFFICIENTLY. | Closed | ASED0124 |  | 126:56:0 | 6
642 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00326 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | MANJUSHA NEVARE | 9423646267 | 13401 | 03-Mar-2018 | 10:28 |  |  | 05-Mar-2018 | 11:30 | NOT WORKING | RECTIFIED THE BATTERY BACKUP PROBLEM, NOW MACHINE IS WORKING CONDITION. | Closed | AKT03068 |  | 49:2:0 | 3
643 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00326 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | MAJUSHA NIWARE | 9423646267 | 6398 | 23-Aug-2017 | 08:55 | 25-Aug-2017 | 02:45 | 05-Sep-2017 | 11:24 | NOT WORKING | REPAIRED POWER SUPPLY MODULE AND CHECKED,BUT VENTIKITS NOT AVAILABLE WE RECOMMEND USER TO BUY VENTIKITS | Closed | AKT03068 |  | 314:29:0 | 14
644 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00326 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | SCOTT NURSE | 9764971169 | 41430 | 03-Apr-2020 | 08:00 |  |  | 24-Apr-2020 | 16:40 | NOT WORKING | REQUIRED FRONT PANEL ASSEMBLY. PO SENT TO VENDOR, BUT AS PER MAIL CONFIRMATION FROM VENDOR IT WILL TAKE 90 DAYS FOR DELIVERY OF THIS MATERIAL DUE TO LOCKDOWN. LOCKDOWN DUE TO COVID 19- NO SPARES/TRANSPORT/COURIER POSSIBLE-PRESENTLY CLOSING THE WO. WILL OPEN THE CALL AGAIN & REPAIR THE ASSET ONCE THE LOCKDOWN IS OPENED BY THE GOVERNMENT. | Closed | AKT03068 |  | 512:40:0 | 22
645 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00326 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | SCOTT NURSE | 9764971169 | 29886 | 13-Jul-2019 | 08:08 |  |  | 18-Jul-2019 | 16:10 | NOT WORKING | CLEANED THE OXYGEN AND FLOW SENSOR. INSPIRATORY AND EXPIRATORY VOLUME IS CHECKED AND FOUND TO BE OK. TIDAL VOLUME AND PEEP IS MONITORED AND ITS OK. AIR AND OXYGEN FAILURE IS SORTED OUT. INTENSIVE CARE VENTILATOR IS NOW WORKING. | Closed | AKT03068 |  | 128:2:0 | 6
646 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00326 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | MANJUSHA NEVARE | 9423646267 | 14956 | 16-Apr-2018 | 09:15 |  |  | 27-Oct-2018 | 10:16 | NOT WORKING | CHECK AND FOUND MAIN CONTROLLER BOARD IS DEFECTIVE. ORDERED THE SAME. SINCE IT TAKES ALMOST 3 MONTHS TIME FROM COMPANY TO BE HANDED AFTER QUOTATION AND PAYMENT SO AT PRESENT WE ARE CLOSING THIS CALL. AFTER NEW CONTROL BOARD WILL BE RECEIVED BY THE SUPPLIER, AGAIN THE CALL WILL BE RAISED AND IT WILL BE ATTENDED AND FIXED. | Closed | AKT03068 |  | 4657:1:0 | 195
647 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00328 | ECG UNIT 3 CHANNEL PORTABLE | GHARME NURSE | 9511971047 | 10604 | 15-Dec-2017 | 10:25 |  |  | 15-Dec-2017 | 22:12 | NOT WORKING | REPLACED 10 LEAD ECG CABLE | Closed | RM1000146 |  | 11:47:0 | 1
648 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00329 | NEBULIZER ELECTRIC | Rohani | 9420128032 | 41580 | 10-Apr-2020 | 08:30 |  |  | 10-Apr-2020 | 10:00 | NOT WORKING | CHECK AND FOUND MACHINE PUMP IS NOT WORKING SO FILL OILING AND DUST REMOVE. SO PRESSURE IS GONE AND MACHINE IS WORKING. | Closed | 171212 |  | 1:30:0 | 1
649 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00330 | BP APPARATUS ANEROID | PAYAL PANDEY | 8329393206 | 20265 | 11-Sep-2018 | 10:00 |  |  | 14-Sep-2018 | 11:57 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND VALVE .CHANGED THE NEW BLADDER AND VALVE. B.P APPARTUS IS NOW WORKING | Closed | NO INFO |  | 73:57:0 | 4
650 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00331 | BP APPARATUS ANEROID | PAYAL PANDEY | 8329393206 | 20266 | 11-Sep-2018 | 10:00 |  |  | 14-Sep-2018 | 11:13 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BULB AND BLADDER.CHANGED THE NEW BULB AND BLADDER.B.P APPRATUS IS NOW WORKING | Closed | NO INFO |  | 73:13:0 | 4
651 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00332 | BP APPARATUS ANEROID | PAYAL PANDEY | 8329393206 | 20264 | 11-Sep-2018 | 10:00 |  |  | 14-Sep-2018 | 11:08 | NOT WORKING | CHECK AND FOUND LEAKAGE IN VALVE.CHANGED THE NEW VALVE.B.P APPRATUS IS NOW WORKING | Closed | NO INFO |  | 73:8:0 | 4
652 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00333 | BP APPARATUS ANEROID | C S GADDAM | 9850236795 | 34340 | 24-Oct-2019 | 09:05 |  |  | 24-Oct-2019 | 11:07 | NOT WORKING | CHECK AND FOUND BLADDER, BULB, CUFF AND VALVE IS MISSING. FIXED THE NEW BLADDER, BULB, CUFF AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 2:2:0 | 1
653 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00334 | BP APPARATUS ANEROID | TRISHNA POTPHODE | 9422474142 | 18497 | 27-Jul-2018 | 08:31 |  |  | 27-Jul-2018 | 17:00 | NOT WORKING | CHECK AND FOUND CUFF, BLADDER,BULB IS MISSING. NEW CUFF, BLADDER, VALVE,AND BULB REPLACED. NOW MACHINE WORKING SATISFACTORY. | Closed | 98134235 |  | 8:29:0 | 1
654 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00335 | BP APPARATUS ANEROID | PAYAL PANDEY | 8329393206 | 20262 | 11-Sep-2018 | 10:00 |  |  | 14-Sep-2018 | 11:11 | NOT WORKING | CHECK AND FOUND BLADDER ,CUFF BULB AND VALVE IS MISSING .FIXED THE NEW BLADDER CUFF BULB AND VALVE.B.P APPRATUS IS NOW WORKING . | Closed | NO INFO |  | 73:11:0 | 4
655 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00336 | BP APPARATUS ANEROID | PAYAL PANDE | 8329393206 | 12912 | 15-Feb-2018 | 08:50 |  |  | 15-Feb-2018 | 22:08 | NOT WORKING | REPLACED THE BLADDER, BULB | Closed | NO INFO |  | 13:18:0 | 1
656 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00336 | BP APPARATUS ANEROID | KHOBRAGADE SISTER | 9764574066 | 10708 | 19-Dec-2017 | 08:45 |  |  | 19-Dec-2017 | 23:43 | NOT WORKING | REPLACED THE BLADDER, CUFF, BULB, METAL VALVE | Closed | NO INFO |  | 14:58:0 | 1
657 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00337 | BP APPARATUS ANEROID | DR SATELEVE | 7387157815 | 12839 | 13-Feb-2018 | 10:25 |  |  | 14-Feb-2018 | 10:49 | NOT WORKING | REPLACED THE METAL VALVE | Closed | NO INFO |  | 24:24:0 | 2
658 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00338 | BP APPARATUS ANEROID | TRISHNA POTPHODE | 9422474142 | 18375 | 24-Jul-2018 | 09:15 |  |  | 25-Jul-2018 | 12:15 | NOT WORKING | CHECK AND FOUND PROBLEM REPORTED IN CUFF AND MURCURY. REPLACED CUFF AND RE FILL MURCURY. NOW MACHINE WORKING SATISFACTORY | Closed | 98134326 |  | 27:0:0 | 2
659 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00339 | BP APPARATUS ANEROID | PAYAL PANDEY | 8380886773 | 9574 | 23-Nov-2017 | 08:55 |  |  | 24-Nov-2017 | 16:31 |  | REPLACED BLADDER AND COTTON CUFF | Closed | NO INFO |  | 31:36:0 | 2
660 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00340 | BP APPARATUS ANEROID | SANGITA | 7276114184 | 26729 | 12-Apr-2019 | 08:10 |  |  | 15-Apr-2019 | 10:12 | NOT WORKING | CHECK AND FOUND BP CUFF BLADDER IS LEAK NEED TO CHANGE. NOW REPLACE WITH NEW ONE NOW MACHINE IS WORKING GOOD | Closed | 45398 |  | 74:2:0 | 4
661 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00340 | BP APPARATUS ANEROID | SANGITA TOTALWAR | 7276114182 | 32450 | 14-Sep-2019 | 08:08 |  |  | 16-Sep-2019 | 15:52 | NOT WORKING | CHECK AND FOUND VALVE IS LEAK. ALSO CUFF IS WORN OUT. CHANGED THE DEFECTIVE VALVE AND CUFF WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 45398 |  | 55:44:0 | 3
662 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00341 | BP APPARATUS ANEROID | TRISHNA POTPHODE | 9422474142 | 18498 | 27-Jul-2018 | 08:35 |  |  | 27-Jul-2018 | 17:00 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN. REPLACED GLASS TUBE AND FILL MERCURY. NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 8:25:0 | 1
663 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00342 | BP APPARATUS ANEROID | PAYAL PANDEY | 8380886773 | 9572 | 23-Nov-2017 | 08:55 |  |  | 24-Nov-2017 | 13:51 | NOT WORKING | REPLACED BLADDER, CUFF, BULB | Closed | 98134231 |  | 28:56:0 | 2
664 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00343 | BP APPARATUS ANEROID | KHOBRAGADE SISTER | 9764574066 | 10710 | 19-Dec-2017 | 08:45 |  |  | 20-Dec-2017 | 23:48 | NOT WORKING | REPLACED THE BLADDER, CUFF, BULB, METAL VALVE | Closed | NO INFO |  | 39:3:0 | 2
665 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00344 | BP APPARATUS ANEROID | DR SATELEVE | 7387157815 | 12838 | 13-Feb-2018 | 10:25 |  |  | 14-Feb-2018 | 11:39 | NOT WORKING | REPLACED THE METAL VALVE, COTTON CUFF | Closed | 98135355 |  | 25:14:0 | 2
666 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00344 | BP APPARATUS ANEROID | PAYAL PANDEY | 8380886773 | 9575 | 23-Nov-2017 | 08:55 |  |  | 24-Nov-2017 | 16:24 | NOT WORKING | REPLACED BLADDER AND BULB | Closed | 98135355 |  | 31:29:0 | 2
667 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00346 | BP APPARATUS ANEROID | SHILPA LANJEWAR | 8275331823 | 32732 | 19-Sep-2019 | 08:47 |  |  | 19-Sep-2019 | 15:29 | NOT WORKING | CHECK AND FOUND BLADDER, BULB AND VALVE IS HAVING LEAKAGE. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | RF1000550 |  | 6:42:0 | 1
668 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00346 | BP APPARATUS ANEROID | TRISHNA POTPHODE | 9422474142 | 18374 | 24-Jul-2018 | 09:15 |  |  | 25-Jul-2018 | 12:30 | NOT WORKING | CHECK AND FOUND BLADDER LEAK, MURCURY IS EMPTY, VALVE IS NOT WORKING. FILL MURCURY REPLACED BLADDER AND VALVE. NOW MACHINE IS IN WORKING CONDITION. | Closed | RF1000550 |  | 27:15:0 | 2
669 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00348 | BP APPARATUS ANEROID | DR SATELEVE | 7387157815 | 12837 | 13-Feb-2018 | 10:25 |  |  | 14-Feb-2018 | 21:30 | NOT WORKING | REPLACED THE METAL VALVE | Closed | RM1000177 |  | 35:5:0 | 2
670 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00348 | BP APPARATUS ANEROID | KHOBRAGADE SISTER | 9764574066 | 10709 | 19-Dec-2017 | 08:45 |  |  | 20-Dec-2017 | 23:53 | NOT WORKING | REPLACED THE BLADDER, CUFF, BULB, METAL VALVE | Closed | RM1000177 |  | 39:8:0 | 2
671 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00349 | BP APPARATUS ANEROID | TRISHNA POTPHODE | 9422474142 | 18373 | 24-Jul-2018 | 09:15 |  |  | 25-Jul-2018 | 12:30 | NOT WORKING | CHECK AND FOUND CUFF, BLADDER, BULB,GLASS TUBE AND VALVE IS MISSING. INSERT ALL WITH NEW ONE. FILL THE MURCURY. NOW B.P. APRATUS WORKING FINE. | Closed | NO INFO |  | 27:15:0 | 2
672 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00351 | BP APPARATUS ANEROID | PRIYA MESHRAM | 9665175373 | 34363 | 24-Oct-2019 | 09:35 |  |  | 24-Oct-2019 | 11:10 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE OF BLADDER AND VALVE. CHANGED THE OLD DEFECTIVE BLADDER AND VALVE WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 1:35:0 | 1
673 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00353 | BP APPARATUS ANEROID | ANAJEVAR | 8275331823 | 40163 | 27-Feb-2020 | 08:18 |  |  | 27-Feb-2020 | 10:00 | NOT WORKING | CHECK AND FOUND BLADER, CUFF AND VALVE IS LEAK AND GLASS TUBE IS DAMEGE. CHANGE THE BLADER, CUFF AND VALVE AND GLASS TUBE WITH NEW ONE. MACHINE IS WORKING NOW AT PRESENT AND HANDOVER USER IN WORKING CONDITION. | Closed | RM0104071 |  | 1:42:0 | 1
674 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00353 | BP APPARATUS ANEROID | PAYAL PANDEY | 8329393206 | 20261 | 11-Sep-2018 | 10:00 |  |  | 14-Sep-2018 | 11:06 | NOT WORKING | CHECK AND FOUND LEAKAG IN BLADDER . CHANGED NEW BLADDER .B.P APPRATUS IS NOW WORKING | Closed | RM0104071 |  | 73:6:0 | 4
675 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00354 | BP APPARATUS ANEROID | PRIYA MESHRAM | 9665175373 | 34364 | 24-Oct-2019 | 09:35 |  |  | 24-Oct-2019 | 11:11 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN BLADDER, BULB AND VALVE. CHANGED THE DEFECTIVE SPARES WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 1:36:0 | 1
676 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00355 | BP APPARATUS ANEROID | PAYAL PANDEY | 8329393206 | 20263 | 11-Sep-2018 | 10:00 |  |  | 14-Sep-2018 | 11:04 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER .CHANGED THE NEW BLADDER. B.P APPRATUS IS NOW WORKING | Closed | RM0104073 |  | 73:4:0 | 4
677 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00357 | BP APPARATUS ANEROID | MANDA DHUPARE | 9423173691 | 22522 | 30-Nov-2018 | 09:14 |  |  | 01-Dec-2018 | 11:05 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER BULB AND VALVE. CUFF IS ALSO TORN OUT. CHANGED THE NEW CUFF BLADDER VALVE AND BULB MERCURY IS ALSO REFILLED B. P APPARATUS IS NOW WORKING | Closed | NO INFO |  | 25:51:0 | 2
678 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00357 | BP APPARATUS ANEROID | PORNIMA POTDUKHE | 9881409188 | 14937 | 16-Apr-2018 | 08:25 |  |  | 17-Apr-2018 | 03:55 | NOT WORKING | REPLACED NEW BLADDER AND BULB | Closed | NO INFO |  | 19:30:0 | 2
679 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00359 | BP APPARATUS ANEROID | PAYAL PANDE | 8329393206 | 12918 | 15-Feb-2018 | 09:15 |  |  | 15-Feb-2018 | 22:11 | NOT WORKING | REPLACED THE METAL VALVE | Closed | NO INFO |  | 12:56:0 | 1
680 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00364 | BP APPARATUS ANEROID | GUNJAN PALI | 9226201417 | 24565 | 02-Feb-2019 | 08:02 |  |  | 02-Feb-2019 | 16:20 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN. FIXED THE NEW GLASS TUBE AND REFILLED THE MERCURY. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 8:18:0 | 1
681 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00364 | BP APPARATUS ANEROID | GUNJAL NURSE | 9226201417 | 24599 | 04-Feb-2019 | 09:01 |  |  | 04-Feb-2019 | 14:20 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING. | Closed | NO INFO |  | 5:19:0 | 1
682 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00364 | BP APPARATUS ANEROID | SONALI NURSE | 8657733630 | 27524 | 11-May-2019 | 09:37 |  |  | 15-May-2019 | 16:53 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS LEAK. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 103:16:0 | 5
683 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00364 | BP APPARATUS ANEROID | SANGITA TATALWAR | 7276114184 | 14284 | 28-Mar-2018 | 08:30 |  |  | 29-Mar-2018 | 10:15 | NOT WORKING | FILLED WITH NEW MERCURY AND REPLACED NEW GLASS TUBE INTO A MACHINE NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 25:45:0 | 2
684 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00364 | BP APPARATUS ANEROID | SANGITA | 7276114184 | 26730 | 12-Apr-2019 | 08:10 |  |  | 15-Apr-2019 | 10:16 | NOT WORKING | CHECK AND FOUND BP APPRATUS BULB IS LEAK NEED TO CHANGE . NOW CHANGE THE BULB MACHINE IS WORKING GOOD | Closed | NO INFO |  | 74:6:0 | 4
685 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00364 | BP APPARATUS ANEROID | SONALI FHATKAL | 8208297987 | 24846 | 10-Feb-2019 | 09:05 |  |  | 11-Feb-2019 | 16:10 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING. | Closed | NO INFO |  | 31:5:0 | 2
686 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00364 | BP APPARATUS ANEROID | SANGITA TOTALWAR | 7276114182 | 32449 | 14-Sep-2019 | 08:08 |  |  | 16-Sep-2019 | 14:49 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS LEAK. CHANGED THE OLD DEFECTIVE BLADDER AND BULB WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 54:41:0 | 3
687 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00364 | BP APPARATUS ANEROID | P I NAGDEVE | 9403379340 | 10743 | 19-Dec-2017 | 09:45 |  |  | 20-Dec-2017 | 23:57 | NOT WORKING | REPLACED THE GLASS TUBE | Closed | NO INFO |  | 38:12:0 | 2
688 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00366 | BP APPARATUS ANEROID | PAYAL PANDEY | 8380886773 | 9573 | 23-Nov-2017 | 08:55 |  |  | 24-Nov-2017 | 16:38 | NOT WORKING | REPLACED METAL VALVE, BLADDER, TUBE | Closed | 605019B283 |  | 31:43:0 | 2
689 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00368 | BP APPARATUS ANEROID | PATIL | 9420866915 | 11291 | 04-Jan-2018 | 09:38 |  |  | 04-Jan-2018 | 20:11 | NOT WORKING | REPLACED THE BLADDER, TUBE, BULB WITH VALVE | Closed | 8370 |  | 10:33:0 | 1
690 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00368 | BP APPARATUS ANEROID | SHENDE | 9421714662 | 34272 | 23-Oct-2019 | 09:06 |  |  | 23-Oct-2019 | 10:58 | NOT WORKING | CHECK AND FOUND BLADDER IS HAVING LEAKAGE. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE SPARES WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 8370 |  | 1:52:0 | 1
691 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00373 | Fumigation Machine | HARSHA BHUSARI | 9996022114 | 32441 | 13-Sep-2019 | 09:21 |  |  | 13-Sep-2019 | 16:09 | NOT WORKING | CHECK AND FOUND POWER PLUCK IS DEFECTIVE. FIXED THE NEW POWER PLUCK AND POWER SUPPLY ISSUE IS RESOLVED. FUMIGATION MACHINE IS NOW WORKING. | Closed | NO INFO |  | 6:48:0 | 1
692 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00376 | BP APPARATUS ANEROID | MAYA WASNIK | 8149503946 | 35029 | 12-Nov-2019 | 09:49 |  |  | 12-Nov-2019 | 16:26 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN BLADDER AND VALVE.CUFF IS ALSO WORN OUT.CHANGED THE DEFECTIVE CUFF, BLADDER AND VALVE WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 6:37:0 | 1
693 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00376 | BP APPARATUS ANEROID | WASNIK NURSE | 8149503946 | 13693 | 12-Mar-2018 | 08:36 |  |  | 13-Mar-2018 | 15:51 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB WITH CONNECTOR. NOW MACHINE IS WORKING CONDITION | Closed | NO INFO |  | 31:15:0 | 2
694 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00376 | BP APPARATUS ANEROID | RAMTEKE NURSE | 8149503946 | 7160 | 16-Sep-2017 | 09:11 |  |  | 17-Sep-2017 | 11:40 | NOT WORKING | CLEAN PROPERLY THE GLASS TUBE THEN OK | Closed | NO INFO |  | 26:29:0 | 2
695 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00376 | BP APPARATUS ANEROID | M WASNIK | 8149503946 | 32168 | 06-Sep-2019 | 08:05 |  |  | 07-Sep-2019 | 15:12 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER, BULB AND VALVE. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 31:7:0 | 2
696 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00376 | BP APPARATUS ANEROID | MAYA WASNIK | 8149503946 | 15067 | 19-Apr-2018 | 09:08 |  |  | 20-Apr-2018 | 09:33 | NOT WORKING | REPLACED NEW BLADDER,CUFF .NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 24:25:0 | 2
697 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00391 | MONITOR | BAVANKAR | 9422132247 | 18543 | 28-Jul-2018 | 08:43 |  |  | 06-Nov-2018 | 11:10 | NOT WORKING | CHECK AND FOUND SPO2 PROB IS DEFECTIVE .CHANGE THE SAME NOW MONITOR IS WORKING | Closed | NO INFO |  | 2426:27:0 | 102
698 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00391 | MONITOR | HEMLATA | 9096022114 | 41410 | 01-Apr-2020 | 09:07 |  |  | 01-Apr-2020 | 14:14 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS HAVING LEAKAGE. FIXED THE NEW NIBP CUFF. NOW, BP IS BEING MONITORED AND MEASURED. MONITOR IS NOW WORKING. | Closed | NO INFO |  | 5:7:0 | 1
699 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00398 | ADULT WEIGHING MACHINE | RAJANI SAKURE | 8698396820 | 39946 | 24-Feb-2020 | 09:06 |  |  | 25-Feb-2020 | 16:14 | NOT WORKING | CHECK AND FOUND MACHINE IS NOT TURNING 'ON'. FIXED THE LOOSE CONNECTION FROM POWER SUPPLY BOARD. POWER SUPPLY FAILURE ISSUE RESOLVED. ADULT WEIGHING MACHINE IS NOW WORKING. | Closed | NO INFO |  | 31:8:0 | 2
700 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00399 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | RAJANI SAKURE | 8698396820 | 27583 | 14-May-2019 | 08:20 |  |  | 15-May-2019 | 16:47 | NOT WORKING | CHECK AND FOUND LOAD CELL WIRE IS OPENED WHICH IS TO BE CONNECTED TO CONTROLLER BOARD OF MACHINE. FIXED THE LOAD CELL WIRE TO THE CONTROLLER BOARD AND NOW WEIGHT IS EXACTLY DISPLAYED ON SCREEN. WEIGHING MACHINE IS NOW WORKING. | Closed | NO INFO |  | 32:27:0 | 2
701 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00399 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | RAJANI PATIL | 8698396820 | 16991 | 20-Jun-2018 | 09:55 |  |  | 21-Jun-2018 | 04:30 | NOT WORKING | CHECK AND FOUND WEIGHING MACHINE NOT SHOWING STABLE WEIGHT. FOUND SCRATCH PEN INSERTED UNDER LOAD CELL. REMOVE SCETCH PEN. NOW WEIGHING MACHINE WORKING SATISFACTORY. | Closed | NO INFO |  | 18:35:0 | 2
702 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00399 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | RAJANI NURSE | 8698396020 | 22228 | 19-Nov-2018 | 09:13 |  |  | 22-Dec-2018 | 11:05 | NOT WORKING | CHECK AND FOUND LOAD CELL PLUCK WIRE IS INTERNALLY BROKEN.FIXED THE WIRE AND CONNECTION OF LOAD CELL FLICK ELECTRIC BODY WEIGHING MACHINE IS NOW WORKING | Closed | NO INFO |  | 793:52:0 | 34
703 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00400 | NEBULIZER ELECTRIC | PRIYA SAHANE | 9325073518 | 29864 | 12-Jul-2019 | 08:50 |  |  | 15-Jul-2019 | 15:08 | NOT WORKING | CHECK AND FOUND MASK TUBE IS FAULTY. AS IT IS CONSUMABLE ITEM, IT WONT BE PROVIDED BY OUR COMPANY.REST, THE MACHINE IS WORKING NORMALLY. | Closed | NO INFO |  | 78:18:0 | 4
704 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00407 | PULSE OXIMETER LINE POWERED | SANGITA TOTALWAR | 7276117184 | 21467 | 19-Oct-2018 | 08:12 |  |  | 26-Oct-2018 | 15:00 | NOT WORKING | CHECK AND FOUND SPO2 PROBE PIN IS BENT AND CANNOT BE FIXED TO FEMALE CONNECTOR. SPO2 PROBE PIN IS STRAIGHTENED AND PROPERLY FIXED TO FEMALE CONNECTOR. SPO2 IS BEING MONITORED AND MEASURED. PULSE OXIMETER IS WORKING. | Closed | 0402168A.5AP012 |  | 174:48:0 | 8
705 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00423 | BP APPARATUS ANEROID | ARCHANA VASNIK | 8378035077 | 12027 | 20-Jan-2018 | 09:22 |  |  | 20-Jan-2018 | 18:07 | NOT WORKING | REPLACED THE BLADDER,TUBING AND CUFF | Closed | NO INFO |  | 8:45:0 | 1
706 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00423 | BP APPARATUS ANEROID | S NARVATE | 9960144881 | 23626 | 05-Jan-2019 | 08:04 |  |  | 05-Jan-2019 | 11:00 | NOT WORKING | CHECK AND FOUND BULB IS LEAK CHANGED THE BULB. BP APPRATUS IS NOW WORKING | Closed | NO INFO |  | 2:56:0 | 1
707 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00423 | BP APPARATUS ANEROID | S MARVATE | 9960144881 | 27693 | 17-May-2019 | 08:20 |  |  | 18-May-2019 | 10:45 | NOT WORKING | CHECK AND FOUND BLADDER & BULB IS LEAK. CUFF IS WORN OUT. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 26:25:0 | 2
708 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00423 | BP APPARATUS ANEROID | S MARVATE | 9960144881 | 32124 | 05-Sep-2019 | 08:26 |  |  | 07-Sep-2019 | 14:10 | NOT WORKING | CHECK AND FOUND LEAKAGE IN VALVE AND BULB. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 53:44:0 | 3
709 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00423 | BP APPARATUS ANEROID | S MARGATE | 9960144881 | 15390 | 30-Apr-2018 | 09:10 |  |  | 30-Apr-2018 | 15:02 | NOT WORKING | BLADDER IS TORN ,CHANGED NEW ONE BLADDER AND THE METAL CONNECTOR | Closed | NO INFO |  | 5:52:0 | 1
710 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00424 | RADIANT WARMER | ANITA LINGE | 9049217699 | 10127 | 06-Dec-2017 | 10:20 |  |  | 09-Dec-2017 | 11:25 | NOT WORKING | SOLDERED THE AIR AND SKIN PROBE CONNECTOR,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 73:5:0 | 4
711 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00425 | RADIANT WARMER | ANITA | 9049217699 | 35329 | 18-Nov-2019 | 09:16 |  |  | 19-Nov-2019 | 14:15 | NOT WORKING | CHECK AND FOUND POWER CABLE IS BURNT. FIXED THE NEW POWER CABLE. POWER SUPPLY BOARD IS CHECKED. CUT-OFF OF HEATER AT HIGH AND LOW BODY TEMPERATURE IS CHECKED. MACHINE IS NOW WORKING. | Closed | NO INFO |  | 28:59:0 | 2
712 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00425 | RADIANT WARMER | ANITA LINGE | 9049217699 | 24810 | 08-Feb-2019 | 09:10 |  |  | 11-Feb-2019 | 14:27 | NOT WORKING | CHECK AND FOUND POWER SOCKET OF RADIANT WARMER IS BURNT. CHANGED THE OLD WITH THE NEW POWER SOCKET. POWER SUPPLY PROBLEM IS RESOLVED.RADAINT WARMER IS NOW WORKING. | Closed | NO INFO |  | 77:17:0 | 4
713 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00425 | RADIANT WARMER | ANJALI LIMJE | 9049217699 | 33749 | 12-Oct-2019 | 09:10 |  |  | 12-Oct-2019 | 15:45 | NOT WORKING | CHECK AND FOUND SKIN PROBE TEMPERATURE COULD NOT BE SENSED THROUGH DISPLAY CONTROLLER BOARD AND DUE TO WHICH HEATER DOESN'T TURNS ON. FIXED THE LOOSE CONNECTION OF SKIN PROBE THROUGH DISPLAY CONTROLLER BOARD. NOW, SKIN TEMPERATURE IS SENSED AND MONITORED. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO |  | 6:35:0 | 1
714 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00425 | RADIANT WARMER | ANITA LINGE | 9049217699 | 10126 | 06-Dec-2017 | 10:20 |  |  | 08-Dec-2017 | 10:51 | NOT WORKING | REPLACED NEW ELECTRONIC CHOKE INTO AN MACHINE, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 48:31:0 | 3
715 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00427 | BP APPARATUS ANEROID | ANITA NURSE | 9049217699 | 38624 | 28-Jan-2020 | 08:45 |  |  | 28-Jan-2020 | 10:44 | NOT WORKING | CHECK AND FOUND BLADDER AND VALVE IS HAVING LEAKAGE. FIXED THE NEW BLADDER AND VALVE. BP APPARATUS IS NOW WORKING. | Closed | A3C100221 |  | 1:59:0 | 1
716 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00427 | BP APPARATUS ANEROID | ANJALI LIMJE | 9049217699 | 33750 | 12-Oct-2019 | 09:10 |  |  | 12-Oct-2019 | 14:52 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE OF BLADDER AND BULB. FIXED THE OLD DEFECTIVE SPARES WITH THE NEW ONE. B.P APPARATUS IS NOW WORKING. | Closed | A3C100221 |  | 5:42:0 | 1
717 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00427 | BP APPARATUS ANEROID | Sunita Marbate | 9960144881 | 43766 | 02-Jul-2020 | 08:51 |  |  | 02-Jul-2020 | 11:45 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND CUFF IS WORN OUT. FIXED THE NEW BLADDER AND CUFF. BP APPARATUS IS NOW WORKING. | Closed | A3C100221 |  | 2:54:0 | 1
718 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00427 | BP APPARATUS ANEROID | ANITA LINGE | 9049217699 | 10125 | 06-Dec-2017 | 10:20 |  |  | 07-Dec-2017 | 11:57 | NOT WORKING | REPLACED THE BLADDER AND CUFF | Closed | A3C100221 |  | 25:37:0 | 2
719 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00427 | BP APPARATUS ANEROID | Mashnoon Shaikh | 9834314650 | 41728 | 19-Apr-2020 | 09:00 |  |  | 19-Apr-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS BULB IS DAMAGE. THEN CHANGE THE NEW BULB NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | A3C100221 |  | 8:0:0 | 1
720 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00427 | BP APPARATUS ANEROID | S NARVATE | 9960144881 | 23627 | 05-Jan-2019 | 08:04 |  |  | 07-Jan-2019 | 12:01 | NOT WORKING | CHECK AND FOUND BLADDER & VALVE IS LEAK CHANGED THE NEW BLADDER AND VALVE. NO APPRATUS IS NOW WORKING | Closed | A3C100221 |  | 51:57:0 | 3
721 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00427 | BP APPARATUS ANEROID | S MARVATE | 9960144881 | 37630 | 07-Jan-2020 | 09:19 |  |  | 07-Jan-2020 | 10:11 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS HAVING LEAKAGE. FIXED THE OLD DEFECTIVE BLADDER AND BULB WITH THE NEW ONE. B.P APPARATUS IS NOW WORKING. | Closed | A3C100221 |  | 0:52:0 | 1
722 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00427 | BP APPARATUS ANEROID | Anita Limje | 9049217699 | 42008 | 01-May-2020 | 08:00 |  |  | 02-May-2020 | 11:33 | NOT WORKING | CHECK AND FOUND BLADDER & BULB IS HAVING LEAKAGE. CUFF IS ALSO WORN OUT. FIXED THE NEW BLADDER, BULB AND CUFF. BP APPARATUS IS NOW WORKING. | Closed | A3C100221 |  | 27:33:0 | 2
723 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00432 | CARDIOTOCOGRAPHY MACHINE | Anita Limje | 9049217699 | 42035 | 04-May-2020 | 08:58 |  |  | 04-May-2020 | 14:27 | NOT WORKING | CHECK AND FOUND POWER ADAPTER IS FAULTY AND DUE TO WHICH MACHINE COULD NOT GET CHARGED AND TURNED ON. FIXED THE NEW POWER ADAPTER OF RATING 15V, 3A. POWER SUPPLY FAILURE ISSUE IS RESOLVED. CARDIOTOCOGRAPHY MACHINE IS NOW WORKING. | Closed | JA3C100231 |  | 5:29:0 | 1
724 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00433 | BP APPARATUS ANEROID | Anita Limje | 9049217699 | 42007 | 01-May-2020 | 08:00 |  |  | 02-May-2020 | 10:31 | NOT WORKING | CHECK AND FOUND BLADDER, BULB AND VALVE IS HAVING LEAKAGE. FIXED THE NEW BLADDER, BULB AND VALVE. BP APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 26:31:0 | 2
725 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00443 | BP APPARATUS ANEROID | ANITA NIMJE | 9049217699 | 21599 | 24-Oct-2018 | 08:40 |  |  | 27-Oct-2018 | 14:03 | NOT WORKING | CHECK AND FOUND VALVE AND BULB IS LEAK. CHANGED THE NEW BULB AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 77:23:0 | 4
726 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00443 | BP APPARATUS ANEROID | ANITA | 9249217699 | 27442 | 09-May-2019 | 08:35 |  |  | 13-May-2019 | 15:21 | NOT WORKING | CHECK AND FOUND VALVE IS LEAK AND CUFF IS WORN OUT. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 102:46:0 | 5
727 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00443 | BP APPARATUS ANEROID | ANITA NURSE | 9049217699 | 26613 | 09-Apr-2019 | 08:15 |  |  | 11-Apr-2019 | 15:04 | MERCURY ISSUE | CHECK AND FOUND BLADDER AND BULB IS LEAK. CUFF IS ALSO WORN OUT. FIXED THE NEW BLADDER, BULB AND CUFF. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 54:49:0 | 3
728 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00443 | BP APPARATUS ANEROID | ANITA LINGE | 9049217699 | 24811 | 08-Feb-2019 | 09:10 |  |  | 11-Feb-2019 | 14:33 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS LEAK. CUFF IS ALSO WORN OUT. CHANGED THE NEW BULB, VALVE, BLADDER AND CUFF. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 77:23:0 | 4
729 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00444 | DIGITAL BP APPARATUS | ANITA NURSE | 9049217699 | 38625 | 28-Jan-2020 | 08:45 |  |  | 28-Jan-2020 | 13:47 | NOT WORKING | CHECK AND FOUND PRESSURE COULD NOT BE ATTAINED THROUGH NIBP CUFF. TUBING ACROSS INFLATING MOTOR IS OPEN. FIXED THE TUBING. NOW, NIBP CUFF IS INFLATED AND B.P IS BEING MEASURED. DIGITAL BP APPARATUS IS NOW WORKING. | Closed | 197 |  | 5:2:0 | 1
730 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00446 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 6787 | 08-Sep-2017 | 08:05 | 09-Sep-2017 | 19:34 | 11-Sep-2017 | 19:34 | NOT WORKING | REPLACED THE 18W FLUORESCENT LAMP | Closed | PH306158 |  | 83:29:0 | 4
731 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00446 | PHOTOTHERAPHY UNIT | JYOTI CHAUDHARI | 7588768076 | 10037 | 04-Dec-2017 | 09:50 |  |  | 05-Dec-2017 | 11:49 | NOT WORKING | REPLACED UV TUBE | Closed | PH306158 |  | 25:59:0 | 2
732 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00446 | PHOTOTHERAPHY UNIT | COUDHRE NURSE | 7588768076 | 7935 | 10-Oct-2017 | 09:07 |  |  | 11-Oct-2017 | 15:36 | NOT WORKING | REPLACE THE STARTER | Closed | PH306158 |  | 30:29:0 | 2
733 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00450 | NEBULIZER ELECTRIC | JOYTI CHOUDHARI | 7588768076 | 21894 | 05-Nov-2018 | 09:07 |  |  | 10-Nov-2018 | 15:11 | STEAM PROBLEM | CHECK AND FOUND MACHINE IS WORKING. | Closed | NO INFO |  | 126:4:0 | 6
734 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00452 | SYRINGE PUMP | CHAUDARY NURSE | 7588768076 | 6982 | 12-Sep-2017 | 09:42 |  |  | 13-Sep-2017 | 19:20 | NOT WORKING | RESET THE SETTING AND GUIDE THE TECHNICIAN | Closed | MR850AEU |  | 33:38:0 | 2
735 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00456 | INTENSIVE CARE VENTILATOR (NEONATAL OR PEDEATRIC) | SCOTT NURSE | 9764971169 | 29887 | 13-Jul-2019 | 08:09 |  |  | 18-Jul-2019 | 18:13 | NOT WORKING | ALL MODES AND PARAMETERS OF VENTILATION IS CHECKED AND MONITORED. OXYGEN AND FLOW SENSOR IS CLEANED. PEEP AND TIDAL VOLUME IS CHECKED AND READING AND DIFFERENCE OF EXHALATION AND INHALATION IS ABSOLUTE AND OK. INTENSIVE CARE VENTILATION IS WORKING NORMALLY. | Closed | ASED0143 |  | 130:4:0 | 6
736 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00463 | SUCTION PUMP PORTABLE, ELECTRONIC | C S GADDAM | 9850236795 | 34339 | 24-Oct-2019 | 09:05 |  |  | 24-Oct-2019 | 11:05 | NOT WORKING | CHECK AND FOUND MOTOR IS NOT TURNING 'ON'. CONNECTING WIRE ACROSS MOTOR IS OPEN. FIXED THE CONNECTION AND SUCTION MOTOR IS WORKING AND ALSO THE MACHINE. DESIRED SUCTION PRESSURE IS GENERATED. | Closed | NO INFO |  | 2:0:0 | 1
737 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00463 | SUCTION PUMP PORTABLE, ELECTRONIC | K P RAMTEKE | 9518511380 | 15186 | 23-Apr-2018 | 09:14 |  |  | 03-Sep-2018 | 14:18 | NOT WORKING | CHECKED AND FOUND SUCTION MOTOR IS DEFECTIVE AND ASSEMBLY IS MISSING. ORDER HAS BEEN RAISED FOR NEW MOTOR AND ASSEMBLY BUT SAME MOTOR AND ASSEMBLY COULD BE FOUND ANYWHERE. SO, THIS SUCTION MACHINE IS NOT REPAIRABLE. UNDER PROCESS FOR CONDEMNATION. | Closed | NO INFO |  | 3197:4:0 | 134
738 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00481 | DEFIBRILLATOR (AED/DEFIB) | SHEELA BAVANKAR | 94221322247 | 19857 | 31-Aug-2018 | 09:05 |  |  | 01-Sep-2018 | 10:05 | NOT WORKING | CHECK AND FOUND CHEST PADDLE IS DETACHED.FIXED THE METALIC CHEST PADDLE.DEFIBRILLATOR IS NOW WORKING. | Closed | NO INFO |  | 25:0:0 | 2
739 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00483 | PULSE OXIMETER LINE POWERED | TRISHNA POTPODE | 9422474142 | 19182 | 14-Aug-2018 | 09:36 |  |  | 21-Nov-2018 | 11:24 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS RAT BITTER ALONG WITH NIBP CUFF. SPO2 MODULE IS BROKEN. POWER ADAPTER IS ALSO RAT BITTER. CHANGED THE NEW SPO2 PROBE NIBP CUFF AND SPO2 MODULE. POWER ADAPTER IS REPAIRED. NOW MACHINE IS WORKING | Closed | NA |  | 2377:48:0 | 100
740 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00483 | PULSE OXIMETER LINE POWERED | PRIYA SAHANE | 9325073518 | 29865 | 12-Jul-2019 | 08:50 |  |  | 15-Jul-2019 | 15:11 | NOT WORKING | CHECK AND FOUND POWER ADAPTER OF 9V IS DEFECTIVE. CHANGED THE OLD DEFECTIVE POWER ADAPTER WITH THE NEW ONE. PULSE OXIMETER IS NOW WORKING. | Closed | NA |  | 78:21:0 | 4
741 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00484 | BP APPARATUS ANEROID | SHRUTI TAVALE | 9623779643 | 33198 | 28-Sep-2019 | 09:16 |  |  | 28-Sep-2019 | 10:33 | NOT WORKING | CHEAK AND FOUND BLADDER IS LEAK. CHANGE THE NEW BLADDER . B. P. APPRAUTUS IS NOW WORKING AT PRESENT. | Closed | NA |  | 1:17:0 | 1
742 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00493 | PULSE OXIMETER LINE POWERED | PRATIMA MESHRAM | 9404581302 | 24804 | 08-Feb-2019 | 09:07 |  |  | 08-Feb-2019 | 14:04 | NOT WOKING | CHECK AND FOUND SPO2 PROBE IS NOT WORKING. CHANGED THE OLD WITH NEW 5 PIN, SINGLE LOCK NEONATAL SPO2 PROBE. PULSE OXIMETER IS NOW WORKING. | Closed | NA |  | 4:57:0 | 1
743 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00493 | PULSE OXIMETER LINE POWERED | CHAUDHARI | 7588768076 | 42144 | 11-May-2020 | 08:10 |  |  | 12-May-2020 | 10:55 | NOT WORKING | CHECK AND FOUND FEMALE SPO2 CONNECTOR IS DETACHED AND WENT INSIDE THE MACHINE. FIXED THE SPO2 FEMALE CONNECTOR. SPO2 PROBE IS BEING FIXED IN FEMALE CONNECTOR AND CHECKED. PULSE OXIMETER IS NOW WORKING. | Closed | NA |  | 26:45:0 | 2
744 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00495 | RADIANT WARMER | JYOTI CHAUDHARI | 7588768076 | 35855 | 28-Nov-2019 | 08:02 |  |  | 29-Nov-2019 | 04:11 | NOT WORKING | CHECK AND FOUND POWER CORD IS FAULTY. FIXED THE NEW POWER CORD. CUT-OFF OF HEATER AT HIGH AND LOW BODY TEMPERATURE IS CHECKED AND FOUND TO BE OK. RADIANT WARMER IS NOW WORKING. | Closed | NA |  | 20:9:0 | 2
745 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00503 | RO PLANT FOR HAEMODIALYSIS | HEMA BORZAR | 9767558503 | 38574 | 27-Jan-2020 | 09:20 |  |  | 27-Jan-2020 | 16:45 | NOT WORKING | CHECK AND FOUND VALVE OF SAND CARBON FILTER IS DEFECTIVE AND DUE TO WHICH THERE IS LEAKAGE OF WATER FROM IT. ALSO CONTROL PANEL IS NOT WORKING. FIXED THE NEW VALVE OF SAND CARBON FILTER AND CONTROL PANEL. RO PLANT IS NOW WORKING. | Closed | NA |  | 7:25:0 | 1
746 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00503 | RO PLANT FOR HAEMODIALYSIS | NILESH | 9960704591 | 29420 | 03-Jul-2019 | 09:38 |  |  | 15-Jul-2019 | 06:45 | NOT WORKING | CHECK AND FOUND MEMBRANES OF RO IS DEGRADED AND DUE TO WHICH OUTPUT WATER FROM RO IS CONTAMINATED. ALSO PIPING IS LOOSELY CONNECTED CAUSING LEAKAGE PROBLEM. CHANGED BOTH THE MEMBRANE THROUGH VENDOR AND PIPING IS FIXED . NEW LEAKAGE PROBLEM IS SOLVED AND WATER OBTAINED FROM RO IS OF GOOD QUALITY. RO PLANT IS NOW WORKING. | Closed | NA |  | 285:7:0 | 13
747 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00503 | RO PLANT FOR HAEMODIALYSIS | HEMA | 9767558503 | 40938 | 16-Mar-2020 | 09:12 |  |  | 16-Mar-2020 | 16:46 | NOT WORKING | CHECK AND FOUND PRESSURE PUMP IS NOT WORKING AS ITS INTERNAL WIRING IS SHORT. NEW INTERNAL WIRING IS BEING DONE AND PRESSURE PUMP IS FIXED. NOW, RO WATER IS EJECTED OUT AND IS COLLECTED IN STORAGE TANK. RO PLANT IS NOW WORKING. | Closed | NA |  | 7:34:0 | 1
748 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00504 | RO PLANT FOR HAEMODIALYSIS | HEMA BORJHAR | 9767558503 | 37543 | 04-Jan-2020 | 09:06 |  |  | 04-Jan-2020 | 17:21 | NOT WORKING | CHECK AND FOUND HIGH PRESSURE PUMP IS NOT WORKING AS WATER COULD NOT BE EJECTED OUT TO STORAGE TANK. TDS MEASURED OF FILTERED WATER IS VERY HIGH. BOTH THE MEMBRANES ARE CHANGED,SAND CARBON FILTERS IS CHANGED AND NEW HIGH PRESSURE PUMP IS INSTALLED. NOW, DESIRED AMOUNT OF WATER IS EJECTED TO STORAGE TANK AND TDS OF FILTERED WATER IS NORMAL. | Closed | NA |  | 8:15:0 | 1
749 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00507 | OPERATING MICROSCOPE | JITENDRA PADWAL | 8149913123 | 22604 | 04-Dec-2018 | 08:17 |  |  | 05-Dec-2018 | 11:03 | NOT WORKING | CHECK AND FOUND OPERATING MICROSCOPE, ILLUMINATION OR LAMP POWER IS VERY LOW. CHANGE THE WITH NEW ILLUMINATION LAMP, NOW MACHINE IS WORKING | Closed | NA |  | 26:46:0 | 2
750 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00510 | FUMIGATION MACHINE | MEENA | 9421714662 | 34180 | 22-Oct-2019 | 08:20 |  |  | 23-Oct-2019 | 15:10 | NOT WORKING | CHECK AND FOUND SUCTION TUBES IS SHORT AND HAVING LEAKAGE. CHANGED THE NEW TUBINGS AND PIPING IS CLEANED. NOW, FUMIGATION MACHINE IS WORKING WITH PROPER FORMATION OF FOGG. | Closed | NO INFO |  | 30:50:0 | 2
751 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00516 | SHORT WAVE DIATHERMY | DR PRITI | 9764882444 | 17249 | 28-Jun-2018 | 08:05 |  |  | 03-Jul-2018 | 15:45 | NOT WORKING | CHECK AND FOUND DISPLAY IS NOT WORKING, DISPLAY TIMER CARD AND FILAMENT. RECTIFIED ALL PROBLEMS. NOW MACHINE IS WORKING SATISFACTORY. | Closed | NA |  | 127:40:0 | 6
752 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00516 | SHORT WAVE DIATHERMY | DR CHOLE | 9764882444 | 38742 | 30-Jan-2020 | 08:40 |  |  | 30-Jan-2020 | 11:27 | NOT WORKING | CHECK AND FOUND HEATING PADS IS BURNT AND DUE TO WHICH IT IS NOT WORKING. FIXED THE NEW HEATING PAD. CHECKED THE MACHINE ON DIFFERENT POWER AND INTENSITY MODE. SHORT WAVE DIATHERMY MACHINE IS NOW WORKING. | Closed | NA |  | 2:47:0 | 1
753 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00517 | SHORT WAVE DIATHERMY | DR CHOLE | 9764882444 | 38743 | 30-Jan-2020 | 08:41 |  |  | 30-Jan-2020 | 11:21 | NOT WORKING | CHECK AND FOUND HEATING PAD IS BURNT AND DUE TO WHICH IT IS NOT WORKING. FIXED THE NEW HEATING PAD. CHECKED THE MACHINE ON DIFFERENT POWER AND INTENSITY MODE. SHORT WAVE DIATHERMY MACHINE IS NOW WORKING. | Closed | NA |  | 2:40:0 | 1
754 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00518 | SHORT WAVE DIATHERMY | DR PRITI CHOLE | 9764882444 | 38848 | 01-Feb-2020 | 09:12 |  |  | 01-Feb-2020 | 11:34 | NOT WORKING | CHECK AND FOUND HEATING PADS IS BURNT AND DUE TO WHICH IT IS NOT WORKING. FIXED THE NEW HEATING PADS. CHECKED THE MACHINE ON DIFFERENT POWER AND INTENSITY MODE. SHORT WAVE DIATHERMY MACHINE IS NOW WORKING. | Closed | NA |  | 2:22:0 | 1
755 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00518 | SHORT WAVE DIATHERMY | DR PRITI | 9764882444 | 17250 | 28-Jun-2018 | 08:05 |  |  | 03-Jul-2018 | 15:45 | NOT WORKING | CHECK AND FOUND DISPLAY IS NOT WORKING, PROBLEM WITH DISPLAY TIMER CARD AND FILAMENT. RECTIFIED ALL PROBLEMS, NOW MACHINE WORKING SATISFACTORY. | Closed | NA |  | 127:40:0 | 6
756 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00519 | SHORT WAVE DIATHERMY | DR PRITI | 9764882444 | 17251 | 28-Jun-2018 | 08:05 |  |  | 03-Jul-2018 | 15:40 | NOT WORKING | CHECK AND FOUND DISPLAY TIMER CARD IS MISSING IN THIS MACHINE ALSO PROBLEM WITH INTERNAL WIRING. FIX THE PROBLEM NOW MACHINE IS WORKING. THEY CAN USE THIS MACHINE BY MANUAL TIMING. | Closed | NA |  | 127:35:0 | 6
757 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00519 | SHORT WAVE DIATHERMY | DR PRITI CHOLE | 9764882444 | 38849 | 01-Feb-2020 | 09:12 |  |  | 01-Feb-2020 | 10:36 | NOT WORKING | CHECK AND FOUND HEATING PADS IS NOT WORKING. FIXED THE NEW HEATING PADS. CHECKED THE MACHINE ON DIFFERENT POWER AND INTENSITY MODE. SHORT WAVE DIATHERMY MACHINE IS NOW WORKING. | Closed | NA |  | 1:24:0 | 1
758 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00521 | ULTRASONIC THERAPY UNIT | DR PRITI | 9764882444 | 17252 | 28-Jun-2018 | 08:05 |  |  | 03-Jul-2018 | 15:50 | NOT WORKING | CHECK AND FOUND DISPLAY IS NOT WORKING. CHECK INTERNAL WIRING FOUND WRONG. FIX THE WIRING PROBLEM. NOW MACHINE IS WORKING SATISFACTORY. | Closed | 1275 |  | 127:45:0 | 6
759 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00525 | DIGITAL MUSCLE STIMULATOR | DR PRITI | 9764882444 | 17253 | 28-Jun-2018 | 08:05 |  |  | 04-Aug-2018 | 11:10 | NOT WORKING | CHECK AND FOUND ,THE INTENSITY OF EQUIPMENT NOT SATISFACTORY PROVIDING LOW FREQUENCY IN HIGH VALUE CHECK PRESATE FOUND FAULTY REPLACE SAME NOW EQUIPMENT IS WORKING | Closed | NA |  | 891:5:0 | 38
760 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00529 | INFRARED LAMP | DR PREETI | 9764882444 | 40843 | 13-Mar-2020 | 08:25 |  |  | 13-Mar-2020 | 11:39 | NOT WORKING | CHECK AND FOUND INFRARED BULB IS FUSE. REPLACE THE DEFECTIVE IR BULB WITH THE NEW ONE. INFRARED LAMP IS NOW WORKING. | Closed | NO INFO |  | 3:14:0 | 1
761 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00530 | INFRARED LAMP | DR PREETI | 9764882444 | 40844 | 13-Mar-2020 | 08:25 |  |  | 13-Mar-2020 | 11:50 | NOT WORKING | CHECK AND FOUND INFRARED BULB IS FUSE. REPLACE THE DEFECTIVE IR BULB WITH THE NEW ONE. BULB HOLDER AND STAND IS BEING FIXED. INFRARED LAMP IS NOW WORKING. | Closed | NO INFO |  | 3:25:0 | 1
762 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00533 | BP APPARATUS ANEROID | MANDA DHUPARE | 9423173691 | 22521 | 30-Nov-2018 | 09:14 |  |  | 01-Dec-2018 | 12:25 | NOT WORKING | CHECK AND FOUND NO MERCURY IS PRESENT BLADDER AND VALVE HAVE LEAKAGE. REFILLED THE MERCURY CHANGED THE NEW BLADDER AND VALVE.B.P APPRATUS IS NOW WORKING | Closed | NO INFO |  | 27:11:0 | 2
763 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00536 | BP APPARATUS ANEROID | MANDA DHUPARE | 9423173691 | 22519 | 30-Nov-2018 | 09:14 |  |  | 01-Dec-2018 | 01:15 | NOT WORKING | CHECK AND FOUND BULB VALVE AND BLADDER IS LEAK. CUFF IS ALSO WORN.CHANGED THE NEW BULB VALVE.,BLADDER AND CUFF.B.P APPRATUS IS NOW WORKING | Closed | NA |  | 16:1:0 | 2
764 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00536 | BP APPARATUS ANEROID | B.V.AYALWAR | 9405949739 | 33500 | 07-Oct-2019 | 08:20 |  |  | 07-Oct-2019 | 11:52 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF, BULB AND VALVE ARE MISSING. FIXED THE NEW BLADDER, CUFF, BULB AND VALVE. B.P APPARATUS IS NOW WORKING. | Closed | NA |  | 3:32:0 | 1
765 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00537 | LARYNGOSCOPE | PORNIMA POTDUKE | 9881409188 | 27988 | 28-May-2019 | 09:25 |  |  | 30-May-2019 | 11:04 | NOT WORKING | CHECK AND FOUND BATTERY IS NOT PRESENT AND DUE TO WHICH BULB IS NOT GLOWING. FIXED THE BATTERY AND BULB IS GLOWN.LARYNGOSCOPE IS NOW WORKING. | Closed | NO INFO |  | 49:39:0 | 3
766 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00537 | LARYNGOSCOPE | KANCHAN MUL | 9763230995 | 35761 | 26-Nov-2019 | 08:32 |  |  | 26-Nov-2019 | 11:14 | NOT WORKING | LYRIGOSCOP BULB IS NOT WORKING. BULB IS FUSE, SO BULB IS CHANGE AND LYARINGOSCOP IS NOW WORKING AT PRESENT. | Closed | NO INFO |  | 2:42:0 | 1
767 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00541 | Suction pump Portable, Electronic | DHUPARE NURSE | 9423113691 | 35977 | 30-Nov-2019 | 08:20 |  |  | 30-Nov-2019 | 18:46 | NOT WORKING | CHECK AND FOUND POWER ON/OFF SWITCH IS FAULTY. FIXED THE NEW POWER ON/OFF SWITCH. POWER SUPPLY ISSUE IS RESOLVED. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO |  | 10:26:0 | 1
768 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00543 | ECG UNIT 3 CHANNEL PORTABLE | Smt. Maya Wasnik | 8149503946 | 43765 | 02-Jul-2020 | 08:50 |  |  | 02-Jul-2020 | 10:45 | NOT WORKING | CHECK AND FOUND DISPLAY IS AUTOMATICALLY TURNING 'OFF' IN BETWEEN WHILE TAKING THE ECG. FIXED THE LOOSE CONNECTION BETWEEN DISPLAY CONTROLLER BOARD AND MAIN INTERFACING PCB BOARD. FLUCTUATIONS AND TURNING 'OFF' & 'ON' OF DISPLAY SCREEN IS RESOLVED. ECG MACHINE IS NOW WORKING. | Closed | NA |  | 1:55:0 | 1
769 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00545 | DIGITAL BP APPARATUS | N M CHAKALE | 9422824065 | 38265 | 22-Jan-2020 | 09:18 |  |  | 22-Jan-2020 | 11:36 | NOT WORKING | CHECK AND FOUND INFLATING INTERNAL MOTOR IS NOT WORKING. FIXED THE NEW INFLATING INTERNAL MOTOR. NOW, CUFF IS INFLATED. B.P IS BEING MONITORED. ALSO NEW POWER ADAPTER IS HANDED OVER. DIGITAL BP APPARATUS IS NOW WORKING. | Closed | 197 |  | 2:18:0 | 1
770 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00557 | BP APPARATUS ANEROID | SUNANDA | 8007445066 | 28521 | 13-Jun-2019 | 08:48 |  |  | 13-Jun-2019 | 16:39 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT. BULB IS ALSO LEAK. CHANGED THE DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NA |  | 7:51:0 | 1
771 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00559 | BP APPARATUS ANEROID | SUNANDA | 8007445066 | 28520 | 13-Jun-2019 | 08:48 |  |  | 13-Jun-2019 | 17:41 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT. CHANGED THE NEW CUFF. B. P APPARATUS IS NOW WORKING. | Closed | NA |  | 8:53:0 | 1
772 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00559 | BP APPARATUS ANEROID | S. N Umbarkar | 8007445066 | 42822 | 02-Jun-2020 | 09:10 |  |  | 02-Jun-2020 | 09:45 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER, BULB AND VALVE. FIXED THE NEW BLADDER, BULB AND VALVE. BP APPARATUS IS NOW WORKING. | Closed | NA |  | 0:35:0 | 1
773 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00567 | DIGITAL BP APPARATUS | HEMLATA | 9096022114 | 41411 | 01-Apr-2020 | 09:07 |  |  | 01-Apr-2020 | 11:03 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS HAVING LEAKAGE. FIXED THE NEW NIBP CUFF. NOW, BP IS BEING MEASURED. DIGITAL BP APPARATUS IS NOW WORKING. | Closed | 15066802420 |  | 1:56:0 | 1
774 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00569 | BP APPARATUS ANEROID | DR APARNA JAKKAL | 8888836389 | 34883 | 08-Nov-2019 | 09:27 |  |  | 08-Nov-2019 | 17:55 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN BLADDER, BULB AND VALVE. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE SPARES WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NA |  | 8:28:0 | 1
775 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00581 | PULSE OXIMETER LINE POWERED | HARSHA NURSE | 9096022114 | 39673 | 18-Feb-2020 | 09:08 |  |  | 18-Feb-2020 | 10:00 | NOT WORKING | CHECK AND FOUND MACHINE 5AMP FUSE IS SHORT. REPLACE THE NEW ONE 5 AMP FUSE. AND MACHINE IS WORKING PROPERLY AT PRESENT. | Closed | NO INFO |  | 0:52:0 | 1
776 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00587 | BINOCULAR MICROSCOPE | DR APARNA | 8888863689 | 26851 | 16-Apr-2019 | 08:40 |  |  | 18-Apr-2019 | 11:03 | NOT WORKING | CHECK AND FOUND BULB IS FUSED. FIXED THE NEW HALOGEN BULB OF RATING 6V,20W.BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | 151185514 |  | 50:23:0 | 3
777 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00587 | BINOCULAR MICROSCOPE | DR APARNA | 8888863689 | 31426 | 19-Aug-2019 | 09:12 |  |  | 21-Aug-2019 | 14:40 | NOT WORKING | CHECK AND FOUND REVOLVING NOSEPIECE IS FAULTY AND DUE TO WHICH OBJECTIVE LENS COULD NOT BE FIXED AT ONE PLACE AFTER CHANGING THE FOCAL LENGTH OF OBJECTIVE LENS. MR HAS BEEN RAISED AND THIS PART IS CURRENTLY NOT AVAILABLE. LATER IF THIS PART WOULD BE AVAILABLE, MR WOULD BE RAISED AND PART WOULD BE FIXED. AT PRESENT, THIS CALL IS BEING CLOSED. | Closed | 151185514 |  | 53:28:0 | 3
778 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00590 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | TRUPTI GHOGATE | 8600356619 | 31489 | 20-Aug-2019 | 09:17 |  |  | 22-Aug-2019 | 15:35 | NOT WORKING | CHECK AND FOUND METALLIC WELDING ACROSS COIL IS DETACHED BECAUSE OF RUSTING WHICH BLOCKS WATER FROM COMING OUT DUE TO WHICH THERE IS LEAKAGE OF WATER FROM FRONT DOOR PANEL. COMPANY TYPE WELDING COULD NOT BE DONE NOW. IT IS BEYOND OUR SCOPE AS DISCUSSED WITH EXPERT. REST, MACHINE IS IN WORKING CONDITION. | Closed | NA |  | 54:18:0 | 3
779 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00598 | BP APPARATUS ANEROID | SARITA NIRAVAN | 7875830247 | 39436 | 13-Feb-2020 | 08:15 |  |  | 14-Feb-2020 | 10:38 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT AND BULB IS HAVING LEAKAGE. FIXED THE NEW CUFF AND BULB. BP APPARATUS IS NOW WORKING. | Closed | 1.5053204539e+014 |  | 26:23:0 | 2
780 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00598 | BP APPARATUS ANEROID | SARITA NIRAVAN | 7875830247 | 36891 | 21-Dec-2019 | 09:10 |  |  | 23-Dec-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN VALVE IS LEAK. CHANGE THE NEW VALVE NOW MACHINE IS IN WORKING CONDITION. | Closed | 1.5053204539e+014 |  | 50:50:0 | 3
781 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00599 | BP APPARATUS ANEROID | SARITA NIRAVAN | 7875830247 | 36892 | 21-Dec-2019 | 09:10 |  |  | 23-Dec-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER. CHANGE THE NEW BLADDER NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 50:50:0 | 3
782 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00599 | BP APPARATUS ANEROID | PAYAL SHREEKHANDE | 9175091575 | 34563 | 02-Nov-2019 | 09:10 |  |  | 02-Nov-2019 | 11:09 | NOT WORKING | CHECK AND FOUND BLADDER AND VALVE IS LEAK. FIXED THE NEW BLADDER AND VALVE. B. P. APPRATUS IS NOW WORKING. | Closed | NO INFO |  | 1:59:0 | 1
783 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00599 | BP APPARATUS ANEROID | SARITA NIRAVAN | 7875830247 | 39437 | 13-Feb-2020 | 08:15 |  |  | 14-Feb-2020 | 11:40 | NOT WORKING | CHECK AND FOUND BULB IS HAVING LEAKAGE. FIXED THE NEW BULB. BP APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 27:25:0 | 2
784 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00599 | BP APPARATUS ANEROID | RAMTAKE SISTER | 9518511380 | 18841 | 06-Jul-2018 | 08:01 |  |  | 06-Jul-2018 | 17:30 | NOT WORKING | CHECK AND FOUND BLADDER IS FAULTY. REPLACED SAME . NOW MACHINE WORKING SATISFACTORY | Closed | NO INFO |  | 9:29:0 | 1
785 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00600 | BP APPARATUS ANEROID | PAYAL SHREEKHANDE | 9175091575 | 34564 | 02-Nov-2019 | 09:10 |  |  | 02-Nov-2019 | 10:39 | NOT WORKING | CHECK AND FOUND BLADER, CUFF AND BULB IS LEAK. FIXED THE NEW BLADDER, CUFF AND BULB. B. P. APPRAUTUS IS NOW WORKING. | Closed | NO INFO |  | 1:29:0 | 1
786 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00600 | BP APPARATUS ANEROID | HARSHA BHUSARI | 9096022114 | 34466 | 30-Oct-2019 | 08:39 |  |  | 30-Oct-2019 | 16:50 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF, BULB AND VALVE IS MISSING. FIXED THE NEW BLADDER, CUFF, BULB AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 8:11:0 | 1
787 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00601 | BP APPARATUS ANEROID | PAYAL SHREEKHANDE | 9175091575 | 34565 | 02-Nov-2019 | 09:10 |  |  | 02-Nov-2019 | 10:52 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF ,BULB AND EXTENSION TUBE IS LEAK. FIXED THE NEW BLADDER, CUFF, BULB AND EXTENSION TUBE. B. P APPRAUTUS IS NOW WORKING. | Closed | NO INFO |  | 1:42:0 | 1
788 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00602 | BP APPARATUS ANEROID | PAYAL SHREEKHANDE | 9175091575 | 34566 | 02-Nov-2019 | 09:10 |  |  | 02-Nov-2019 | 10:01 | NOT WORKING | CHECK AND FOUND BLADDER, BULB IS LEAK AND VALVE IS MISSING. FIXED THE NEW BLADDER, BULB AND VALVE. B. P APPRAUTUS IS NOW WORKING. | Closed | NA |  | 0:51:0 | 1
789 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00604 | BP APPARATUS ANEROID | PAYAL SHREEKHANDE | 9175091575 | 34567 | 02-Nov-2019 | 09:10 |  |  | 02-Nov-2019 | 10:17 | NOT WORKING | CHECK AND FOUND VALVE AND BULB IS LEAK. FIXED THE NEW VALVE AND BULB. B.P. APPARATUS IS NOW WORKING. | Closed | RM0104069 |  | 1:7:0 | 1
790 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00616 | PULSE OXIMETER LINE POWERED | TRUSHNA | 9422474142 | 31596 | 21-Aug-2019 | 09:40 |  |  | 22-Aug-2019 | 14:21 | NOT WORKING | CHECK AND FOUND DISPLAY CONTROLLER BOARD IS FAULTY. SPO2 PROBE(NEONATAL) IS BROKEN AND POWER ADAPTER IS ALSO FAULTY.CHANGED WITH NEW DISPLAY CONTROLLER BOARD, NEONATAL SPO2 PROBE AND POWER ADAPTER. MONITOR IS NOW WORKING. | Closed | BG1608400034 |  | 28:41:0 | 2
791 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00617 | BP APPARATUS ANEROID | APARNA | 8888863689 | 34796 | 07-Nov-2019 | 09:02 |  |  | 07-Nov-2019 | 15:27 | NOT WORKING | CHECK AND FOUND SPEED REGULATOR IS NOW WORKING AND USER CANNOT SET THE SPEED OF MOTOR AS DESIRED. CHANGED THE DEFECTIVE SPEED REGULATOR WITH THE NEW ONE. CENTRIFUGE MACHINE IS NOW WORKING. | Closed | RM1000166 |  | 6:25:0 | 1
792 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00618 | BP APPARATUS ANEROID | SONALI | 8208297987 | 44206 | 15-Jul-2020 | 08:20 |  |  |  |  | NOT WORKING |  | Open | 254024
793 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00618 | BP APPARATUS ANEROID | Jungan | 9226201417 | 42672 | 28-May-2020 | 08:10 |  |  | 29-May-2020 | 09:30 | NOT WORKING | CHECK AND FOUND BLADDER AND VALVE IS HAVING LEAKAGE. FIXED THE NEW BLADDER AND VALVE. BP APPARATUS IS NOW WORKING. | Closed | 254024 |  | 25:20:0 | 2
794 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00618 | BP APPARATUS ANEROID | SANGITA TOTALWAR | 7276117184 | 21503 | 20-Oct-2018 | 08:00 |  |  | 25-Oct-2018 | 14:15 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN. NO MERCURY IS PRESENT. LEAKAGE IN BLADDER AND VALVE. CHANGED THE NEW GLASS TUBE, BLADDER AND VALVE. REFILLING OF MERCURY IS DONE. NOW MACHINE IS WORKING. | Closed | 254024 |  | 126:15:0 | 6
795 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00619 | BP APPARATUS ANEROID | SANGEETA NURSE | 7276118184 | 36888 | 21-Dec-2019 | 09:09 |  |  | 23-Dec-2019 | 12:10 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BULB AND VALVE IS LEAK. THEN CHANGE THE NEW BULB AND VALVE NOW MACHINE IS IN WORKING CONDITION. | Closed | RM1000166 |  | 51:1:0 | 3
796 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00619 | BP APPARATUS ANEROID | Jungan | 9226201417 | 42685 | 28-May-2020 | 09:20 |  |  | 29-May-2020 | 10:00 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT AND BULB & VALVE IS HAVING LEAKAGE. FIXED THE NEW CUFF, BULB AND VALVE. BP APPARATUS IS NOW WORKING. | Closed | RM1000166 |  | 24:40:0 | 2
797 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00622 | BP APPARATUS ANEROID | SHENDE | 9421714662 | 34270 | 23-Oct-2019 | 09:06 |  |  | 23-Oct-2019 | 11:55 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS HAVING LEAKAGE. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE SPARES WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 2:49:0 | 1
798 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00624 | BP APPARATUS ANEROID | SHENDE | 9421714662 | 34271 | 23-Oct-2019 | 09:06 |  |  | 23-Oct-2019 | 11:01 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN BLADDER, BULB, VALVE AND CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE BLADDER, BULB, VALVE AND CUFF WITH THE NEW ONE. B.P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 1:55:0 | 1
799 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00626 | PULSE OXIMETER LINE POWERED | S MARGATE NURSE | 9960144881 | 24175 | 22-Jan-2019 | 09:01 |  |  | 24-Jan-2019 | 16:16 | CONNECTION ISSUE | CHECK AND FOUND 3 PIN-TOP POWER PLUCK IS BROKEN.CHANGED THE NEW 3-PIN TOP POWER PLUCK. PULSE OXIMETER IS NOW WORKING. | Closed | 61161474 |  | 55:15:0 | 3
800 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00626 | PULSE OXIMETER LINE POWERED | S BARMATE | 9960144881 | 31617 | 22-Aug-2019 | 09:00 |  |  | 23-Aug-2019 | 10:37 | NOT WORKING | CHECK AND FOUND POWER PLUCK ISSUE AND DUE TO WHICH PULSE OXIMETER COULD NOT TURN 'ON'.REPAIRED THE POWER PLUCK. PULSE OXIMETER IS NOW WORKING. | Closed | 61161474 |  | 25:37:0 | 2
801 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00626 | PULSE OXIMETER LINE POWERED | S MARVATE | 9960144881 | 35988 | 30-Nov-2019 | 08:45 |  |  | 30-Nov-2019 | 14:27 | NOT WORKING | CHECK AND FOUND POWER ADAPTER IS FAULTY. FIXED AND REPAIRED THE POWER ADAPTER. POWER SUPPLY ISSUE IS BEING RESOLVED. PULSE OXIMTER IS NOW WORKING. | Closed | 61161474 |  | 5:42:0 | 1
802 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00626 | PULSE OXIMETER LINE POWERED | S MARVATE | 9960144881 | 27691 | 17-May-2019 | 08:20 |  |  | 18-May-2019 | 12:53 | NOT WORKING | CHECK AND FOUND POWER PLUCK OF THE CABLE IS OPENED INSIDE. FIXED THE WIRE OF THE PLUCK. PULSE OXIMETER IS NOW WORKING. | Closed | 61161474 |  | 28:33:0 | 2
803 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00627 | BP APPARATUS ANEROID | S MARVATE | 9960144881 | 32123 | 05-Sep-2019 | 08:25 |  |  | 07-Sep-2019 | 12:08 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN BLADDER AND VALVE. ALSO CUFF IS WORN OUT. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 173094 |  | 51:43:0 | 3
804 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00627 | BP APPARATUS ANEROID | S MARVATE | 9960144881 | 27692 | 17-May-2019 | 08:20 |  |  | 18-May-2019 | 11:30 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK. CHANGED THE OLD DEFECTIVE BLADDER WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 173094 |  | 27:10:0 | 2
805 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00627 | BP APPARATUS ANEROID | S MARVATE | 9960144881 | 37631 | 07-Jan-2020 | 09:19 |  |  | 07-Jan-2020 | 10:40 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND VALVE. FIXED THE OLD DEFECTIVE BLADDER AND VALVE WITH THE NEW ONE. B.P APPARATUS IS NOW WORKING. | Closed | 173094 |  | 1:21:0 | 1
806 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00627 | BP APPARATUS ANEROID | Mashnoon Shaikh | 9834314650 | 41729 | 19-Apr-2020 | 09:00 |  |  | 19-Apr-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS BULB IS DAMAGE. THEN CHANGE THE NEW BULB NOW MACHINE IS IN GOOD WORKING CONDITION | Closed | 173094 |  | 8:0:0 | 1
807 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00627 | BP APPARATUS ANEROID | Sunita Marbate | 9960144881 | 43767 | 02-Jul-2020 | 08:52 |  |  | 02-Jul-2020 | 13:30 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER, BULB AND VALVE. FIXED THE NEW BLADDER, BULB & VALVE. BP APPARATUS IS NOW WORKING. | Closed | 173094 |  | 4:38:0 | 1
808 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00628 | DIGITAL BP APPARATUS | RAHATE | 9822694627 | 34740 | 06-Nov-2019 | 09:18 |  |  | 08-Nov-2019 | 12:40 | NOT WORKING | CHECK AND FOUND BLADDER, BULB, VALVE AND CUFF IS MISSING. FIXED THE NEW BLADDER, BULB, VALVE AND CUFF. B. P APPARATUS IS NOW WORKING. | Closed | 612106 |  | 51:22:0 | 3
809 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00629 | BP APPARATUS ANEROID | DR APARNA JAKKAL | 8888836389 | 34884 | 08-Nov-2019 | 09:27 |  |  | 08-Nov-2019 | 16:57 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN BLADDER, BULB AND VALVE. CHANGED THE OLD DEFECTIVE BLADDER, BULB AND VALVE WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 1506680089 |  | 7:30:0 | 1
810 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00646 | PULSE OXIMETER LINE POWERED | SONALI | 8208297987 | 43802 | 03-Jul-2020 | 08:44 |  |  |  |  | NOT WORKING |  | WIP | BQ1608400528
811 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00649 | NEBULIZER ELECTRIC | Maya Wasnik | 8149503946 | 43691 | 30-Jun-2020 | 09:16 |  |  | 30-Jun-2020 | 14:30 | NOT WORKING | CHECK AND FOUND MOTOR IS JAMMED. OILING OF MOTOR AND ATTACHED AXLE FOR PISTON IS BEING DONE. AIR PRESSURE IS NOW GENERATED WITH FORMATION OF FUMES. NEBULIZER MACHINE IS NOW WORKING. | Closed | NGM966358 |  | 5:14:0 | 1
812 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00651 | NEBULIZER ELECTRIC | Rohani | 9420128032 | 41581 | 10-Apr-2020 | 08:35 |  |  | 10-Apr-2020 | 10:00 | NOT WORKING | CHECK AND FOUND WIRE ACROSS ON/OFF SWITCH IS BURNT DUE TO SHORT CIRCUIT. FIXED THE WIRE AND REPAIRED IT. NEBULIZER MACHINE IS NOW WORKING. | Closed | NGM221911 |  | 1:25:0 | 1
813 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00651 | NEBULIZER ELECTRIC | Maya Wasnik | 8149503946 | 43690 | 30-Jun-2020 | 09:15 |  |  | 30-Jun-2020 | 10:45 | NOT WORKING | CHECK AND FOUND AIR PRESSURE IS NOT GENERATED AS PISTON IS DETACHED FROM THE MOTOR. FIXED THE PISTON THROUGH THE AXLE OF MOTOR. NOW, AIR PRESSURE IS GENERATED THROUGH CONTINOUS MOVEMENT OF PISTON. NEBULIZER MACHINE IS NOW WORKING. | Closed | NGM221911 |  | 1:30:0 | 1
814 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00672 | PULSE OXIMETER LINE POWERED | SUPRIYA | 9689282458 | 40171 | 27-Feb-2020 | 08:32 |  |  | 02-Mar-2020 | 10:00 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS NOT WORKING. CHANGED THE OLD WITH NEW 5 PIN, SINGLE LOCK NEONATAL SPO2 PROBE. PULSE OXIMETER IS NOW WORKING. | Closed | 18010400104 |  | 97:28:0 | 5
815 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00672 | PULSE OXIMETER LINE POWERED | CHAUDHARI | 7588768076 | 42145 | 11-May-2020 | 08:10 |  |  | 12-May-2020 | 10:57 | NOT WORKING | CHECK AND FOUND FAILURE IN POWER SUPPLY AS THERE IS FAULT IN POWER CORD. FIXED THE NEW POWER CORD. FAILURE OF POWER SUPPLY IS RESOLVED. PULSE OXIMETER IS NOW WORKING. | Closed | 18010400104 |  | 26:47:0 | 2
816 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00672 | PULSE OXIMETER LINE POWERED | Jyoti Chaudhari | 7588768076 | 42322 | 16-May-2020 | 09:35 |  |  | 16-May-2020 | 14:25 | NOT WORKING | CHECK AND FOUND SPO2 PROBE NOT WORKING AS IT IS DEFECTIVE. FIXED THE NEW 5 PIN, SINGLE LOCK NEONATAL PROBE. PULSE RATE AND SATURATION OF OXYGEN IS BEING MONITORED AND ITS ACCEPTABLE. PULSE OXIMETER IS NOW WORKING. | Closed | 18010400104 |  | 4:50:0 | 1
817 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00679 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | SUPRIYA | 9689282458 | 40172 | 27-Feb-2020 | 08:35 |  |  | 27-Feb-2020 | 10:00 | NOT WORKING | CHECK AND FOUND TRANFORMER TO MOTHER BOARD WIRE CONECTION IS BRACK. SO FIX THE WIRE AND SOLD RING AND MACHINE IS WORKING NOW AT PRESENT IN ME AND SISTER AND CROSS CHECK OTHER WEIGHING MACHINE SO PROPERLY WORKING. | Closed | 19DTT3491 |  | 1:25:0 | 1
818 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00683 | PHOTOTHERAPY UNIT | Jyoti chaudhari | 7588768076 | 41791 | 22-Apr-2020 | 09:23 |  |  | 02-May-2020 | 11:43 | NOT WORKING | CHECK AND FOUND ALL 15 BLUE LED BULB OF RATING 3.2V-3.6V AND 2-3 WATT IS FUSED. MR HAS BEEN SENT FOR NEW LED BULB BUT DUE TO LOCKDOWN VENDOR COULD NOT SUPPORT FOR DELIVERY OF REQUIRED SPARES. LOCKDOWN DUE TO COVID 19- NO SPARES/TRANSPORT/COURIER POSSIBLE. PRESENTLY CLOSING THE WO. WILL OPEN THE CALL AGAIN AND REPAIR THE ASSET ONCE THE SPARES WILL BE RECEIVED FROM VENDOR AFTER LOCKDOWN WILL BE OPENED BY GOVERNMENT. | Closed | NO INFO |  | 242:20:0 | 11
819 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00687 | PHOTOTHERAPY UNIT | Jyoti chaudhari | 7588768076 | 41736 | 20-Apr-2020 | 08:45 |  |  | 02-May-2020 | 10:30 | NOT WORKING | CHECK AND FOUND ALL 15 BLUE LED BULB OF RATING 3.2V - 3.6V & 2-3 WATT IS FUSED. MR HAS BEEN SENT FOR NEW LED BULB BUT DUE TO LOCKDOWN VENDOR COULD NOT SUPPORT FOR DELIVERY OF REQUIRED SPARES. LOCKDOWN DUE TO COVID 19- NO SPARES/TRANSPORT/COURIER POSSIBLE. PRESENTLY CLOSING THE WO. WILL OPEN THE CALL AGAIN & REPAIR THE ASSET ONCE THE SPARES WILL BE RECEIVED FROM VENDOR AFTER LOCKDOWN WILL BE OPENED BY GOVERNMENT. | Closed | NO INFO |  | 289:45:0 | 13
820 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00693 | PULSE OXIMETER LINE POWERED | Pratima R Meshram | 8668499481 | 42682 | 28-May-2020 | 09:05 |  |  | 28-May-2020 | 13:30 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS FAULTY AS SATURATION OF OXYGEN IN BLOOD COULD NOT BE SENSED. FIXED THE NEW NEONATAL 5 PIN, SINGLE LOCK SPO2 PROBE. READING OF OXYGEN SATURATION AND PULSE RATE IS OBSERVED AND FOUND TO BE ACCEPTABLE. PULSE OXIMETER IS NOW WORKING. | Closed | 18010100017 |  | 4:25:0 | 1
821 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00696 | SYRINGE PUMP | SCOTT NURSE | 9764971169 | 41515 | 07-Apr-2020 | 08:10 |  |  | 09-Apr-2020 | 11:39 | NOT WORKING | CHECK AND FOUND OCCLUSION ALARM OCCURS JUST AFTER STARTING THE DOSE AND THE PUMP STOPS. THERE IS ISSUE WITH OCCLUSION SENSOR, PLUNGER SET AND MOTOR. MR HAS BEEN RAISED FOR NEW OCCLUSION SENSOR, PLUNGER SET AND MOTOR. AS, SOON AS PARTS WILL BE RECEIVED IT WILL BE FIXED AND MACHINE WILL BE MADE FUNCTIONAL. | Closed | 090216B4 |  | 51:29:0 | 3
822 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00702 | NEBULIZER ELECTRIC | C S GADDAM | 9850236795 | 41137 | 20-Mar-2020 | 08:50 |  |  | 21-Mar-2020 | 15:16 | NOT WORKING | THIS CALL IS RAISED BY MISTAKENLY. USER WAS TRYING TO LOG CALL OF DIGITAL BP APPARATUS BUT GAVE THE ASSET NUMBER OF NEBULIZER MACHINE.HOWEVER, CHECKED THE NEBULIZER MACHINE. IT IS WORKING. | Closed | CN-001147 |  | 30:26:0 | 2
823 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00703 | DIGITAL BP APPARATUS | C S GADDAM | 9850236795 | 41138 | 20-Mar-2020 | 08:50 |  |  | 21-Mar-2020 | 11:48 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS HAVING LEAKAGE. FIXED THE NEW NIBP CUFF. DIGITAL BP APPARATUS IS NOW WORKING. | Closed | 1.5066803591e+014 |  | 26:58:0 | 2
824 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00718 | SUCTION PUMP PORTABLE, ELECTRONIC | Harsha Bhusari | 9096022114 | 42823 | 02-Jun-2020 | 09:15 |  |  | 02-Jun-2020 | 10:03 | NOT WORKING | CHECK AND FOUND POWER 'ON/OFF' SWITCH IS FAULTY AND DUE TO WHICH MACHINE IS NOT TURNING ON. FIXED THE NEW 4 PIN, 16 AMPERE POWER 'ON/OFF' SWITCH. POWER SUPPLY FAILURE ISSUE IS RESOLVED. SUCTION MACHINE IS NOW WORKING. | Closed | SL2018100009 |  | 0:48:0 | 1
825 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00718 | SUCTION PUMP PORTABLE, ELECTRONIC | Harsha Bhusari | 9096022114 | 42071 | 06-May-2020 | 09:10 |  |  | 06-May-2020 | 16:18 | NOT WORKING | CHECK AND FOUND INPUT TUBING CONNECTOR OF CONTAINER IS BROKEN.FIXED THE INPUT TUBING CONNECTOR. ALL TUBINGS ARE CLEANED BY DIPPING AND CLEANING IN HOT WATER. SUCTION PRESSURE IS CHECKED AND ITS OK. SUCTION MACHINE IS NOW WORKING. | Closed | SL2018100009 |  | 7:8:0 | 1
826 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00719 | CENTRIFUGE (BENCH TOP) | ARUN MULMULE | 9422144593 | 40776 | 12-Mar-2020 | 08:25 |  |  | 14-Mar-2020 | 15:00 | NOT WORKING | CHECK AND FOUND KNOB OF SPEED REGULATOR IS LOOSE. FIXED THE SPEED REGULATOR AND ARMATURE IS CLEANED. CENTRIFUGE MACHINE IS NOW WORKING WITH DESIRED RPM SPEED AT DIFFERENT SET RPM. | Closed | NO INFO |  | 54:35:0 | 3
827 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00721 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | A L Mulmule | 9422144593 | 42010 | 01-May-2020 | 08:10 |  |  | 02-May-2020 | 16:01 | NOT WORKING | CHECK AND FOUND ASPIRATION COULD NOT TAKE PLACE. CLEANED THE TUBINGS AND FLOW CELL. PROBLEM IS RESOLVED AND READING OBTAINED IS ACCEPTABLE. SEMI AUTOMATED BIOCHEMISTRY ANALYSER IS NOW WORKING. | Closed | NO INFO |  | 31:51:0 | 2
828 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00748 | DIALYSIS MACHINE | Hema Borjhar: | 9767558503 | 44276 | 16-Jul-2020 | 08:14 |  |  |  |  | NOT WORKING |  | Open | 8SXA0KDG
829 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00748 | DIALYSIS MACHINE | Hema Borjhar | 9767558503 | 42093 | 07-May-2020 | 09:12 |  |  | 07-May-2020 | 14:30 | NOT WORKING | V39 ALARM DURING LAST MINUTE OF RINSE AND HOT DISINFECTION. POSITIVE & NEGATIVE TEST FAILURE WITH HIGH CONDUCTIVITY ISSUE. CLEANED THE V39 VALVE AND V39 FAILURE ISSUE IS RESOLVED. TMP CALIBRATION ALONG WITH MIXING SYSTEM CALIBRATION IS DONE. POSITIVE AND NEGATIVE PRESSURE TEST IS PASSED AND CONDUCTIVITY ACHIEVED IS NORMAL. HOT DISINFECTION AND FRONT CLEANING IS GIVEN. DIALYSIS MACHINE IS NOW WORKING. | Closed | 8SXA0KDG |  | 5:18:0 | 1
830 | BHANDARA | DH BHANDARA | MH/R1/BHA/DH/BHA/00756 | PH METER | DR APARNA JAKKAL | 8888863689 | 40272 | 29-Feb-2020 | 09:08 |  |  | 29-Feb-2020 | 11:05 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE ISSUE. FIXED THE NEW FUSE AND REFIXED THE CONNECTION THROUGH SMPS BOARD. POWER SUPPLY FAILURE ISSUE RESOLVED. PH METER IS NOW WORKING. | Closed | 200804172 |  | 1:57:0 | 1
831 | BHANDARA | DTC BHANDARA | MH/R1/BHA/DTC/BHA/00002 | WATER BATH | RAJU KARANDE | 9423114657 | 4045 | 26-May-2017 | 09:30 |  |  | 30-May-2017 | 12:52 | NOT WORKING | REPLACE NEW SOCKET AND LEAKAGE PROBLEM WAS ARRESTED | Closed | 30876 | LOW | 99:22:0 | 5
832 | BHANDARA | DTC BHANDARA | MH/R1/BHA/DTC/BHA/00034 | SUCTION PUMP PORTABLE, ELECTRONIC | MINAKSHI VIRUTKAR | 8208050243 | 35773 | 26-Nov-2019 | 08:51 |  |  | 27-Nov-2019 | 16:55 | NOT WORKING | CHECK AND FOUND WINDING OF THE MOTOR IS SHORT AND HYDRAULIC VACUUM PUMP IS FAULTY AS PRESSURE COULD NOT BE GENERATED. NEW WINDING AND NEW SHEETS OF VACUUM PUMP IS FIXED THROUGH VENDOR. NOW MOTOR IS WORKING WITH DESIRED SUCTION PRESSURE. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 32:4:0 | 2
833 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00001 | BP APPARATUS ANEROID | SANDHYA UIKE | 8788308205 | 38857 | 01-Feb-2020 | 09:12 |  |  | 01-Feb-2020 | 10:02 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT. THERE IS LEAKAGE IN VALVE AND THE CONNECTING TUBE. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B.P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 0:50:0 | 1
834 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00001 | BP APPARATUS ANEROID | DR SHENDE | 9960495284 | 13271 | 27-Feb-2018 | 09:06 |  |  | 01-Mar-2018 | 11:54 | NOT WORKING | REPLACED BLADDER,CUFF, BULB WITH CONNECTOR. NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 50:48:0 | 3
835 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00001 | BP APPARATUS ANEROID | DR SANDIP JAGNADE | 9922811406 | 23951 | 15-Jan-2019 | 08:04 |  |  | 17-Jan-2019 | 12:16 | NOT WORKING | CHECK AND FOUND B.P APPRATUS BLADDER WAS LEAK. CHANGE THE BLADDER NOW MACHINE IS WORKING GOOD | Closed | NO INFO | LOW | 52:12:0 | 3
836 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00001 | BP APPARATUS ANEROID | DR LANGEWAR | 9423112300 | 9707 | 25-Nov-2017 | 09:20 |  |  | 27-Nov-2017 | 11:20 | NOT WORKING | REPLACED NEW BLADDER, BULB WITH CONNECTOR. NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 50:0:0 | 3
837 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00001 | BP APPARATUS ANEROID | DR MANOJ | 9960495284 | 16267 | 30-May-2018 | 08:50 |  |  | 05-Jun-2018 | 11:37 | NOT WORKING | REPLACED THE BLADDER | Closed | NO INFO | LOW | 146:47:0 | 7
838 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00011 | SUCTION PUMP PORTABLE, ELECTRONIC | DR LANGEWAR | 9423112300 | 9708 | 25-Nov-2017 | 09:20 |  |  | 27-Nov-2017 | 13:04 |  | ADJUST VACUUM CONTROL SECTION PROPERLY | Closed | NO INFO | MEDIUM | 51:44:0 | 3
839 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00012 | RADIANT WARMER | DR SANDIP JAGNADE | 9922811406 | 23949 | 15-Jan-2019 | 08:04 |  |  | 17-Jan-2019 | 12:13 | NOT WORKING | CHECK AND FOUND POWER SUPPLY IS LOSE. FIXED THE POWER SUPPLY PROPERLY AFTER FIXED THE POWER SUPPLY MACHINE IS WORKING | Closed | NO INFO | HIGH | 52:9:0 | 3
840 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00012 | RADIANT WARMER | SANDHYA UIKE | 8788308205 | 38856 | 01-Feb-2020 | 09:12 |  |  | 03-Feb-2020 | 10:00 | NOT WORKING | MACHINE CHECK AND FOUND LIGHT IS BURNT, SO REPLACE THE LIGHT AND NOW MACHINE IS WORKING PROPERLY. | Closed | NO INFO | HIGH | 48:48:0 | 3
841 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00015 | Fumigation Machine | DR SANDIP JAGNADE | 9922811406 | 23950 | 15-Jan-2019 | 08:04 |  |  | 21-Jan-2019 | 11:48 | NOT WORKING | CHECK AND FOUND MACHINE MOTOR IS BURNOUT NEED TO REWINDING THE MOTOR. MOTOR IS REWINDING NOW FUMIGATION MACHINE IS WORKING GOOD | Closed | NO INFO | LOW | 147:44:0 | 7
842 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00016 | Fumigation Machine | DR JAGNADE | 8459927636 | 22071 | 14-Nov-2018 | 08:25 |  |  | 17-Nov-2018 | 13:36 | NOT WORKING | CHECK AND FOUND POWER CABLE IS DAMAGE. CHANGED THE NEW POWER CABLE. FUMIGATION MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 77:11:0 | 4
843 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00018 | BP APPARATUS ANEROID | DR MANOJ | 9960495284 | 16265 | 30-May-2018 | 08:50 |  |  | 05-Jun-2018 | 11:11 | NOT WORKING | REPLACED WITH NEW BLADDER AND METAL CONNECTOR | Closed | NO INFO |  | 146:21:0 | 7
844 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00019 | BP APPARATUS ANEROID | DR MANOJ | 9960495284 | 16266 | 30-May-2018 | 08:50 |  |  | 05-Jun-2018 | 11:21 | NOT WORKING | REPLACED NEW METAL CONNECTOR | Closed | NO INFO |  | 146:31:0 | 7
845 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00020 | FETAL DOPPLER | DR LANGEWAR | 9423112300 | 9709 | 25-Nov-2017 | 09:20 |  |  | 27-Nov-2017 | 13:06 | NOT WORKING | SOLDERED SPEAKER SECTION PROPERLY,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO |  | 51:46:0 | 3
846 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00020 | FETAL DOPPLER | YEGESHWARI ALAMADE | 9763188039 | 19876 | 31-Aug-2018 | 09:16 |  |  | 03-Sep-2018 | 10:15 | NOT WORKING | CHECK AND FOUND POWER CORD IS FAULTY.CHANGED THE NEW POWER CORD.FOETAL DOPPLER IS NOW WORKING. | Closed | NO INFO |  | 72:59:0 | 4
847 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00022 | DIGITAL BP APPARATUS | SANDHYA UIKE | 8788308205 | 38852 | 01-Feb-2020 | 09:13 |  |  | 01-Feb-2020 | 10:45 | NOT WORKING | CHECK AND FOUND CUFF IS LEAK. FIXED THE NEW CUFF. DIGITAL B. P APPARATUS IS NOW WORKING. | Closed | 197 |  | 1:32:0 | 1
848 | BHANDARA | PHC ANDHALGAON | MH/R1/BHA/PHC/ADG/00026 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | SANDHYA UIKE | 8788308205 | 38855 | 01-Feb-2020 | 09:12 |  |  | 01-Feb-2020 | 10:00 | NOT WORKING | CHECK AND FOUND COIL IS BURNT AND DUE TO WHICH WATER COULD NOT BE BOILED. CHANGED THE HEATING COIL OF 3KW.AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO |  | 0:48:0 | 1
849 | BHANDARA | PHC ASGAON | MH/R1/BHA/PHC/AGN/00003 | BP APPARATUS ANEROID | DR.SHETTIWAR | 9764804796 | 33461 | 05-Oct-2019 | 09:25 |  |  | 05-Oct-2019 | 09:48 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS LEAK. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 98136204 | LOW | 0:23:0 | 1
850 | BHANDARA | PHC ASGAON | MH/R1/BHA/PHC/AGN/00013 | FETAL DOPPLER | ANISHA TEMBHURNE | 9890875367 | 12285 | 29-Jan-2018 | 09:07 |  |  | 01-Feb-2018 | 11:31 | NOT WORKING | REPLACED NEW 3 PIN FETAL DOPPLER PROBE, NOW MACHINE IS WORKING CONDITION | Closed | DFM2K804011/11 | MEDIUM | 74:24:0 | 4
851 | BHANDARA | PHC ASGAON | MH/R1/BHA/PHC/AGN/00019 | BP APPARATUS ANEROID | DR.SHETTIWAR | 9764804796 | 33462 | 05-Oct-2019 | 09:25 |  |  | 05-Oct-2019 | 09:46 | NOT WORKING | CHECK AND FOUND BLADDER, BULB, CUFF AND VALVE IS MISSING. FIXED THE NEW BLADDER, BULB, CUFF AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 0:21:0 | 1
852 | BHANDARA | PHC ASGAON | MH/R1/BHA/PHC/AGN/00023 | DIGITAL BP APPARATUS | Pravin | 9923634604 | 43829 | 04-Jul-2020 | 08:30 |  |  | 04-Jul-2020 | 10:00 | NOT WORKING | CHECK AND FOUND ERROR IS SHOWING WHILE INFLATING THE CUFF. B.P.CUFF IS FOUND LEAK. REPLACED WITH NEW ONE. NOW MACHINE IS WORKING. | Closed | NO INFO |  | 1:30:0 | 1
853 | BHANDARA | PHC ASGAON | MH/R1/BHA/PHC/AGN/00023 | DIGITAL BP APPARATUS | DR SETTIWAL | 9764804796 | 35661 | 23-Nov-2019 | 09:09 |  |  | 24-Nov-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN CUFF. CHANGE THE NEW CUFF NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER | Closed | NO INFO |  | 26:51:0 | 2
854 | BHANDARA | PHC ASGAON | MH/R1/BHA/PHC/AGN/00024 | BP APPARATUS ANEROID | DR.SHETTIWAR | 9764804796 | 33463 | 05-Oct-2019 | 09:25 |  |  | 05-Oct-2019 | 09:43 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT. FIXED THE NEW CUFF. B.P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 0:18:0 | 1
855 | BHANDARA | PHC BARWHA | MH/R1/BHA/PHC/BRW/00002 | BP APPARATUS ANEROID | KALPANA | 8605481718 | 24915 | 12-Feb-2019 | 09:20 |  |  | 13-Feb-2019 | 15:15 | NOT WORKING | CHECK AND FOUND BLADDER, BULB AND VALVE IS LEAK. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW PARTS. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 29:55:0 | 2
856 | BHANDARA | PHC BARWHA | MH/R1/BHA/PHC/BRW/00004 | BINOCULAR MICROSCOPE |  |  | 8523 | 30-Oct-2017 | 09:26 |  |  | 30-Oct-2017 | 20:06 | NOT WORKING | REPLACED THE POWER CARD | Closed | NO INFO | MEDIUM | 10:40:0 | 1
857 | BHANDARA | PHC BARWHA | MH/R1/BHA/PHC/BRW/00007 | BP APPARATUS ANEROID | KALPANA | 8605481718 | 24916 | 12-Feb-2019 | 09:20 |  |  | 13-Feb-2019 | 17:18 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT AND BULB, VALVE, BLADDER IS LEAK. FIXED THE NEW CUFF,BLADDER,VALVE, BULB. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 31:58:0 | 2
858 | BHANDARA | PHC BARWHA | MH/R1/BHA/PHC/BRW/00021 | LOW COST GLUCOMETER | A G SATARKAR | 9420620227 | 15568 | 04-May-2018 | 09:00 |  |  | 04-May-2018 | 16:56 | NOT WORKING | LITHIUM BATTERY WAS OUR OF CHARGE REPLACED NEW BATTERY | Closed | 55311791984 | MEDIUM | 7:56:0 | 1
859 | BHANDARA | PHC BARWHA | MH/R1/BHA/PHC/BRW/00023 | BP APPARATUS ANEROID | A G SATARKAR | 9420620227 | 15567 | 04-May-2018 | 09:00 |  |  | 04-May-2018 | 16:50 | NOT WORKING | REPLACED WITH NEW GLASS TUBE AND FILLED MERCURY | Closed | NO INFO |  | 7:50:0 | 1
860 | BHANDARA | PHC BARWHA | MH/R1/BHA/PHC/BRW/00023 | BP APPARATUS ANEROID | KALPANA | 8605481718 | 24917 | 12-Feb-2019 | 09:20 |  |  | 13-Feb-2019 | 14:11 | NOT WORKING | CHECK AND FOUND CUFF, BLADDER, VALVE AND BULB IS MISSING. FIXED THE NEW CUFF, BLADDER, VALVE AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 28:51:0 | 2
861 | BHANDARA | PHC BARWHA | MH/R1/BHA/PHC/BRW/00024 | BP APPARATUS ANEROID | KALPANA | 8605481718 | 24918 | 12-Feb-2019 | 09:20 |  |  | 13-Feb-2019 | 16:21 | NOT WORKING | CHECK AND FOUND BLADDER AND VALVE IS LEAK. CHANGED THE OLD DEFECTIVE PARTS WITH NEW PARTS. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 31:1:0 | 2
862 | BHANDARA | PHC BARWHA | MH/R1/BHA/PHC/BRW/00027 | BP APPARATUS ANEROID | KALPANA | 8605481718 | 29514 | 04-Jul-2019 | 09:21 |  |  | 08-Jul-2019 | 13:55 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF, BULB AND VALVE IS MISSING. FIXED THE NEW BLADDER, CUFF, BULB, VALVE AND CONNECTING TUBE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 100:34:0 | 5
863 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00001 | BP APPARATUS ANEROID | DR SHENDE | 9960495284 | 8105 | 14-Oct-2017 | 09:02 |  |  | 16-Oct-2017 | 20:35 | NOT WOKRING | REPLACE THE BLADDER | Closed | RM0105693 | LOW | 59:33:0 | 3
864 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00001 | BP APPARATUS ANEROID | S M KAVLE | 9422153049 | 11313 | 04-Jan-2018 | 10:10 |  |  | 05-Jan-2018 | 13:04 | NOT WORKING | REPLACED THE BLADDER,TUBING, TUBING CONNECTOR AND CUFF | Closed | RM0105693 | LOW | 26:54:0 | 2
865 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00001 | BP APPARATUS ANEROID | DR SHENDE | 9960495284 | 28356 | 08-Jun-2019 | 08:15 |  |  | 11-Jun-2019 | 16:31 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND EXTENSION TUBE. CHANGED THE OLD DEFECTIVE PART WITH THE NEW. B. P APPARATUS IN NOW WORKING. | Closed | RM0105693 | LOW | 80:16:0 | 4
866 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00001 | BP APPARATUS ANEROID | DR MANOJ SHINDE | 9960495284 | 18765 | 03-Aug-2018 | 08:07 |  |  | 10-Aug-2018 | 09:30 | NOT WORKING | CHECK AND FOUND BP BLADDER LEAKAGE. CUFF IS WORN OUT BULB IS ALSO WORN OUT CHANGED THE NEW CUFF BLADDER VALVE AND BULB BP APPRATUS NOW WORKING | Closed | RM0105693 | LOW | 169:23:0 | 8
867 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00002 | BP APPARATUS ANEROID | SMITA NURSE | 9422153049 | 25999 | 18-Mar-2019 | 09:26 |  |  | 18-Mar-2019 | 21:46 | NOT WOKRING | CHECK AND FOUND CONNECTING TUBE IS DEGRADED AND CONNECTOR IS ALSO BROKEN. CHANGED THE NEW CONNECTING TUBE AND CONNECTOR. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 12:20:0 | 1
868 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00004 | ADULT WEIGHING MACHINE | SMITA KALE | 9422153049 | 16715 | 14-Jun-2018 | 08:42 |  |  | 15-Jun-2018 | 15:07 | NOT WORKING | POWER SUPPLY BOARD NOT FOUND IN THE MACHINE | Closed | NO INFO | LOW | 30:25:0 | 2
869 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00006 | NEBULIZER ELECTRIC | DR SHENDE | 9960495284 | 20666 | 25-Sep-2018 | 08:17 |  |  | 26-Sep-2018 | 12:03 | NOT WORKING | CHECK AND FOUND NEBULISER FAN IS NOT ROTATING PROPERLY FAN CLEAN AND FIXED PROPERLY.NOW MACHINE IS WORKING GOOD | Closed | NO INFO | MEDIUM | 27:46:0 | 2
870 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00006 | NEBULIZER ELECTRIC | S M KAVLE | 9422153049 | 23020 | 15-Dec-2018 | 08:20 |  |  | 18-Dec-2018 | 10:36 | NOT WORKING | CHECK AND FOUND MOTOR IS JAMMED AND ITS NOT ROTATING. OILING OF MOTOR IS DONE.NOW THE MOTOR IS ROTATING NORMALLY. NEBULIZER MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 74:16:0 | 4
871 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00006 | NEBULIZER ELECTRIC | SMITA KALE | 9422153049 | 11133 | 30-Dec-2017 | 09:05 |  |  | 04-Jan-2018 | 11:36 | NOT WORKING | SOLDERING THE WIRE CONNECTION | Closed | NO INFO | MEDIUM | 122:31:0 | 6
872 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00008 | BP APPARATUS ANEROID | S M KAVLE | 9422153049 | 11314 | 04-Jan-2018 | 10:10 |  |  | 05-Jan-2018 | 13:08 | NOT WORKING | REPLACED THE TUBING AND BULB | Closed | NO INFO | LOW | 26:58:0 | 2
873 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00008 | BP APPARATUS ANEROID | DR SHENDE | 9960495284 | 28357 | 08-Jun-2019 | 08:20 |  |  | 11-Jun-2019 | 16:42 | NOT WORKING | CHECK AND FOUND LEAKAGE IN EXTENSION TUBE. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 80:22:0 | 4
874 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00011 | CENTRIFUGE (BENCH TOP) | NEETA KAVLE | 9422153049 | 23690 | 07-Jan-2019 | 09:15 |  |  | 08-Jan-2019 | 12:38 | NOT WORKING | CHECK AND FOUND CENTRIFUGE MACHINE CARBON BRUSH IS TOTALLY END. INSTALL THE NEW CARBON BRUSH. NOW MACHINE IS WORKING GOOD. | Closed | NO INFO | MEDIUM | 27:23:0 | 2
875 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00012 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | SMITA NURSE | 9422153049 | 30945 | 07-Aug-2019 | 08:35 |  |  | 12-Aug-2019 | 12:00 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT AND ALSO POWER CABLE IS MISSING. FIXED THE NEW HEATING COIL OF 2KW AND ALSO THE POWER CABLE. VERTICAL AUTO CLAVE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 123:25:0 | 6
876 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00019 | Fumigation Machine | DR D K LAKSHTIVAR | 9860954196 | 20756 | 26-Sep-2018 | 08:40 |  |  | 03-Oct-2018 | 11:10 | NOT WORKING | CHECK AND FOUND FUMIGATION MOTOR IS BARN OUT SO WE WILL REWINDING THIS MOTOR NOW MACHINE IS WORKING | Closed | NO INFO | LOW | 170:30:0 | 8
877 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00019 | Fumigation Machine | SITA KALAVE | 9422153049 | 17032 | 22-Jun-2018 | 08:17 |  |  | 09-Jul-2018 | 12:30 | NOT WORKING | CHECK AND FOUND PROBLEM REPORTED IN MOTOR . REPAIR MOTOR FROM LOCAL VENDOR . AFTER REPAIR THE MACHINE WORKING FINE TESTED OK | Closed | NO INFO | LOW | 412:13:0 | 18
878 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00020 | SHADOWLESS LAMP STAND TYPE | DR SHENDE | 9960495284 | 8104 | 14-Oct-2017 | 09:02 |  |  | 16-Oct-2017 | 20:37 | NOT WORKING | REPLACE THE 12V/50W BULB AND HOLDER | Closed | NO INFO | MEDIUM | 59:35:0 | 3
879 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00022 | FETAL DOPPLER | SMITA | 9422153049 | 34455 | 30-Oct-2019 | 08:13 |  |  | 31-Oct-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BATTERY CELL AND INTERNAL POWER CONNECTION. MAKE THE POWER CONNECTION AND CHANGE NEW BATTERY CELL NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | 0912016FQ0658 |  | 32:47:0 | 2
880 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00028 | ADULT WEIGHING MACHINE | SMITA KALVALE | 9422153049 | 19101 | 13-Aug-2018 | 09:42 |  |  | 16-Aug-2018 | 11:21 | NOT WORKING | CHECK AND FOUND POWER SUPPLY OF ADULT WEIGHINGMACHINE WAS MISSING FIXED NEW POWER SUPPLY NOW ADULT WEIGHING MACHINE IS WORKING | Closed | NO INFO |  | 73:39:0 | 4
881 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00028 | ADULT WEIGHING MACHINE | SMITA KALE | 9422153049 | 16716 | 14-Jun-2018 | 08:45 |  |  | 15-Jun-2018 | 03:22 | NOT WORKING | LOCATED THE MARKER IN ZERO POSITION. NOW MACHINE WORKING | Closed | NO INFO |  | 18:37:0 | 2
882 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00037 | BP APPARATUS ANEROID | SMITA NURSE | 9422153049 | 26000 | 18-Mar-2019 | 09:26 |  |  | 18-Mar-2019 | 21:47 | NOT WORKING | CHECK AND FOUND BULB IS LEAK. CHANGED THE NEW BULB. B. P APPARATUS IS NOW WORKING. | Closed | 91835311 |  | 12:21:0 | 1
883 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00037 | BP APPARATUS ANEROID | SMITA NURSE | 9422153049 | 30944 | 07-Aug-2019 | 08:35 |  |  | 08-Aug-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER AND VALVE. CHANGE THE NEW VALVE NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | 91835311 |  | 27:25:0 | 2
884 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00037 | BP APPARATUS ANEROID | S M KAVLE | 9422153049 | 23021 | 15-Dec-2018 | 08:20 |  |  | 18-Dec-2018 | 11:38 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK AND CUFF IS WORN OUT. CHANGED THE NEW BLADDER AND CUFF. B. P APPARATUS IS NOW WORKING. | Closed | 91835311 |  | 75:18:0 | 4
885 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00041 | BP APPARATUS ANEROID | S M KAVLE | 9422153049 | 23022 | 15-Dec-2018 | 08:20 |  |  | 18-Dec-2018 | 12:41 | NOT WORKING | CHECK AND FOUND CUFF, BLADDER, VALVE AND BULB IS MISSING. FIXED THE NEW CUFF, BLADDER, VALVE AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | 17M0120 |  | 76:21:0 | 4
886 | BHANDARA | PHC BETALA | MH/R1/BHA/PHC/BTL/00042 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | SMITA NURSE | 9422153049 | 30946 | 07-Aug-2019 | 08:35 |  |  | 12-Aug-2019 | 12:00 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT. FIXED THE NEW HEATING COIL OF 2 KW. VERTICAL AUTO CLAVE. MACHINE IS NOW WORKING. | Closed | NO INFO |  | 123:25:0 | 6
887 | BHANDARA | PHC BHUYAR | MH/R1/BHA/PHC/BUR/00004 | BP APPARATUS ANEROID | DR KAIKADE | 9423115444 | 35354 | 19-Nov-2019 | 08:15 |  |  | 21-Nov-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS AIR LEAK IN BLADDER. CHANGE THE NEW BLADDER NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | LOW | 51:45:0 | 3
888 | BHANDARA | PHC BHUYAR | MH/R1/BHA/PHC/BUR/00006 | NEBULIZER ELECTRIC | DR KAIKADE | 9423115444 | 35356 | 19-Nov-2019 | 08:15 |  |  | 21-Nov-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN POWER CABLE. SHORT OUT THIS PROBLEM NOW MACHINE IS IN WORKING CONDITION. | Closed | NUM681519 | MEDIUM | 51:45:0 | 3
889 | BHANDARA | PHC BHUYAR | MH/R1/BHA/PHC/BUR/00013 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | PRITEE CHUTE | 9822852880 | 36620 | 17-Dec-2019 | 08:11 |  |  | 18-Dec-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN POWER CABLE IS FAULTY. THEN CHANGE THE NEW POWER CABLE NOW MACHINE IS IN WORKING CONDITION | Closed | NO INFO | MEDIUM | 27:49:0 | 2
890 | BHANDARA | PHC BHUYAR | MH/R1/BHA/PHC/BUR/00017 | SUCTION PUMP PORTABLE, ELECTRONIC | MESHRAM | 8007752059 | 24710 | 06-Feb-2019 | 09:13 |  |  | 07-Feb-2019 | 13:29 | NOT WORKING | CHECK AND FOUND INTERNAL TUBING ACROSS MOTOR IS OPENED. ALSO PRESSURE GAUGE IS NOT WORKING. FIXED THE INTERNAL TUBING ACROSS MOTOR AND PRESSURE GAUGE IS REPAIRED. NOW, SUCTION PRESSURE IS GENERATED WITH PRESSURE GAUGE DEPICTING CORRECT PRESSURE. | Closed | NO INFO | MEDIUM | 28:16:0 | 2
891 | BHANDARA | PHC BHUYAR | MH/R1/BHA/PHC/BUR/00017 | SUCTION PUMP PORTABLE, ELECTRONIC | PRITEE | 9822852880 | 38416 | 23-Jan-2020 | 09:53 |  |  | 23-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN MOTOR IS SHORT CIRCUITED THEN MAKE NEW WINDING NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 7:7:0 | 1
892 | BHANDARA | PHC BHUYAR | MH/R1/BHA/PHC/BUR/00027 | DIGITAL BP APPARATUS | DR KAIKADE | 9423115444 | 35355 | 19-Nov-2019 | 08:15 |  |  | 21-Nov-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS BATTERY ISSUE. CONNECT THE NEW BATTERY CELL NOW MACHINE IS WORKING CONDITION HAND OVER TO USER. | Closed | 20160610471VGI |  | 51:45:0 | 3
893 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00003 | NEBULIZER ELECTRIC | MOHAN RAKADE | 9021215950 | 25634 | 07-Mar-2019 | 09:05 |  |  | 08-Mar-2019 | 11:09 | NOT WORKING | CHECK AND FOUND POWER ON/OFF SWITCH IS FAULTY. FIXED THE NEW ON/OFF SWITCH. NEBULISER MACHINE IS NOW WORKING | Closed | PCB77694 | MEDIUM | 26:4:0 | 2
894 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00006 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | DR CHOUHAN | 9604807659 | 9436 | 21-Nov-2017 | 09:05 |  |  | 24-Nov-2017 | 09:30 | NOT WORKING | REPLACED NEW 6V BATTERY AND SOLDERED VARIABLE RESISTOR ON MICROCONTROLLER BOARD,NOW MACHINE IS WORKING CONDITION | Closed | 9091682 | MEDIUM | 72:25:0 | 4
895 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00006 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | DR SACHIN CHAVAN | 9604867698 | 4328 | 06-Jun-2017 | 01:00 | 08-Jun-2017 | 11:57 | 09-Jun-2017 | 12:30 | NOT WORKING | CALIBRATION PROCESS DONE PROPERLY,NOW WEIGHT SHOWING PROPERLY AND MACHINE IS WORKING GOOD CONDITION. | Closed | 9091682 | MEDIUM | 83:30:0 | 4
896 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00006 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | DR CHAVAN | 9604807698 | 7224 | 18-Sep-2017 | 08:35 |  |  | 19-Sep-2017 | 09:30 | NOT WORKING | REPLACED NEW 6V 4.5AH BATTERY, NOW ERROR PROBLEM GET RECTIFIED,MACHINE IS WORKING CONDITION. | Closed | 9091682 | MEDIUM | 24:55:0 | 2
897 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00009 | ADULT WEIGHING MACHINE | MOHAN RAKADE | 9021215950 | 25636 | 07-Mar-2019 | 09:05 |  |  | 08-Mar-2019 | 11:02 | NOT WORKING | CHECK AND FOUND MACHINE POWER CORD IS BURNT OUT. FIXED THE NEW POWER CORD. ADULT WEIGHING MACHINE IS NOW WORKING | Closed | NO INFO |  | 25:57:0 | 2
898 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00010 | NEBULIZER ELECTRIC | MOHAN RAKADE | 9021215950 | 25633 | 07-Mar-2019 | 09:05 |  |  | 08-Mar-2019 | 11:06 | NOT WORKING | CHECK AND FOUND ON/OFF SWITCH FAULTY. REPLACE THE POWER CORD AND FIXED THE NEW ON /OFF SWITCH. NOW NEBULISER MACHINE IS WORKING | Closed | NO INFO |  | 26:1:0 | 2
899 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00016 | HB METER | MOHAN RAKHADE | 9021215950 | 27898 | 24-May-2019 | 09:05 |  |  | 29-May-2019 | 16:17 | NOT WORKING | CHECK AND FOUND ERROR 02 IS DISPLAYED ON SCREEN. CLEANED THE OPTRONIC SENSOR AND ERROR 02 IS ELEMINATED. HB METER IS NOW WORKING. | Closed | 1414013401 |  | 127:12:0 | 6
900 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00021 | BP APPARATUS ANEROID | PATALE | 9422587160 | 25020 | 14-Feb-2019 | 09:03 |  |  | 18-Feb-2019 | 17:32 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS LEAK. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 104:29:0 | 5
901 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00040 | DIGITAL BP APPARATUS | MOHAN RAKADE | 9975034610 | 22751 | 08-Dec-2018 | 09:35 |  |  | 14-Dec-2018 | 11:05 | READING ISSUE | CHECK AND FOUND DIGITAL BP APPRATUS NIBP IS DEFECTIVE. CHANGE THE SAME WITH NEW ONE NOW MACHINE IS WORKING | Closed | NA |  | 145:30:0 | 7
902 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00040 | DIGITAL BP APPARATUS | PATALE | 9422587160 | 25021 | 14-Feb-2019 | 09:03 |  |  | 19-Feb-2019 | 16:24 | NOT WORKING | CHECK AND FOUND MERCURY IS SPILLS OUT. ALSO BLADDER IS LEAK. FIXED THE BLADDER AND REFILLED THE MERCURY. B. P APPARATUS IS NOW WORKING. | Closed | NA |  | 127:21:0 | 6
903 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00040 | DIGITAL BP APPARATUS | DR TRISHNA | 9421715677 | 38814 | 31-Jan-2020 | 09:18 |  |  | 31-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN ADAPTER PORT . SHORT OUT THIS PROBLEM NOW MACHINE IS IN WORKING CONDITION. | Closed | NA |  | 7:42:0 | 1
904 | BHANDARA | PHC CHULHAD | MH/R1/BHA/PHC/CLD/00046 | BP APPARATUS ANEROID | PATALE | 9422587160 | 25022 | 14-Feb-2019 | 09:03 |  |  | 18-Feb-2019 | 19:34 | NOT WORKING | CHECK AND FOUND BLADDER,BULB AND VALVE IS LEAK. FIXED THE NEW BLADDER, BULB AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NA |  | 106:31:0 | 5
905 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00001 | RADIANT WARMER | DR KAIKADE | 9057580945 | 5581 | 24-Jul-2017 | 09:10 |  |  | 27-Jul-2017 | 11:19 | NOT WORKING | REPLACE 3PIN TEMPERATURE PROBE | Closed | NO INFO | HIGH | 74:9:0 | 4
906 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00003 | SUCTION PUMP PORTABLE, ELECTRONIC | DR KAIKADE | 7057580945 | 12518 | 03-Feb-2018 | 09:15 |  |  | 03-Feb-2018 | 19:28 | NOT WORKING | RECONNECT THE TUBES | Closed | NO INFO | MEDIUM | 10:13:0 | 1
907 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00003 | SUCTION PUMP PORTABLE, ELECTRONIC | DAHIWALE | 8275117813 | 1770 | 20-Feb-2017 | 11:00 |  |  | 22-Feb-2017 | 14:30 | NOT WORKING | REFIXED THE TUBING CONNECTION | Closed | NO INFO | MEDIUM | 51:30:0 | 3
908 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00003 | SUCTION PUMP PORTABLE, ELECTRONIC | ANITA | 9673444543 | 28375 | 10-Jun-2019 | 08:15 |  |  | 14-Jun-2019 | 15:02 | NOT WORKING | CHECK AND FOUND SUCTION OUTLET FILTER IS BROKEN. REPAIRED THE SAME. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 102:47:0 | 5
909 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00007 | Fumigation Machine | DAHIWALE | 8225117813 | 2864 | 31-Mar-2017 | 09:25 |  |  | 03-Apr-2017 | 11:30 | NOT WORKING | REPLACED BEARING AND FAN,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | LOW | 74:5:0 | 4
910 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00016 | BINOCULAR MICROSCOPE | DR KAIKADE | 94233115444 | 5318 | 14-Jul-2017 | 09:20 |  |  | 15-Jul-2017 | 11:30 | NOT WORKING | REARRANGED THE SLIDE FIXER ARRANGEMENT, NOW MACHINE IS WORKING CONDITION. | Closed | 08-J-8281 | MEDIUM | 26:10:0 | 2
911 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00016 | BINOCULAR MICROSCOPE | MAHADEV | 8405897090 | 29359 | 02-Jul-2019 | 08:14 |  |  | 04-Jul-2019 | 10:43 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. CHANGED THE NEW HALOGEN BULB OF RATING 12V,50W.BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | 08-J-8281 | MEDIUM | 50:29:0 | 3
912 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00016 | BINOCULAR MICROSCOPE | SHIVANKAR YOGRAJ | 9405987090 | 40556 | 06-Mar-2020 | 09:10 |  |  | 08-Mar-2020 | 10:00 | NOT WORKING | CHECK AND FOUND TRANSFORMER IS FAULTY. REPLACE THE 6V-3A TRANSFORMER AND FIX THE HALLOGEN BULB IN HOLDER. NOW MACHINE IS WORKING. | Closed | 08-J-8281 | MEDIUM | 48:50:0 | 3
913 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00016 | BINOCULAR MICROSCOPE | YOGRAJ SHIVANKAR | 7066342042 | 6508 | 28-Aug-2017 | 09:15 | 30-Aug-2017 | 19:44 | 29-Aug-2017 | 11:59 | NOT WORKING | REPLACE THE 6V TRANSFORMER | Closed | 08-J-8281 | MEDIUM | 26:44:0 | 2
914 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00016 | BINOCULAR MICROSCOPE | DR KAIKADE | 7057580945 | 12519 | 03-Feb-2018 | 09:15 |  |  | 03-Feb-2018 | 14:06 | NOT WORKING | REPLACED THE LAMP AND TRIGGERING PCB | Closed | 08-J-8281 | MEDIUM | 4:51:0 | 1
915 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00019 | BP APPARATUS ANEROID | DR SHANKAR | 9423115444 | 9188 | 16-Nov-2017 | 09:10 |  |  | 17-Nov-2017 | 16:13 | NOT WORKING | REPLACED BLADDER AND METAL CONNECTOR | Closed | NO INFO | LOW | 31:3:0 | 2
916 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00019 | BP APPARATUS ANEROID | DR KAIKADE | 94233115444 | 5317 | 14-Jul-2017 | 09:20 |  |  | 15-Jul-2017 | 11:30 | NOT WORKING | REPLACED NEW BLADDER,CUFF,BULB WITH CONNECTOR . NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | LOW | 26:10:0 | 2
917 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00019 | BP APPARATUS ANEROID | DR KAIKADE | 9423115444 | 17398 | 02-Jul-2018 | 10:49 |  |  | 06-Jul-2018 | 15:49 | NOT WORKING | CHECK THE MACHINE PROBLEM REPORTED IN B.P. BULB OF MACHINE, REPLACED IT WITH NEW ONE NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | LOW | 101:0:0 | 5
918 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00019 | BP APPARATUS ANEROID | ANITA MARGALE | 9673444543 | 28596 | 14-Jun-2019 | 09:25 |  |  | 14-Jun-2019 | 16:31 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK AND EXTENSION CORD IS DEGRADED. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 7:6:0 | 1
919 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00019 | BP APPARATUS ANEROID | KISHOR DAHIWALE | 8275117813 | 918 | 12-Jan-2017 | 12:11 |  |  | 16-Jan-2017 | 09:35 | NOT WORKING | BLADDER AND BULB REPLACED | Closed | NO INFO | LOW | 93:24:0 | 5
920 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00020 | BP APPARATUS ANEROID | KISHOR DAHIWALE | 8275117813 | 919 | 12-Jan-2017 | 12:00 |  |  | 16-Jan-2017 | 09:40 | NOT WORKING | BLADDER AND CUFF REPLACED | Closed | 1978 | LOW | 93:40:0 | 5
921 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00020 | BP APPARATUS ANEROID | DR KAIKADE | 9423115444 | 17397 | 02-Jul-2018 | 10:42 |  |  | 06-Jul-2018 | 15:45 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT. CHANGE WITH NEW ONE . NOW B.P. APRATUS WORKING SATISFACTORY. | Closed | 1978 | LOW | 101:3:0 | 5
922 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00020 | BP APPARATUS ANEROID | DR KAIKADE | 9423115444 | 5165 | 10-Jul-2017 | 09:02 |  |  | 11-Jul-2017 | 15:15 | NOT WORKING | REPLACED NEW BLADDER ,BULB WITH CONNECTOR,NOW MACHINE IS WORKING CONDITION | Closed | 1978 | LOW | 30:13:0 | 2
923 | BHANDARA | PHC DIGHORI | MH/R1/BHA/PHC/DIG/00039 | NEBULIZER ELECTRIC | DR KAIKADE | 9423115444 | 24312 | 25-Jan-2019 | 09:07 |  |  | 28-Jan-2019 | 15:30 | NOT WORKING | CHECK AND FOUND TUBING IS FAULTY. AS IT IS CONSUMABLE ITEM SO IT IS NOT PROVIDE BY THE COMPANY. SO, HOSPITAL AUTHORITY HAS TO ARRANGE THE TUBING AND MASK. REST, MACHINE IS WORKING. | Closed | NO INFO |  | 78:23:0 | 4
924 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00001 | BP APPARATUS ANEROID | P D SAHARE | 9421816679 | 13184 | 23-Feb-2018 | 10:10 |  |  | 23-Feb-2018 | 15:26 | NOT WORKING | REPLACED THE METAL VALVE | Closed | NO INFO | LOW | 5:16:0 | 1
925 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00001 | BP APPARATUS ANEROID | G T REHPADE | 9423413779 | 16226 | 29-May-2018 | 08:10 |  |  | 04-Jun-2018 | 09:29 | NOT WORKING | REPLACED WITH NEW BLADDER AND CUFF. NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO | LOW | 145:19:0 | 7
926 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00002 | ADULT WEIGHING MACHINE | G T REHPADE | 9423413779 | 16228 | 29-May-2018 | 08:14 |  |  | 04-Jun-2018 | 09:25 | NOT WORKING | CHANGED THE 3AMP FUSE . NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO | LOW | 145:11:0 | 7
927 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00007 | BP APPARATUS ANEROID | P D SAHARE | 9421816679 | 13185 | 23-Feb-2018 | 10:10 |  |  | 23-Feb-2018 | 15:23 | NOT WORKING | REPLACED THE METAL VALVE, BULB | Closed | 585 | LOW | 5:13:0 | 1
928 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00010 | BP APPARATUS ANEROID | G T REHPADE | 9423413779 | 16227 | 29-May-2018 | 08:12 |  |  | 04-Jun-2018 | 09:41 | NOT WORKING | REPLACED WITH NEW BLADDER, METAL CONNECTOR, BULB. NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO | LOW | 145:29:0 | 7
929 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00010 | BP APPARATUS ANEROID | DR SAVITA MALDONGRE | 8806681174 | 5936 | 05-Aug-2017 | 10:32 |  |  | 05-Aug-2017 | 17:37 | NOT WORKING | REPLACED THE BLADDER | Closed | NO INFO | LOW | 7:5:0 | 1
930 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00010 | BP APPARATUS ANEROID | J J REHPADE | 9423413779 | 11506 | 10-Jan-2018 | 09:21 |  |  | 11-Jan-2018 | 11:36 | NOT WORKING | REPLACED THE BULB AND METAL CONNECTOR | Closed | NO INFO | LOW | 26:15:0 | 2
931 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00011 | NEBULIZER ELECTRIC | VIJAYA SHENDE | 9765035530 | 33184 | 28-Sep-2019 | 09:08 |  |  | 29-Sep-2019 | 11:30 | NOT WORKING | CHECK AND FOUND ON/OFF SWITCH IS FAULTY AND DUE TO WHICH MACHINE IS NOT TURNING ON. CHANGED THE DEFECTIVE ON/OFF SWITCH WITH THE NEW ONE. MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 26:22:0 | 2
932 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00018 | BINOCULAR MICROSCOPE | P D SAHARE | 9421816679 | 13186 | 23-Feb-2018 | 10:10 |  |  | 23-Feb-2018 | 15:20 | NOT WORKING | REPLACED THE 6V 20 WITH LAMP | Closed | NO INFO | MEDIUM | 5:10:0 | 1
933 | BHANDARA | PHC DHARGAON | MH/R1/BHA/PHC/DRG/00020 | CENTRIFUGE (BENCH TOP) | P R ANDEL | 9049217695 | 13126 | 22-Feb-2018 | 09:16 |  |  | 26-Feb-2018 | 12:02 | NOT WORKING | CHECKED AND FOUND TEST TUBE HOLDER GET DAMAGED AND MOTOR ALSO DAMAGED COMPLETELY, IT REQUIRED EXTENSIVE REPAIR,SO WE RECOMMENDED FOR CONDEMNED. | Closed | TC48020 | MEDIUM | 98:46:0 | 5
934 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00001 | BP APPARATUS ANEROID | MINAKSHI NURSE | 800764216 | 23648 | 05-Jan-2019 | 09:10 |  |  | 15-Jan-2019 | 14:24 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS LEAK. CHANGED THE NEW BLADDER,VALVE AND BULB. B. P APPRATUS IS NOW WORKING. | Closed | 7773 | LOW | 245:14:0 | 11
935 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00002 | BP APPARATUS ANEROID | NITIN VAIRAGADKAR | 9403801044 | 13879 | 16-Mar-2018 | 09:20 |  |  | 19-Mar-2018 | 14:42 | NOT WORKING | REPLACED NEW BLADDER, BULB WITH CONNECTOR NOW MACHINE IS WORKING CONDITION. | Closed | 2542 | LOW | 77:22:0 | 4
936 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00002 | BP APPARATUS ANEROID | NITIN | 9403801044 | 10339 | 11-Dec-2017 | 10:15 |  |  | 15-Dec-2017 | 22:46 | NOT WORKING | REPLACED THE BLADDER, METAL VALVE, BULB | Closed | 2542 | LOW | 108:31:0 | 5
937 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00002 | BP APPARATUS ANEROID | NITIN VAIRAGALEKAR | 9403801044 | 18319 | 23-Jul-2018 | 08:37 |  |  | 25-Jul-2018 | 16:30 | NOT WORKING | CHECK THE MACHINE, FOUND BLADDER,CUFF,BULB FAULTY. REPLACED SAME. NOW MACHINE WORKING SATISFACTORY. | Closed | 2542 | LOW | 55:53:0 | 3
938 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00002 | BP APPARATUS ANEROID | NITIN WARALEKAR | 9403801044 | 16727 | 14-Jun-2018 | 09:04 |  |  | 19-Jun-2018 | 11:02 | NOT WORKING | REPLACED THE METAL CONNECTOR | Closed | 2542 | LOW | 121:58:0 | 6
939 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00002 | BP APPARATUS ANEROID | DR ANSARI | 7620409720 | 19033 | 10-Aug-2018 | 09:51 |  |  | 11-Aug-2018 | 10:15 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK. CUFF IS ALSO WORN OUT. CHANGED THE NEW BLADDER AND CUFF. B. P INSTRUMENT IS NOW WORKING. | Closed | 2542 | LOW | 24:24:0 | 2
940 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00008 | ADULT WEIGHING MACHINE | NITIN VAIRAGADKAR | 9403801044 | 13881 | 16-Mar-2018 | 09:20 |  |  | 19-Mar-2018 | 13:49 | NOT WORKING | REPLACED NEW POWER WIRE 6A, MACHINE IS WORKING CONDITION | Closed | NO INFO | LOW | 76:29:0 | 4
941 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00008 | ADULT WEIGHING MACHINE | NITIN | 9403801044 | 10338 | 11-Dec-2017 | 10:15 |  |  | 15-Dec-2017 | 22:38 | NOT WORKING | REPLACED THE 6V/4.5A RECHARGEABLE BATTERY | Closed | NO INFO | LOW | 108:23:0 | 5
942 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00015 | FETAL DOPPLER | NITIN | 9403801044 | 10337 | 11-Dec-2017 | 10:15 |  |  | 15-Dec-2017 | 22:52 | NOT WORKING | REPLACED THE 12V CORE TRANSFORMER | Closed | NO INFO | MEDIUM | 108:37:0 | 5
943 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00015 | FETAL DOPPLER | DR ANSARI | 7620409720 | 20023 | 04-Sep-2018 | 09:25 |  |  | 10-Sep-2018 | 14:00 | NOT WORKING | CHECK AND FOUND TRANSFORMER IS NOT WORKING.ALSO DOPPLER PROBE IS DEFECTIVE.CHANGED THE NEW TRANSFORMER(12V OUTPUT) AND NEW PROBE.FETAL DOPPLER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 148:35:0 | 7
944 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00017 | RADIANT WARMER | NITIN VAIRAGE | 9970364436 | 18630 | 30-Jul-2018 | 09:48 |  |  | 03-Sep-2018 | 11:05 | NOT WORKING | CHECK AND FOUND DISPLAY BOARD IS DEFECTIVE.SMPS IS ALSO DEFECTIVE. REPAIRED THE SMPS BOARD.CHANGED THE NEW DISPLAY BOARD. CUT-OFF OF HEATER AT LOW AND HIGH BODY TEMPERATURE IS CHECKED AND FOUND TO BE OK WARMER IS WORKING | Closed | NO INFO | HIGH | 841:17:0 | 36
945 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00017 | RADIANT WARMER | NITIN | 9970361437 | 36595 | 16-Dec-2019 | 09:17 |  |  | 19-Dec-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND MAKE PROPER CONNECTION OF SKIN TEMPERATURE SENSOR AND AIR TEMPERATURE SENSOR. NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | HIGH | 74:43:0 | 4
946 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00018 | BP APPARATUS ANEROID | NITIN WARALEKAR | 9403801044 | 16728 | 14-Jun-2018 | 09:06 |  |  | 19-Jun-2018 | 11:27 | NOT WORKING | REPLACED THE BLADDER | Closed | NO INFO | LOW | 122:21:0 | 6
947 | BHANDARA | PHC DEVHADI | MH/R1/BHA/PHC/DVD/00025 | BP APPARATUS ANEROID | NITIN VAIRAGADKAR | 9403801044 | 13880 | 16-Mar-2018 | 09:20 |  |  | 19-Mar-2018 | 15:47 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB WITH CONNECTOR NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 78:27:0 | 4
948 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00001 | BP APPARATUS ANEROID | S G NEWARE | 9923592482 | 5984 | 09-Aug-2017 | 10:00 |  |  | 10-Aug-2017 | 15:19 | NOT WORKING | REPLACE THE RUBBER BLADDER | Closed | NO INFO | LOW | 29:19:0 | 2
949 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00007 | CENTRIFUGE (BENCH TOP) | JAYASHREE NURSE | 9130628265 | 40618 | 07-Mar-2020 | 09:13 |  |  | 09-Mar-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN MOTOR CONNECTION IS LOOSED THEN MAKE PROPER CONNECTION NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | TC48132 | MEDIUM | 50:47:0 | 3
950 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00010 | HB METER | ASHWINI | 9960039468 | 35713 | 25-Nov-2019 | 08:09 |  |  | 25-Nov-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS ERROR 02 IS SHOWING.CLEAN THE SENSOR NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 8:51:0 | 1
951 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00014 | SUCTION PUMP PORTABLE, ELECTRONIC | ASWIN | 9960059468 | 36619 | 17-Dec-2019 | 08:10 |  |  | 17-Dec-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS MACHINE IS NOT GIVING A PRESSURE THEN FILLING OIL IN MOTOR NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 8:50:0 | 1
952 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00014 | SUCTION PUMP PORTABLE, ELECTRONIC | DR BHOIR | 9158570963 | 20397 | 15-Sep-2018 | 08:30 |  |  | 20-Sep-2018 | 12:11 | NOT WORKING | CHECK AND FOUND TUBING ARE NOT PROPERLY ATTACHED.TUBINGS ARE ALSO JAMMED WITH DUST.CLEANED THE TUBING AND FIXED IT PROPERLY.DESIRED SUCTION PRESSURE IS GENERATED NOW.SUCTION PUMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 123:41:0 | 6
953 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00020 | NEBULIZER ELECTRIC | NAGARIKA NURSE | 7788183706 | 39158 | 07-Feb-2020 | 09:06 |  |  | 08-Feb-2020 | 17:00 | NOT WORKING | CHECKED AND FOUND PROBLEM IN MOTOR IS NOT WORKING THEN MAKE POWER CONNECTION AND GREESING NOW MACHINE IS IN WORKING CONDITION. | Closed | NUM253227 | MEDIUM | 31:54:0 | 2
954 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00020 | NEBULIZER ELECTRIC | BHOIR | 9158570963 | 22318 | 22-Nov-2018 | 09:01 |  |  | 23-Nov-2018 | 11:05 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING ATTACHED TUBING AND MASK IS FAULTY AND IT IS CONSUMABLE ITEM SO IT WON'T BE PROVIDED BY THE COMPANY HOSPITAL AUTHORITY HAVE TO PURCHASE THE TUBING AND MASK | Closed | NUM253227 | MEDIUM | 26:4:0 | 2
955 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00020 | NEBULIZER ELECTRIC | DR ASHWIN | 9960059468 | 38337 | 23-Jan-2020 | 09:04 |  |  | 23-Jan-2020 | 17:00 | NOT WORKING | CHECKED AND FOUND POWER CABLE BURNED. THEN CONNECT THE NEW POWER CABLE NOW MACHINE IN WORKING CONDITION. | Closed | NUM253227 | MEDIUM | 7:56:0 | 1
956 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00020 | NEBULIZER ELECTRIC | DR ASHWIN | 9158570963 | 21926 | 08-Nov-2018 | 08:06 |  |  | 20-Nov-2018 | 11:55 | STEAM ISSUE | CHECK AND FOUND MACHINE IS WORKING BUT ATTACHED TUBING AND MASK HAS PROBLEM. PROPER FUMES CANNOT BE GENERATED DUE TO DEFECT IN MASK AND TUBING. AS THE TUBING AND MASK IS CONSUMABLE ITEM, IT IS NOT PROVIDED BY THE COMPANY. HOSPITAL AUTHORITY USE TO BUY THE NEW MASK AND TUBING. | Closed | NUM253227 | MEDIUM | 291:49:0 | 13
957 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00020 | NEBULIZER ELECTRIC | DR ASWIN | 9960039468 | 29899 | 13-Jul-2019 | 09:08 |  |  | 15-Jul-2019 | 04:05 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN NEBULIZER MASK IS DAMAGE NEED TO CHANGE MASK. MACHINE IS IN WORKING CONDITION HAND OVER TO THE USER. | Closed | NUM253227 | MEDIUM | 42:57:0 | 3
958 | BHANDARA | PHC EKODI | MH/R1/BHA/PHC/EKD/00037 | HB METER | NAGARIKA NURSE | 7788183706 | 39157 | 07-Feb-2020 | 09:04 |  |  | 08-Feb-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS ERROR 02 WAS SHOWN ON DISPLAY THEN CLEAN THE SENSOR NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | 1710013398 |  | 31:56:0 | 2
959 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00001 | RADIANT WARMER | N R GAJBHAI | 9765770084 | 3411 | 27-Apr-2017 | 09:25 |  |  | 29-Apr-2017 | 09:15 | NOT WORKING | REPLACED THE BLADDER,NOW MACHINE IS WORKING CONDITION | Closed | 17926 | LOW | 47:50:0 | 3
960 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00002 | BP APPARATUS ANEROID | SUREKHA NURSE | 8421312498 | 30055 | 17-Jul-2019 | 08:12 |  |  | 19-Jul-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT FOUND PROBLEM IN BLADDER , CULF AND BULB . CHANGE ALL PART NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | LOW | 51:48:0 | 3
961 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00002 | BP APPARATUS ANEROID | DR ASHISH NAIK | 9921804068 | 3803 | 16-May-2017 | 10:20 |  |  | 18-May-2017 | 02:30 | NOT WORKING | REPLACED NEW 3AMPS FUSE,NOW MACHINE IS WORKING FINE. | Closed | NO INFO | LOW | 40:10:0 | 3
962 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00015 | ADULT WEIGHING MACHINE | DR NIKITA | 9404456609 | 38338 | 23-Jan-2020 | 09:05 |  |  | 23-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN WEIGHT SPRING IS DAMAGE THEN CHANGE THE NEW SPRING NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | HIGH | 7:55:0 | 1
963 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00016 | RADIANT WARMER | BHAGYASHREE | 9399473760 | 37744 | 10-Jan-2020 | 08:10 |  |  | 11-Jan-2020 | 16:42 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE ISSUE. MAIN POWER SUPPLY BOARD IS NOT WORKING. INFORMED THE INCHARGE TO MAKE IT REPAIR BY ELECTRICIAN. CHECKED THE MACHINE THROUGH DIFFERENT BOARD CONNECTION. SKIN TEMPERATURE IS BEING SET AND CUT-OFF OF HEATER AT LOW AND HIGH BODY TEMPERATURE IS CHECKED AND ITS OK. EXAMINATION LIGHT IS BEING FIXED. THOROUGH TRAINING IS GIVEN TO INCHARGE SISTER. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 32:32:0 | 2
964 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00017 | SUCTION PUMP PORTABLE, ELECTRONIC | PALLAVI SISTER | 9168240558 | 17926 | 14-Jul-2018 | 08:15 |  |  | 03-Nov-2018 | 12:00 | NOT WORKING | CHECK AND FOUND MOTOR WINDING IS BARN SO REPLACE WINDING AND NOW MACHINE IS WORKING | Closed | NO INFO | MEDIUM | 2691:45:0 | 113
965 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00019 | SHADOWLESS LAMP STAND TYPE | DR NIKITA | 9404456609 | 38340 | 23-Jan-2020 | 09:08 |  |  | 23-Jan-2020 | 13:57 | NOT WORKING | CHECK AND FOUND BULB IS NOT WORKING. BULB HOLDER IS DEFECTIVE AND REPLACED WITH THE NEW ONE. FIXED THE BULB FRAME PROPERLY TO HOLDER AND IS FIXED. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 4:49:0 | 1
966 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00019 | SHADOWLESS LAMP STAND TYPE | DR PALLAVI PATIL | 9730221969 | 34921 | 09-Nov-2019 | 09:15 |  |  | 11-Nov-2019 | 12:00 | NOT WORKING | CHECK THE EQUIPMENT AND FOUND PROBLEM IN HOLDER. CHANGE THE NEW HOLDER NOW MACHINE IS IN GOOD WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 50:45:0 | 3
967 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00020 | FETAL DOPPLER | N R GAJBHAI | 9765770084 | 3412 | 27-Apr-2017 | 09:25 |  |  | 29-Apr-2017 | 10:34 | NOT WORKING | REPLACED 3 PIN PROBE,NOW HEART BEAT SOUND IS NORMAL HANDOVER MACHINE TO USER GOOD CONDITION. | Closed | NO INFO | MEDIUM | 49:9:0 | 3
968 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00020 | FETAL DOPPLER | DR PALLAVI PATIL | 9730221969 | 34922 | 09-Nov-2019 | 09:15 |  |  | 11-Nov-2019 | 15:23 | NOT WORKING | CHECK AND FOUND POWER CABLE IS DEFECTIVE. FIXED THE NEW POWER CABLE. FETAL DOPPLER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 54:8:0 | 3
969 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00023 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | DR AMIT BHURE | 7588895039 | 35660 | 23-Nov-2019 | 09:12 |  |  | 24-Nov-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS HEATING COIL IS BURNED.THEN CHANGE NEW HEATING COIL NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 26:48:0 | 2
970 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00024 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR AMIT BHURE | 7588895039 | 35302 | 18-Nov-2019 | 08:46 |  |  | 18-Nov-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN HEATING COIL AND POWER CABLE. CHANGE THE NEW POWER CABLE AND HEATING COIL NOW MACHINE IS IN GOOD WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 8:14:0 | 1
971 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00024 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | N R GAJBHAI | 9765770084 | 3413 | 27-Apr-2017 | 09:25 |  |  | 29-Apr-2017 | 12:23 | NOT WORKING | REPLACED 2KW HEATING COIL AND 16A PLUG TOP,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 50:58:0 | 3
972 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00025 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR MEGHNA MANDHARE | 9423485432 | 13607 | 09-Mar-2018 | 08:34 |  |  | 12-Mar-2018 | 14:47 | NOT WORKING | REPLACED NEW 2KW HEATING COIL, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 78:13:0 | 4
973 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00025 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR PALLAVI PATIL | 9730221969 | 33735 | 12-Oct-2019 | 09:00 |  |  | 14-Oct-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN HEATING COIL AND POWER CABLE. CHANGE THE NEW COIL AND POWER CABLE NOW MACHINE IS IN WORKING CONDITION HAND OVER TO THE USER. | Closed | NO INFO | MEDIUM | 51:0:0 | 3
974 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00025 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | M G JAYWAR | 9764207805 | 16385 | 03-Jun-2018 | 09:00 |  |  | 08-Jun-2018 | 11:01 | NOT WORKING | A.C WIRE REINSERTED PROPERLY AND RECTIFIED LEAKAGE ON VACCUM RELEASE SECTION | Closed | NO INFO | MEDIUM | 122:1:0 | 6
975 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00025 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR MEGHA | 9657630156 | 6726 | 06-Sep-2017 | 11:55 |  |  | 08-Sep-2017 | 14:53 | NOT WORKING | RECONNECT THE POWER CABLE PROPERLY | Closed | NO INFO | MEDIUM | 50:58:0 | 3
976 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00025 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | PALLAVI BHAGYASHALI | 9168240558 | 17348 | 30-Jun-2018 | 08:41 |  |  | 03-Jul-2018 | 14:00 | NOT WORKING | CHECK AND FOUND LEAKAGE PROBLEM, SEAL THE LEAKAGE WITH M SEAL HANDOVER TO USER IN WORKING CONDITION | Closed | NO INFO | MEDIUM | 77:19:0 | 4
977 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00026 | BP APPARATUS ANEROID | BAJIPALE NURSE | 9168240558 | 20193 | 10-Sep-2018 | 08:04 |  |  | 17-Sep-2018 | 11:02 | NOT WORKING | CHECK AND FOUND B.P BULB,BLADDER VALVE AND BP CUFF WAS TOTALLY DEFECTIVE. CHANGE ABOVE ALL SPARE WITH NEW ONE . NOW BP APPRATUS ANEROID IS WORKING GOOD. | Closed | NO INFO |  | 170:58:0 | 8
978 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00026 | BP APPARATUS ANEROID | DR MEGHNA MANDHARE | 9423485432 | 13608 | 09-Mar-2018 | 08:34 |  |  | 12-Mar-2018 | 16:54 | NOT WORKING | FILLED 3ML MERCURY AND ALSO REPLACED NEW GLASS TUBE, NOW MACHINE IS WORKING CONDITION | Closed | NO INFO |  | 80:20:0 | 4
979 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00054 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | BHANDARKAR NURSE | 9545488436 | 17760 | 11-Jul-2018 | 10:10 |  |  | 16-Jul-2018 | 17:30 | NOT WORKING | CHECK THE MACHINE, THEIR IS LEAKAGE PROBLEM AND COIL WAS DAMAGE. REPLACED THE COIL AND PACK THE LEAKAGE WITH TEFLON TAPE. NOW MACHINE IS IN WORKING CONDITION, HANDOVER TO USER. | Closed | NO INFO |  | 127:20:0 | 6
980 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00054 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | MISS PALLAVI | 9168240558 | 16439 | 05-Jun-2018 | 09:00 |  |  | 11-Jun-2018 | 01:11 | NOT WORKING | RECTIFIED THE LEAKAGE IN THE VACCUM OUTLET | Closed | NO INFO |  | 136:11:0 | 7
981 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00055 | SUCTION PUMP PORTABLE, ELECTRONIC | BHAJIPALE NURSE | 9168240558 | 18070 | 18-Jul-2018 | 08:45 |  |  | 21-Jul-2018 | 17:30 | NOT WORKING | CHECK THE MACHINE PROBLEM REPORTED IN WIRING. CONNECTED WIRING PROPERLY. NOW MACHINE WORKING SATISFACTORY. HANDOVER TO USER. | Closed | NO INFO |  | 80:45:0 | 4
982 | BHANDARA | PHC GOBARWAHI | MH/R1/BHA/PHC/GBW/00055 | SUCTION PUMP PORTABLE, ELECTRONIC | DR PALLAVI PATIL | 9730221969 | 34923 | 09-Nov-2019 | 09:15 |  |  | 11-Nov-2019 | 12:00 | NOT WORKING | CHECK AND FOUND SUCTION MOTOR IS FAULTY AS IT'S WIRING IS SHORT. MOTOR IS TO BE GIVEN FOR REPAIR TO LOCAL VENDOR. AT PRESENT THIS BREAKDOWN COMPLAINT IS CLOSED.AS SOON AS THE MOTOR WILL BE REPAIRED, AGAIN THE COMPLAINT WILL BE RAISED AND MACHINE WILL BE MADE FUNCTIONAL. | Closed | NO INFO |  | 50:45:0 | 3
983 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00003 | FETAL DOPPLER | DR Y S KAPGATE | 9403073011 | 3061 | 13-Apr-2017 | 09:37 |  |  | 17-Apr-2017 | 10:30 | NOT WORKING | REPLACED 3 PIN FETAL DOPPLER PROBE,NOW HEART BEAT SOUND HEARING CLEARLY. | Closed | NO INFO | MEDIUM | 96:53:0 | 5
984 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00009 | BP APPARATUS ANEROID | SAWRUPA | 96376770415 | 33218 | 30-Sep-2019 | 09:05 |  |  | 01-Oct-2019 | 12:00 | NOT WORKING | CHECK THE EQUIPMENT AND FOUND PROBLEM IN BLADDER. CHANGE THE NEW BLADDER NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | LOW | 26:55:0 | 2
985 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00011 | SUCTION PUMP PORTABLE, ELECTRONIC | DR WAGHMARE | 9075112102 | 16274 | 30-May-2018 | 09:05 |  |  | 30-May-2018 | 15:09 | NOT WORKING | CHECK AND FOUND THAT SUCTION TUBE IS NOT WORKING PROPERLY. REPLACED WITH NEW TUBING TO THE GLASS JAR. NOW THE MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 6:4:0 | 1
986 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00016 | CENTRIFUGE (BENCH TOP) | Dr Hina Raut | 7875911692 | 42643 | 26-May-2020 | 09:12 |  |  | 27-May-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS CARBON BRUSH IS DAMAGE THEN CONNECT THE NEW CARBON BRUSH NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | NO INFO | MEDIUM | 26:48:0 | 2
987 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00018 | NEBULIZER ELECTRIC | DR WAGHMARE | 9075112102 | 18132 | 19-Jul-2018 | 09:02 |  |  | 23-Jul-2018 | 17:30 | NOT WORKING | CHECK AND FOUND MOTOR IS JAMMED. FUMES CAN'T BE MADE. DID OILING NOW MOTOR IS RUNNING NORMALLY , FUMES ARE NOW GENERATED. | Closed | NGM253011 | MEDIUM | 104:28:0 | 5
988 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00018 | NEBULIZER ELECTRIC | N S DEHADE | 9923929207 | 38485 | 25-Jan-2020 | 08:13 |  |  | 26-Jan-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN NEBULIZER MASK IS DAMAGE . NEED TO CHANGE NEW NEBULIZER MASK BY PHC. MACHINE IS WORKING OK. | Closed | NGM253011 | MEDIUM | 27:47:0 | 2
989 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00018 | NEBULIZER ELECTRIC | DR SUJIT MISHRAM | 8668362933 | 24751 | 07-Feb-2019 | 08:10 |  |  | 12-Feb-2019 | 17:01 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING BUT TUBING AND MASK IS DEFECTIVE. SINCE TUBING ALONG WITH MASK IS CONSUMABLE AND IT IS NOT PROVIDED BY COMPANY. HOSPITAL ADMINISTRATION HAVE TO ARRANGE THE CONSUMABLE PART. REST NEBULIZER MACHINE IS WORKING. | Closed | NGM253011 | MEDIUM | 128:51:0 | 6
990 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00021 | HB METER | DR WAGHMARE | 9075112102 | 29570 | 06-Jul-2019 | 08:58 |  |  | 07-Jul-2019 | 08:52 | NOT WORKING | CHECKED THE EQUIPMENT FAULT IS OCCURED IN CHARGING ADAPTER REPLACE THE NEW ADAPTER NOW MACHINE IN WORKING CONDITION HAND OVER TO THE USER. | Closed | 147013295 | MEDIUM | 23:54:0 | 2
991 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00030 | BP APPARATUS ANEROID | DR. WAGHMARE | 9075112102 | 15853 | 17-May-2018 | 11:35 |  |  | 19-May-2018 | 01:49 | NOT WORKING | BLADDER AND EXTENSION TUBE IS TORN . REPLACED WITH NEW BLADDER AND EXTENSION TUBE. NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO | LOW | 38:14:0 | 3
992 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00030 | BP APPARATUS ANEROID | SAWRUPA | 96376770415 | 33219 | 30-Sep-2019 | 09:05 |  |  | 01-Oct-2019 | 12:00 | NOT WORKING | CHECK THE EQUIPMENT AND FOUND PROBLEM IN VALVE , EXTENSION TUBE AND BULB. NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | LOW | 26:55:0 | 2
993 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00030 | BP APPARATUS ANEROID | WAGHMARE | 9075112102 | 29544 | 06-Jul-2019 | 08:35 |  |  | 07-Jul-2019 | 08:40 | NOT WORKING | CHECKED THE EQUIPMENT AND CHANGE BLADDER AND BULP. NOW MACHINE IS IN WORKING CONDITION HANDOVER TO THE USER. | Closed | NO INFO | LOW | 24:5:0 | 2
994 | BHANDARA | PHC GONDUMARI | MH/R1/BHA/PHC/GNM/00033 | BP APPARATUS ANEROID | DR WAGHMARE | 9075112102 | 29571 | 06-Jul-2019 | 08:58 |  |  | 07-Jul-2019 | 08:45 | NOT WORKING | CHECKED THE EQUIPMENT AND CHANGE CULF OF BP APPRATUS NOW MACHINE IS IN WORKING CONDITION HAND OVER TO THE USER. | Closed | NO INFO |  | 23:47:0 | 2
995 | BHANDARA | PHC JAMB | MH/R1/BHA/PHC/JMB/00015 | BP APPARATUS ANEROID | POONAM CHAVAN | 9689850600 | 26035 | 19-Mar-2019 | 08:50 |  |  | 20-Mar-2019 | 14:27 | NOT WORKING | CHECK AND FOUND CUFF, BLADDER, VALVE AND BULB IS MISSING. FIXED THE NEW CUFF, BLADDER, VALVE AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 29:37:0 | 2
996 | BHANDARA | PHC JAMB | MH/R1/BHA/PHC/JMB/00018 | PHOTOTHERAPHY UNIT | POONAM CHAVAN | 9689850600 | 26036 | 19-Mar-2019 | 08:50 |  |  | 20-Mar-2019 | 13:31 | NOT WORKING | CHECK AND FOUND TUBE LIGHT(2) IS NOT GLOWING. TUBE LIGHT IS LOOSELY FIXED. FIXED THE TUBE LIGHT IN THE HOLDER PROPERLY AND STARTER IS ALSO CHANGED. NOW, ALL THE TUBE LIGHT IS GLOWING. PHOTOTHERAPHY UNIT IS NOW WORKING. | Closed | NO INFO |  | 28:41:0 | 2
997 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00001 | BP APPARATUS ANEROID | GIRISH | 9422770964 | 29217 | 28-Jun-2019 | 08:59 |  |  | 29-Jun-2019 | 06:00 | NOT WORKING | CHECK AND FOUND BLADER, BULB AND VALVE IS LEAK CHANGED THE OLD DEFECTIVE PARTS WITH NEW ONE. BP APPRAUTUS IS NOW WORKING. | Closed | NO INFO | LOW | 21:1:0 | 2
998 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00001 | BP APPARATUS ANEROID | GHANSHYAM | 9767637498 | 1758 | 20-Feb-2017 | 09:12 |  |  | 21-Feb-2017 | 12:56 | NOT WORKING | REPLACED BLADDER AND RUBBER CONNECTOR, NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | LOW | 27:44:0 | 2
999 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00001 | BP APPARATUS ANEROID | DR WANKADE | 9422770964 | 27008 | 23-Apr-2019 | 08:20 |  |  | 29-Apr-2019 | 08:10 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS LEAK CHANGE THE OLD NIBP CUFF WITH NEW ONE . NOW MACHINE IS WORKING GOOD | Closed | NO INFO | LOW | 143:50:0 | 7
1000 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00002 | BP APPARATUS ANEROID | GIRISH WANKHEDE | 9422770964 | 28824 | 19-Jun-2019 | 09:36 |  |  | 22-Jun-2019 | 11:08 | NOT WORKING | CHECK AND FOUND BLADDER AND VALVE IS LEAK. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 73:32:0 | 4
1001 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00005 | SUCTION PUMP PORTABLE, ELECTRONIC | DR WANKHEDE | 9422770964 | 2415 | 11-Mar-2017 | 09:15 |  |  | 11-Mar-2017 | 11:30 | NOT WORKING | NEWLY REPLACED THE SUCTION TUBE CONNECTOR SET,NOW MACHINE IS WORKING CONDITION | Closed | ODM0445 | MEDIUM | 2:15:0 | 1
1002 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00007 | SUCTION PUMP PORTABLE, ELECTRONIC | DR GIRISH PANKHADE | 9422770964 | 4110 | 30-May-2017 | 10:55 |  |  | 02-Jun-2017 | 14:31 | NOT WORKING | REPLACED TUBING AND CONNECTOR SET | Closed | NO INFO | MEDIUM | 75:36:0 | 4
1003 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00011 | ADULT WEIGHING MACHINE | DR GIRISH WANKHEDE | 9422770964 | 19046 | 11-Aug-2018 | 09:07 |  |  | 16-Aug-2018 | 12:00 | NOT WORKING | CHECK AND FOUND IN THE ADULT WEIGHING MACHINE IS POWER SUPPLY CONNECTION WAS BROK .IT IS REPAIR AND NOW MACHINE IS WORKING NOW | Closed | NO INFO | LOW | 122:53:0 | 6
1004 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00012 | HB METER | Dr Girish Wankhede | 9422770964 | 41910 | 25-Apr-2020 | 08:50 |  |  | 25-Apr-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS ERROR 02 IS SHOWING ON DISPLAY. THEN PROPERLY CLEAN THE SENSOR NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | 1407013 291 | MEDIUM | 8:10:0 | 1
1005 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00018 | BP APPARATUS ANEROID | Dr Girish Wankhede | 56913 9422770964 | 41909 | 25-Apr-2020 | 08:50 |  |  | 25-Apr-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS BLADDER AND BULB IS LEAK. THEN CHANGE THE NEW BLADDER AND BULB NOW MACHINE IS GOOD WORKING CONDITION. | Closed | NO INFO | LOW | 8:10:0 | 1
1006 | BHANDARA | PHC KUDEGAON | MH/R1/BHA/PHC/KDG/00032 | Fumigation Machine | Dr Girish Wankhede | 9422770964 | 41908 | 25-Apr-2020 | 08:50 |  |  | 25-Apr-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS POWER CABLE AND INTERNAL MOTOR CONNECTION IS DAMAGE THEN MAKE INTERNAL CONNECTION PROPERLY AND CONNECT NEW POWER CABLE. NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | NO INFO |  | 8:10:0 | 1
1007 | BHANDARA | PHC KHAMBA | MH/R1/BHA/PHC/KMB/00001 | BP APPARATUS ANEROID | SACHIN | 8999287489 | 16264 | 30-May-2018 | 08:05 |  |  | 30-May-2018 | 11:31 | NOT WORKING | BLADDER IS TORN AND METAL CONNECTOR IS NOT WORKING PROPERLY. REPLACED WITH NEW BLADDER AND METAL CONNECTOR. NOW THE MACHINE IS IN WORKING CONDITION | Closed | RM1010518 | LOW | 3:26:0 | 1
1008 | BHANDARA | PHC KHAMBA | MH/R1/BHA/PHC/KMB/00018 | LOW COST GLUCOMETER | SACHIN NAIK | 9604790833 | 2190 | 06-Mar-2017 | 12:20 |  |  | 07-Mar-2017 | 15:25 | NOT WORKING | REPLACED THE CELL | Closed | NO INFO | MEDIUM | 27:5:0 | 2
1009 | BHANDARA | PHC KHAMBA | MH/R1/BHA/PHC/KMB/00019 | HB METER | SACHIN NAIK | 9604790833 | 2189 | 06-Mar-2017 | 12:15 |  |  | 07-Mar-2017 | 15:19 | NOT WORKING | REMOVE THE ERROR | Closed | NO INFO | MEDIUM | 27:4:0 | 2
1010 | BHANDARA | PHC KHAMBA | MH/R1/BHA/PHC/KMB/00021 | BP APPARATUS ANEROID | SACHIN RINAIT | 8999287489 | 15810 | 16-May-2018 | 09:20 |  |  | 17-May-2018 | 01:26 | NOT WORKING | BLADDER IS TORN. REPLACED WITH NEW BLADDER. NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO |  | 16:6:0 | 2
1011 | BHANDARA | PHC KHAMBA | MH/R1/BHA/PHC/KMB/00022 | BP APPARATUS ANEROID | SACHIN | 8999287489 | 16263 | 30-May-2018 | 08:05 |  |  | 30-May-2018 | 10:45 | NOT WORKING | BULB AND METAL CONNECTOR IS NOT WORKING PROPERLY. REPLACED WITH NEW BULB AND METAL CONNECTOR. NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO |  | 2:40:0 | 1
1012 | BHANDARA | PHC KHAMBA | MH/R1/BHA/PHC/KMB/00025 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | SACHIN | 8999287489 | 29453 | 04-Jul-2019 | 09:15 |  |  | 08-Jul-2019 | 18:34 | COIL ISSUE | CHECKED THE EQUIPMENT AND FOUND COIL HAS DAMAGE THEN CHANGE THE NEW ELECTRIC COIL NOW MACHINE IS IN WORKING CONDITION HAND OVER TO THE USER. | Closed | NO INFO |  | 105:19:0 | 5
1013 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00001 | BP APPARATUS ANEROID | PADWAL NURSE | 9970907290 | 6111 | 14-Aug-2017 | 09:15 |  |  | 18-Aug-2017 | 13:01 | NOT WORKING | REPLACED THE BLADDER | Closed | 8601 | LOW | 99:46:0 | 5
1014 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00001 | BP APPARATUS ANEROID | PADVALE SISTER | 9970907290 | 20741 | 26-Sep-2018 | 08:10 |  |  | 05-Oct-2018 | 11:25 | NOT WORKING | CHECK AND FOUND BP APPPRATUS BP CUFF,BLADDER IS DEFECTIVE.CHANGE ABOVE ALL THE SPARE NOW BP APPRATUS IS WPRKING | Closed | 8601 | LOW | 219:15:0 | 10
1015 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00002 | ADULT WEIGHING MACHINE | PADAVALE SISTER | 9970907290 | 26295 | 27-Mar-2019 | 08:40 |  |  | 30-Mar-2019 | 02:10 | NOT WORKING | CHEAK AND FOUND THE MACHINE. POWER CABLE IS SHORT, SO ADULT WEIGHT MACHINE IS NOT PROPERLY WORK. THE POWER CABLE IS REPAIR AND THEN MACHINE IS WELL AND PROPERLY WORK AT PRESENT. | Closed | NO INFO | LOW | 65:30:0 | 4
1016 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00003 | ADULT WEIGHING MACHINE | PADAVALE SISTER | 9970907290 | 26296 | 27-Mar-2019 | 08:40 |  |  | 01-Apr-2019 | 15:22 | NOT WORKING | CHECK AND FOUND READING OBTAINED IS NOT CORRECT AND INVERTED. CABLE FROM PANEL BOARD TO DISPLAY IS INVERSELY FIXED AND LOAD CELL IS WIRE IS OPEN. FIXED THE LOAD CELL WIRE AND DISPLAY CABLE IS CORRECTLY FIXED THROUGH PANEL BOARD. NOW, WEIGHING MACHINE IS WORKING. | Closed | NO INFO | LOW | 126:42:0 | 6
1017 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00003 | ADULT WEIGHING MACHINE | PADWAL NURSE | 9970907290 | 18525 | 27-Jul-2018 | 09:15 |  |  | 02-Aug-2018 | 01:05 | NOT WORKING | CHECK AND FOUND TRANSFORMER IS BURNT OUT CHANGE THE NEW STEP DOWN TRANSFORMER OF RATING 12V POWER SUPPLY PROBLEM RESOLVED.WEIGHING MACHINE IS NOW WORKING | Closed | NO INFO | LOW | 135:50:0 | 7
1018 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00007 | BP APPARATUS ANEROID | PADOLE NURSE | 9970907290 | 1219 | 27-Jan-2017 | 10:35 |  |  | 03-Mar-2017 | 11:30 | NOT WORKING | BLADDER REPLACED | Closed | NO INFO | LOW | 840:55:0 | 36
1019 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00010 | BP APPARATUS ANEROID | PADWAL NURSE | 9970907290 | 6112 | 14-Aug-2017 | 09:15 |  |  | 19-Aug-2017 | 11:51 | NOT WOKRING | REPLACED THE BULB, CUFF AND BLADDER | Closed | NO INFO | LOW | 122:36:0 | 6
1020 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00010 | BP APPARATUS ANEROID | PADWAL NURSE | 9970907290 | 18526 | 27-Jul-2018 | 09:15 |  |  | 01-Aug-2018 | 16:30 | NOT WORKING | CHECK AND FOUND DISTANCE PIPE IS LEAKAGE . REPLACED THE PIPE SAME. NOW B.P. APRATUS WORKING SATISFACTORY. | Closed | NO INFO | LOW | 127:15:0 | 6
1021 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00010 | BP APPARATUS ANEROID | PADOLE NURSE | 9970907290 | 12958 | 16-Feb-2018 | 09:40 |  |  | 19-Feb-2018 | 13:11 | NOT WORKING | REPLACED THE CUFF | Closed | NO INFO | LOW | 75:31:0 | 4
1022 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00010 | BP APPARATUS ANEROID | PADOLE NURSE | 9970907290 | 10149 | 07-Dec-2017 | 08:10 |  |  | 11-Dec-2017 | 11:50 | NOT WORKING | REPLACED THE BLADDER AND CUFF | Closed | NO INFO | LOW | 99:40:0 | 5
1023 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00010 | BP APPARATUS ANEROID | SONI PADOLE | 9970907290 | 32970 | 24-Sep-2019 | 09:07 |  |  | 24-Sep-2019 | 15:27 | NOT WORKING | CHECK AND FOUND PROBLEM IN EXTENSION TUBE AND VALVE. CHANGED THE VALVE AND EXTENSION TUBE. NOW, MACHINE IS IN WORKING CONDITION AND HANDED OVER TO USER. | Closed | NO INFO | LOW | 6:20:0 | 1
1024 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00011 | FETAL DOPPLER | PADWAL NURSE | 9970907990 | 18527 | 27-Jul-2018 | 09:15 |  |  | 01-Aug-2018 | 16:30 | NOT WORKING | CHECK AND FOUND FETAL DOPPLER MACHINE PROBE WIRE WAS LOOSE. REDUX THE CONNECTION. NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 127:15:0 | 6
1025 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00011 | FETAL DOPPLER | PADWALE NURSE | 9970907290 | 1699 | 16-Feb-2017 | 09:13 |  |  | 01-Jun-2017 | 14:30 | NOT WORKING | REPLACED NEW 3 PIN PROBE ,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 2525:17:0 | 106
1026 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00012 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | PADOLE NURSE | 9970907290 | 1220 | 27-Jan-2017 | 10:40 | 01-Feb-2017 | 10:30 | 02-Feb-2017 | 12:04 | NOT WORKING | REPLACED 500 WATTS COIL,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 145:24:0 | 7
1027 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00014 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | PADOLE NURSE | 9970907290 | 10152 | 07-Dec-2017 | 08:10 |  |  | 11-Dec-2017 | 11:52 | NOT WORKING | REPLACED THE HEATER COIL | Closed | NO INFO | MEDIUM | 99:42:0 | 5
1028 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00014 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | PADOLE NURSE | 9970907290 | 1221 | 27-Jan-2017 | 10:40 |  |  | 01-Feb-2017 | 09:50 | NOT WORKING | RECHECKED THE POWER CORD | Closed | NO INFO | MEDIUM | 119:10:0 | 6
1029 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00017 | LOW COST GLUCOMETER | PADWAL NURSE | 9970907290 | 6110 | 14-Aug-2017 | 09:15 |  |  | 18-Aug-2017 | 13:06 | NOT WORKING | REPLACED THE LITHIUM BATTERY | Closed | NO INFO | MEDIUM | 99:51:0 | 5
1030 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00017 | LOW COST GLUCOMETER | PADOLE NURSE | 9970907290 | 10150 | 07-Dec-2017 | 08:10 |  |  | 11-Dec-2017 | 11:45 | NOT WORKING | REPLACED THE BATTERY | Closed | NO INFO | MEDIUM | 99:35:0 | 5
1031 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00018 | HB METER | PADOLE NURSE | 9970907290 | 1218 | 27-Jan-2017 | 10:35 |  |  | 02-Feb-2017 | 12:05 | NOT WORKING | BLADDER REPLACED | Closed | NO INFO | MEDIUM | 145:30:0 | 7
1032 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00018 | HB METER | HUMANE NURSE | 9657301191 | 4800 | 24-Jun-2017 | 09:15 |  |  | 25-Jun-2017 | 11:57 | NOT WORKING | REPLACED THE BATTERY | Closed | NO INFO | MEDIUM | 26:42:0 | 2
1033 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00018 | HB METER | PAROLE NURSE | 9970907290 | 27113 | 26-Apr-2019 | 08:40 |  |  | 01-May-2019 | 15:01 | NOT WORKING | CHECK AND FOUND ERROR 02 IS DISPLAYED ON THE SCREEN. CLEANED THE SENSOR WITH COTTON AND WATER. ERROR-02 IS ELEMINATED. HB METER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 126:21:0 | 6
1034 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00018 | HB METER | PADOLE | 9970907290 | 27786 | 21-May-2019 | 08:32 |  |  | 23-May-2019 | 14:35 | NOT WORKING | CHECK AND FOUND OPTRONIC SENSOR IS FAULTY. NEED NEW ONE. MR IS RAISED AND WHEN SENSOR WOULD BE AVAILABLE THEN CALL WILL BE RAISED AGAIN AND MACHINE WILL BE MADE FUNCTIONAL. | Closed | NO INFO | MEDIUM | 54:3:0 | 3
1035 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00019 | LOW COST GLUCOMETER | PADWAL NURSE | 9970907290 | 6108 | 14-Aug-2017 | 09:15 |  |  | 19-Aug-2017 | 11:55 | NOT WORKING | REPLACED THE LITHIUM BATTERY | Closed | NO INFO | MEDIUM | 122:40:0 | 6
1036 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00019 | LOW COST GLUCOMETER | PADOLE NURSE | 9970907290 | 10151 | 07-Dec-2017 | 08:10 |  |  | 11-Dec-2017 | 11:47 | NOT WORKING | REPLACED THE BATTERY | Closed | NO INFO | MEDIUM | 99:37:0 | 5
1037 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00020 | LOW COST GLUCOMETER | PADWAL NURSE | 9970907290 | 6109 | 14-Aug-2017 | 09:15 |  |  | 18-Aug-2017 | 13:13 | NOT WORKING | REPLACED THE LITHIUM BATTERY | Closed | NO INFO | MEDIUM | 99:58:0 | 5
1038 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00023 | DEEP FREEZER | PADWALE NURSE | 9970907290 | 2760 | 24-Mar-2017 | 07:00 |  |  | 29-Mar-2017 | 12:22 | NOT WORKING | REPLACED THE MOTOR & COMPRESSOR | Closed | NO INFO | MEDIUM | 125:22:0 | 6
1039 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00023 | DEEP FREEZER | PADWALE NURSE | 9970907290 | 1697 | 16-Feb-2017 | 09:07 |  |  | 18-Feb-2017 | 01:30 | NOT WORKING | MACHINE WAITING FOR SPARE,MOTOR WANT TO CHANGE | Closed | NO INFO | MEDIUM | 40:23:0 | 3
1040 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00024 | DEEP FREEZER | PADWALE NURSE | 9970907290 | 1698 | 16-Feb-2017 | 09:10 |  |  | 23-Feb-2017 | 01:49 | NOT WORKING | EQUIPMENT FOR CONDEMNATION | Closed | NO INFO | MEDIUM | 160:39:0 | 8
1041 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00026 | BP APPARATUS ANEROID | PADAVALE SISTER | 9970907290 | 15840 | 17-May-2018 | 08:26 |  |  | 22-May-2018 | 11:43 | NOT WORKING | BLADDER WAS TORNED. REPLACED WITH NEW BLADDER. NOW THE MACHINE IS IN WORKING | Closed | NO INFO |  | 123:17:0 | 6
1042 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00026 | BP APPARATUS ANEROID | PAROLE NURSE | 9970907290 | 27114 | 26-Apr-2019 | 08:40 |  |  | 30-Apr-2019 | 15:39 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING. | Closed | NO INFO |  | 102:59:0 | 5
1043 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00026 | BP APPARATUS ANEROID | PADVALE SISTER | 9970907290 | 20742 | 26-Sep-2018 | 08:11 |  |  | 05-Oct-2018 | 11:28 | NOT WORKING | CHECK AND FOUND BP CUFF ,BP BULB ,BP VALVE WAS DEFECTIVE CHANGE ABOVE ALL SPARE NOW BP APPRATUS IS WORKING | Closed | NO INFO |  | 219:17:0 | 10
1044 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00033 | BP APPARATUS ANEROID | PADAVALE SISTER | 9970907290 | 26297 | 27-Mar-2019 | 08:41 |  |  | 01-Apr-2019 | 14:17 | NOT WORKING | CHECK AND FOUND VALVE AND BULB IS LEAK. CHANGED THE SAME WITH THE NEW. B. P APPARATUS IS NOW WORKING. | Closed | 40170936 |  | 125:36:0 | 6
1045 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00034 | BP APPARATUS ANEROID | PADAVALE SISTER | 9970907290 | 26298 | 27-Mar-2019 | 08:40 |  |  | 01-Apr-2019 | 13:03 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN. FIXED THE NEW GLASS TUBE AND MERCURY IS REFILLED. B. P APPARATUS IS NOW WORKING. | Closed | W161207000882 |  | 124:23:0 | 6
1046 | BHANDARA | PHC KHAMARI | MH/R1/BHA/PHC/KMR/00034 | BP APPARATUS ANEROID | PADVALE SISTER | 9970907290 | 20743 | 26-Sep-2018 | 08:11 |  |  | 01-Oct-2018 | 11:48 | NOT WORKING | CHECK AND FOUND CUFF IS LEAK. CHANGED THE NEW NIBP CUFF. B. P APPARATUS IS NOW WORKING. | Closed | W161207000882 |  | 123:37:0 | 6
1047 | BHANDARA | PHC KONDHA | MH/R1/BHA/PHC/KND/00001 | BP APPARATUS ANEROID | ROHIT | 9970724284 | 42791 | 02-Jun-2020 | 08:00 |  |  | 02-Jun-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS BLADDER AND VALVE IS LEAK . THEN CHANGE THE NEW BLADDER AND VALVE NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | LOW | 9:0:0 | 1
1048 | BHANDARA | PHC KONDHA | MH/R1/BHA/PHC/KND/00001 | BP APPARATUS ANEROID | SHAILENDRA | 9689266089 | 21989 | 13-Nov-2018 | 08:02 |  |  | 21-Nov-2018 | 12:23 | LEAKAGE ISSUE | CHECK AND FOUND CUFF, BLADDER, BULB AND VALVE IS MISSING. NEW CUFF, BULB, VALVE AND BLADDER IS FIXED. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 196:21:0 | 9
1049 | BHANDARA | PHC KONDHA | MH/R1/BHA/PHC/KND/00001 | BP APPARATUS ANEROID | DR P V LIPSE | 9423114175 | 4177 | 31-May-2017 | 11:00 |  |  | 02-Jun-2017 | 12:11 | NOT WORKING | REPLACED NEW BLADDER,CUFF AND BULB,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | LOW | 49:11:0 | 3
1050 | BHANDARA | PHC KONDHA | MH/R1/BHA/PHC/KND/00006 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | SHAILENDRA YELAME | 9689266089 | 4221 | 02-Jun-2017 | 09:40 |  |  | 03-Jun-2017 | 03:30 | NOT WORKING | REPLACED NEW 500 WATTS HEATING COIL, NOW MACHINE HAND OVER TO USER AS GOOD WORKING CONDITION. | Closed | NO INFO | MEDIUM | 17:50:0 | 2
1051 | BHANDARA | PHC KONDHA | MH/R1/BHA/PHC/KND/00012 | SUCTION PUMP PORTABLE, ELECTRONIC | Dr borkar: | 9049976792 | 41567 | 09-Apr-2020 | 08:55 |  |  | 10-Apr-2020 | 10:00 | NOT WORKING | CHECK AND FOUND PRESSURE GENERATED IS NILL.PROBLEM FOUND IS OF INTERNAL TUBING. FIXED THE INTERNAL TUBINGS AND NOW DESIRED SUCTION PRESSURE IS GENERATED. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 25:5:0 | 2
1052 | BHANDARA | PHC KONDHA | MH/R1/BHA/PHC/KND/00025 | PHOTOTHERAPHY UNIT | Dr borkar: | 9049976792 | 41568 | 09-Apr-2020 | 09:00 |  |  | 10-Apr-2020 | 10:00 | NOT WORKING | CHECK THE EQUIPMENT AND FOUND PROBLEM IS TUBE LIGHT IS NOT WORKING. MR IS TO BE SEND WHEN CHOKE SPARE IS COMMING THEN WE CAN REPAIR THIS MACHINE MACHINE IS IN NOT WORKING CONDITION. | Closed | PM2K304021/04 | HIGH | 25:0:0 | 2
1053 | BHANDARA | PHC KONDHA | MH/R1/BHA/PHC/KND/00025 | PHOTOTHERAPHY UNIT | DR LEPSE | 9423114175 | 1875 | 23-Feb-2017 | 10:47 |  |  | 27-Feb-2017 | 10:30 | NOT WORKING | REFIXED THE TUBE HOLDER | Closed | PM2K304021/04 | HIGH | 95:43:0 | 5
1054 | BHANDARA | PHC KONDHA | MH/R1/BHA/PHC/KND/00036 | BP APPARATUS ANEROID | ROHIT | 9970724284 | 42792 | 02-Jun-2020 | 08:00 |  |  | 02-Jun-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS BLADDER , CUFF AND VALVE IS DAMAGED. THEN CHANGE THE NEW BLADDER, VALVE AND CUFF NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 9:0:0 | 1
1055 | BHANDARA | PHC KONDHA | MH/R1/BHA/PHC/KND/00039 | BP APPARATUS ANEROID | ROHIT | 9970724284 | 42793 | 02-Jun-2020 | 08:00 |  |  | 02-Jun-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS CUFF AND BLADDER IS DAMAGED THEN CHANGE THE NEW BLADDER AND CUFF NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 9:0:0 | 1
1056 | BHANDARA | PHC KARDI | MH/R1/BHA/PHC/KRD/00001 | BP APPARATUS ANEROID | GANESH BHUYAR | 8275553773 | 5985 | 09-Aug-2017 | 10:05 |  |  | 10-Aug-2017 | 11:30 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB WITH CONNECTOR. NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 25:25:0 | 2
1057 | BHANDARA | PHC KARDI | MH/R1/BHA/PHC/KRD/00001 | BP APPARATUS ANEROID | DR R N MANDHARE | 7197242147 | 4109 | 30-May-2017 | 10:50 |  |  | 02-Jun-2017 | 03:20 | NOT WORKING | REPLACED NEW BLADDER AND BULB WITH CONNECTOR, NOW MACHINE IS WORKING FINE | Closed | NO INFO | LOW | 64:30:0 | 4
1058 | BHANDARA | PHC KARDI | MH/R1/BHA/PHC/KRD/00006 | ADULT WEIGHING MACHINE | GANESH BHUYAR | 7972495201 | 24500 | 31-Jan-2019 | 08:54 |  |  | 05-Feb-2019 | 13:31 | NOT WORKING | CHECK AND FOUND LOAD CELL WIRE IS BROKEN. BODY IS DAMAGED AND BROKEN. FIXED THE LOAD CELL WIRE BACK TO CONTROLLER BOARD. FIXED THE BODY. ADULT WEIGHING MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 124:37:0 | 6
1059 | BHANDARA | PHC KARDI | MH/R1/BHA/PHC/KRD/00010 | RADIANT WARMER | PRASHANT WHIKHE | 7588491464 | 2103 | 02-Mar-2017 | 11:35 |  |  | 03-Mar-2017 | 01:30 | NOT WORKING | REFIXED THE SKIN PROBE CONNECTION,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | HIGH | 13:55:0 | 2
1060 | BHANDARA | PHC KARDI | MH/R1/BHA/PHC/KRD/00012 | BP APPARATUS ANEROID | DR.BANSOD | 9665433239 | 28928 | 22-Jun-2019 | 09:10 |  |  | 22-Jun-2019 | 16:18 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS LEAK. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 7:8:0 | 1
1061 | BHANDARA | PHC KARDI | MH/R1/BHA/PHC/KRD/00018 | NEBULIZER ELECTRIC | GANESH BHUYAR | 7972495201 | 24499 | 31-Jan-2019 | 08:54 |  |  | 05-Feb-2019 | 11:29 | NOT WORKING | CHECK AND FOUND POWER ON/OFF SWITCH IS NOT WORKING. REPLACED THE OLD WITH THE NEW POWER ON/OFF SWITCH. NEBULIZER MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 122:35:0 | 6
1062 | BHANDARA | PHC KARDI | MH/R1/BHA/PHC/KRD/00020 | BP APPARATUS ANEROID | DR.BANSOD | 9665433239 | 28929 | 22-Jun-2019 | 09:10 |  |  | 22-Jun-2019 | 17:21 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT AND VALVE IS LEAK. CHANGED THE OLD CUFF AND VALVE WITH THE NEW ONE. B.P APPARATUS IS NOW WORKING. | Closed | RM010568S | LOW | 8:11:0 | 1
1063 | BHANDARA | PHC KARDI | MH/R1/BHA/PHC/KRD/00020 | BP APPARATUS ANEROID | GANESH BHUYAR | 8275553773 | 5986 | 09-Aug-2017 | 10:15 |  |  | 10-Aug-2017 | 11:25 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB WITH CONNECTOR. NOW MACHINE IS WORKING CONDITION. | Closed | RM010568S | LOW | 25:10:0 | 2
1064 | BHANDARA | PHC KARDI | MH/R1/BHA/PHC/KRD/00023 | ADULT WEIGHING MACHINE | GANESH BHUYAR | 7972495201 | 24501 | 31-Jan-2019 | 08:54 |  |  | 05-Feb-2019 | 16:34 | NOT WORKING | CHECK AND FOUND SMPS OF THE MACHINE IS NOT WORKING. CHANGED THE OLD SMPS OF THE MACHINE WITH THE NEW ONE. ADULT WEIGHING MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 127:40:0 | 6
1065 | BHANDARA | PHC LENDEZARI | MH/R1/BHA/PHC/LDZ/00008 | BP APPARATUS ANEROID | DURWAS | 9657859890 | 31602 | 21-Aug-2019 | 10:00 |  |  | 22-Aug-2019 | 04:16 | NOT WORKING | CHEAKED THE EQUIPMENT AND FOUND PROBLEM IN CUFF, BLADER AND BULF. CHANGE THE BLADER, BULF AND CUFF NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | LOW | 18:16:0 | 2
1066 | BHANDARA | PHC LENDEZARI | MH/R1/BHA/PHC/LDZ/00010 | SHADOWLESS LAMP STAND TYPE | PRAPHUL HUMANE | 9370360901 | 24964 | 13-Feb-2019 | 09:03 |  |  | 14-Feb-2019 | 15:22 | NOT WORKING | CHECK AND FOUND POWER CABLE IS DEFECTIVE. REPAIRED THE POWER CABLE. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 30:19:0 | 2
1067 | BHANDARA | PHC LENDEZARI | MH/R1/BHA/PHC/LDZ/00011 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | DURVA | 9657859890 | 33021 | 24-Sep-2019 | 09:32 |  |  | 26-Sep-2019 | 10:04 | NOT WORKING | CHECK AND FOUND LOAD CELL WIRE IS OPEND WHICH IS TO BE CONNECTED TO CONTROLLER BORAD OF MACHINE. FIXED THE LOAD CELL WIRE TO CONTROLLER BORAD AND NOW WEIGHT IS EXACTLY DISPALAY ON SCREEN. ELECTRIC BABY WEIGHT MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 48:32:0 | 3
1068 | BHANDARA | PHC LENDEZARI | MH/R1/BHA/PHC/LDZ/00013 | FETAL DOPPLER | DURVA | 9657859890 | 33022 | 24-Sep-2019 | 09:32 |  |  | 26-Sep-2019 | 10:38 | NOT WORKING | FETAL DOPPLER MACHINE IS CHECKED AND FOUND BATTERY IS FULLY DAMAGE AND THIS ISSUE FOR MACHINE IS NOT WORKING. SO REPLACED THE NEW ONE BATTERY AND MACHINE IS WORKING AT PRESENT. | Closed | 123569893 | MEDIUM | 49:6:0 | 3
1069 | BHANDARA | PHC LENDEZARI | MH/R1/BHA/PHC/LDZ/00016 | BP APPARATUS ANEROID | DURWAS | 9657859890 | 31601 | 21-Aug-2019 | 10:00 |  |  | 22-Aug-2019 | 04:19 | NOT WORKING | CHEAKED THE EQUIPMENT AND FOUND PROBLEM IN CUFF, VALVE. CHANGE THE CUFF AND VALVE AND NOW MACHINE IS IN WORKING CONDITION AND HAND OVER TO USER. | Closed | NO INFO | LOW | 18:19:0 | 2
1070 | BHANDARA | PHC LENDEZARI | MH/R1/BHA/PHC/LDZ/00030 | BP APPARATUS ANEROID | DR PRASHANT | 9421896752 | 29888 | 13-Jul-2019 | 08:11 |  |  | 15-Jul-2019 | 12:10 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER . REPLACE THE NEW BLADDER AND EXTENSION TUBE NOW MACHINE IS IN WORKING CONDITION HAND OVER TO THE USER. | Closed | NO INFO |  | 51:59:0 | 3
1071 | BHANDARA | PHC LENDEZARI | MH/R1/BHA/PHC/LDZ/00032 | BP APPARATUS ANEROID | PRAPHUL HUMANE | 9370360901 | 24965 | 13-Feb-2019 | 09:03 |  |  | 14-Feb-2019 | 12:19 | NOT WORKING | CHECK AND FOUND BLADDER AND VALVE IS LEAK. CHANGED THE DEFECTIVE PART WITH THE NEW. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 27:16:0 | 2
1072 | BHANDARA | PHC LENDEZARI | MH/R1/BHA/PHC/LDZ/00038 | BP APPARATUS ANEROID | PRAPHUL HUMANE | 9370360901 | 24966 | 13-Feb-2019 | 09:03 |  |  | 14-Feb-2019 | 16:25 | NOT WORKING | CHECK AND FOUND BLADDER, BULB AND VALVE IS LEAK. CHANGED THE CUFF, BLADDER, BULB AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 31:22:0 | 2
1073 | BHANDARA | PHC LENDEZARI | MH/R1/BHA/PHC/LDZ/00041 | NEBULIZER ELECTRIC | DR PRASHANT | 9421896752 | 29889 | 13-Jul-2019 | 08:12 |  |  | 15-Jul-2019 | 12:07 | NOT WORKING | CHECKED THE EQUIPMENT FOUND PROBLEM IN LOOSE CONNECTION RECONNECT THE CONNECTION NOW MACHINE IS IN WORKING CONDITION HAND OVER TO THE USER. | Closed | NO INFO |  | 51:55:0 | 3
1074 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00001 | BP APPARATUS ANEROID | YOGESHWAR PADVALE | 9049472217 | 24493 | 31-Jan-2019 | 08:52 |  |  | 05-Feb-2019 | 10:26 | NOT WORKING | CHECK AND FOUND CONNECTING TUBE, VALVE AND BLADDER IS LEAK. CHANGED THE NEW VALVE, BLADDER AND CONNECTING TUBE. B.P APPARATUS IS NOW WORKING. | Closed | 1837 | LOW | 121:34:0 | 6
1075 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00001 | BP APPARATUS ANEROID | SANDEEP GAJBE | 9049472217 | 3772 | 16-May-2017 | 09:00 |  |  | 19-May-2017 | 12:35 | NOT WORKING | REPLACED THE BULB, CONNECTOR AND BLADDER | Closed | 1837 | LOW | 75:35:0 | 4
1076 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00001 | BP APPARATUS ANEROID | DR LANJEKAR | 9423112300 | 32482 | 14-Sep-2019 | 09:00 |  |  | 14-Sep-2019 | 15:20 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND BULB. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 1837 | LOW | 6:20:0 | 1
1077 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00004 | DEEP FREEZER | MS MANISHA | 9420619210 | 1321 | 31-Jan-2017 | 11:00 |  |  | 06-Feb-2017 | 10:30 | NOT WORKING | TEMPERATURE SET AS PER MACHINE | Closed | NO INFO | MEDIUM | 143:30:0 | 7
1078 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00009 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | PADULE | 9049472217 | 10716 | 19-Dec-2017 | 09:07 |  |  | 20-Dec-2017 | 16:32 | NOT WORKING | REPLACED THE HEATER COIL , POWER CABLE AND PRESSURE GAUGE | Closed | NO INFO | MEDIUM | 31:25:0 | 2
1079 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00012 | BP APPARATUS ANEROID | Y S PADOLE | 9049472217 | 13875 | 16-Mar-2018 | 09:15 |  |  | 19-Mar-2018 | 11:05 | NOT WORKING | REPLACED NEW BULB WITH CONNECTOR NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 73:50:0 | 4
1080 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00012 | BP APPARATUS ANEROID | Y S PADOLE | 9049472217 | 12066 | 22-Jan-2018 | 09:21 |  |  | 24-Jan-2018 | 14:28 | NOT WORKING | REPLACED THE BULP | Closed | NO INFO | LOW | 53:7:0 | 3
1081 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00016 | BP APPARATUS ANEROID | SANDEEP GAJBE | 9049472217 | 3773 | 16-May-2017 | 09:02 |  |  | 19-May-2017 | 12:30 | NOT WORKING | REPLACED THE BLADDER | Closed | NO INFO | LOW | 75:28:0 | 4
1082 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00016 | BP APPARATUS ANEROID | Y S PADOLE | 9049472217 | 13876 | 16-Mar-2018 | 09:15 |  |  | 19-Mar-2018 | 09:07 | NOT WORKING | REPLACED NEW CUFF AND BULB WITH CONNECTOR. NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 71:52:0 | 4
1083 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00016 | BP APPARATUS ANEROID | PADULE | 9049472217 | 10715 | 19-Dec-2017 | 09:07 |  |  | 20-Dec-2017 | 16:37 | NOT WORKING | REPLACED THE BLADDER | Closed | NO INFO | LOW | 31:30:0 | 2
1084 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00019 | NEBULIZER ELECTRIC | Y S PADOLE | 9049472217 | 22447 | 28-Nov-2018 | 09:03 |  |  | 01-Dec-2018 | 11:05 | TUBING DAMAGE | CHECK AND FOUND NEBULISER TUBING IS BREAK.NEED TO CHANGE.TUBING IS NOT PROVIDE COMPANY TO THE HOSPITAL THIS IS NOT IN CONTRACT OTHERWISE MACHINE IS WORKING GOOD | Closed | NO INFO | MEDIUM | 74:2:0 | 4
1085 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00021 | HB METER | PARODE | 9049472217 | 43571 | 27-Jun-2020 | 09:03 |  |  | 29-Jun-2020 | 09:01 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS ERROR 02 WAS SHOWN ON DISPLAY THEN CLEAN THE SENSOR NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 47:58:0 | 3
1086 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00021 | HB METER | Y S PADVALE | 9049472217 | 23023 | 15-Dec-2018 | 08:25 |  |  | 21-Dec-2018 | 11:18 | NOT WORKING | CHECK AND FOUND HB METER POWER SUPPLY ADAPTER IS TOTALLY DAMAGE. CHANGE WITH NEW POWER ADAPTER. NOW HB METER IS WORKING GOOD | Closed | NO INFO | MEDIUM | 146:53:0 | 7
1087 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00021 | HB METER | Y S PADOLE | 9049472217 | 36297 | 07-Dec-2019 | 09:02 |  |  | 09-Dec-2019 | 16:20 | NOT WORKING | CHECK AND FOUND ERROR IS DISPLAYED ON SCREEN WHEN MACHINE IS TURNED 'ON'. CLEANED THE SENSOR. ERROR IS REMOVED. READING IS TAKEN AND RESULTS OBTAINED IS SATISFACTORY. HB METER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 55:18:0 | 3
1088 | BHANDARA | PHC MOHADURA | MH/R1/BHA/PHC/MHD/00023 | NEBULIZER ELECTRIC | Y S PAROLE | 9049482217 | 40087 | 26-Feb-2020 | 09:01 |  |  | 26-Feb-2020 | 10:00 | NOT WORKING | CHECK AND FOUND NEBULIZER MACHINE IS CHECKED PROBLEM IN MASK AND TUBING. VAPOR IS NOT COMING AND MACHINE IS CHEAK PRESENT IN MO SIR AND SISTER NOW WORKING POSITION. | Closed | NO INFO |  | 0:59:0 | 1
1089 | BHANDARA | PHC MURMADI | MH/R1/BHA/PHC/MRD/00003 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | B H MADANKAR | 9579882691 | 30654 | 31-Jul-2019 | 09:29 |  |  | 01-Aug-2019 | 14:45 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT AND POWER CORD IS DEFECTIVE. CHANGED THE OLD COIL AND POWER CORD WITH THE NEW ONE. STERILIZER MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 29:16:0 | 2
1090 | BHANDARA | PHC MURMADI | MH/R1/BHA/PHC/MRD/00014 | DIGITAL BP APPARATUS | B S MADHANKAR | 9579882691 | 5637 | 25-Jul-2017 | 09:50 |  |  | 27-Jul-2017 | 13:33 | NOT WORKING | REPLACED BP PLASTIC CONNECTOR | Closed | BPD41234 | LOW | 51:43:0 | 3
1091 | BHANDARA | PHC MURMADI | MH/R1/BHA/PHC/MRD/00015 | BP APPARATUS ANEROID | B S MADHANKAR | 9579882691 | 5636 | 25-Jul-2017 | 09:50 |  |  | 26-Jul-2017 | 18:38 | NOT WORKING | REPLACE RUBBER BLADDER | Closed | NO INFO | LOW | 32:48:0 | 2
1092 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00001 | ADULT WEIGHING MACHINE | KULDIP MHASKE | 9021959441 | 6632 | 05-Sep-2017 | 09:15 |  |  | 07-Sep-2017 | 12:59 | NOT WORKING | REPLACE NEW 6A PLUG TOP,AND PROPERLY CLEAN THE LOAD CELL CABLE | Closed | NO INFO | LOW | 51:44:0 | 3
1093 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00002 | BP APPARATUS ANEROID | GIRHEPNUJE | 9272341645 | 3801 | 16-May-2017 | 10:16 |  |  | 18-May-2017 | 12:30 | NOT WORKING | REPLACED NEW BLUB AND BLADDER,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 50:14:0 | 3
1094 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00002 | BP APPARATUS ANEROID | KULDIP MHASKE | 9021959442 | 24979 | 13-Feb-2019 | 09:30 |  |  | 16-Feb-2019 | 12:15 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS MISSING. FIXED THE NEW BLADDER, VALVE AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 74:45:0 | 4
1095 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00002 | BP APPARATUS ANEROID | GM BANSODE | 8007608264 | 35393 | 19-Nov-2019 | 08:12 |  |  | 19-Nov-2019 | 17:00 | NOT WORKING | CHECK THE EQUIPMENT AND FOUND PROBLEM IN EXTENSION TUBE IS LEAK. CHANGE THE NEW EXTENSION TUBE NOW MACHINE IS WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | LOW | 8:48:0 | 1
1096 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00003 | BP APPARATUS ANEROID | DR BANGAR | 7798549718 | 1217 | 27-Jan-2017 | 10:15 |  |  | 31-Jan-2017 | 12:30 | NOT WORKING | REPLACED BLADDER AND CUFF | Closed | NO INFO | LOW | 98:15:0 | 5
1097 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00008 | FETAL DOPPLER | KULDIP MHASKE | 9021959442 | 24983 | 13-Feb-2019 | 09:30 |  |  | 16-Feb-2019 | 12:34 | NOT WORKING | CHECK AND FOUND FEMALE CONNECTOR OF POWER CABLE IS LOOSENED FORM INSIDE. FIXED THE FEMALE CONNECTOR AND CONNECTED THE POWER CABLE. NOW MACHINE IS WORKING. | Closed | DFM2K804011/19 | MEDIUM | 75:4:0 | 4
1098 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00008 | FETAL DOPPLER | MASKE | 9021959442 | 36969 | 23-Dec-2019 | 08:17 |  |  | 23-Dec-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN PCB BOARD IS SHORT CIRCUITED THEN REPAIR THIS BOARD NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | DFM2K804011/19 | MEDIUM | 8:43:0 | 1
1099 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00012 | SUCTION PUMP PORTABLE, ELECTRONIC | KULDEEP MHASKE | 9021959442 | 10618 | 16-Dec-2017 | 09:20 |  |  | 20-Dec-2017 | 09:24 | NOT WORKING | ACTUAL MACHINE IS BABY WEIGHT MACHINE SO WE REARRANGED LOAD CELL ARRANGEMENTS, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 96:4:0 | 5
1100 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00017 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | KULDEEP MASKE | 9021959442 | 17332 | 29-Jun-2018 | 09:30 |  |  | 05-Jul-2018 | 16:55 | NOT WORKING | CHECK THE MACHINE PROBLEM REPORTED IN WIRING, AND BUSH WAS ABSENT. FITTED NEW BUSH AND FITTED WIRING CORRECTLY NOW MACHINE IS IN WORKING CONDITION, GIVING WEIGHT ACCURATELY. | Closed | NO INFO | MEDIUM | 151:25:0 | 7
1101 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00018 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | GIRHEPNUJE | 9272341645 | 3802 | 16-May-2017 | 10:16 |  |  | 18-May-2017 | 12:00 | NOT WORKING | REPLACED NEW 500 WATTS HEATING COIL,NOW HEATING DONE PROPERLY MACHINE HANDOVER TO USER AS GOOD WORKING CONDITION | Closed | NO INFO | MEDIUM | 49:44:0 | 3
1102 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00018 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | KULDEEP MHASKE | 9021959442 | 10619 | 16-Dec-2017 | 09:20 |  |  | 20-Dec-2017 | 11:27 | NOT WORKING | REPLACED NEW 2KW HEATING COIL, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 98:7:0 | 5
1103 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00019 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | KULDIP MHASKE | 9021959442 | 24982 | 13-Feb-2019 | 09:30 |  |  | 16-Feb-2019 | 15:20 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT. CHANGED THE NEW ELEMENT(1.5 KW). AUTOCLAVE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 77:50:0 | 4
1104 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00019 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | KULDEEP MHASKE | 9021959442 | 10621 | 16-Dec-2017 | 09:20 |  |  | 20-Dec-2017 | 10:29 | NOT WORKING | REPLACED NEW 2KW HEATING COIL, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 97:9:0 | 5
1105 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00046 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | KULDEEP MHASKE | 9021959442 | 10620 | 16-Dec-2017 | 09:20 |  |  | 20-Dec-2017 | 11:31 | NOT WORKING | REPLACED NEW 2KW HEATING COIL, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 98:11:0 | 5
1106 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00046 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | GM BANSODE | 8007608264 | 35392 | 19-Nov-2019 | 08:12 |  |  | 19-Nov-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN HEATING COIL AND POWER CABLE.CHANGE THE NEW POWER CABLE AND HEATING COIL. NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO |  | 8:48:0 | 1
1107 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00049 | BP APPARATUS ANEROID | KULDIP MHASKE | 9021959442 | 24981 | 13-Feb-2019 | 09:30 |  |  | 16-Feb-2019 | 14:17 | NOT WORKING | CHECK AND FOUND BLADDER AND VALVE IS LEAK. CHANGED THE OLD DEFECTIVE PART WITH THE NEW. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 76:47:0 | 4
1108 | BHANDARA | PHC NAKADONGARI | MH/R1/BHA/PHC/NKD/00052 | DIGITAL BP APPARATUS | KULDIP MHASKE | 9021959442 | 24980 | 13-Feb-2019 | 09:30 |  |  | 16-Feb-2019 | 15:23 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF, BULB AND VALVE IS MISSING. FIXED THE NEW BLADDER, CUFF, BULB AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NA |  | 77:53:0 | 4
1109 | BHANDARA | PHC PAHELA | MH/R1/BHA/PHC/PHL/00007 | RADIANT WARMER | A S KULSANGE | 8412903940 | 10822 | 21-Dec-2017 | 09:18 |  |  | 21-Dec-2017 | 13:56 | NOT WORKING | INTERCHANGE THE HEATER COIL SUPPLY | Closed | NO INFO | HIGH | 4:38:0 | 1
1110 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00001 | BP APPARATUS ANEROID | DR KULDEEP | 9423383685 | 35721 | 25-Nov-2019 | 08:25 |  |  | 27-Nov-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS VALVE IS LEAK. THEN CHANGE THE VALVE NOW MACHINE IS IN WORKING CONDITION. | Closed | RM1000474 | LOW | 51:35:0 | 3
1111 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00001 | BP APPARATUS ANEROID | R. K. Dhatrak | 9049704300 | 44132 | 13-Jul-2020 | 08:37 |  |  | 13-Jul-2020 | 10:00 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF, VALVE AND BULB IS MISSING. FIXED THE NEW BLADDER, CUFF, VALVE AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | RM1000474 | LOW | 1:23:0 | 1
1112 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00006 | SYRINGE NEEDLE DESTROYER | UMESH RAUT | 9765929465 | 5609 | 24-Jul-2017 | 10:30 |  |  | 25-Jul-2017 | 18:45 | NOT WORKING | REPLACE THE POWER CARD | Closed | NO INFO | LOW | 32:15:0 | 2
1113 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00007 | RADIANT WARMER | DR KULDIP | 9423383685 | 38336 | 23-Jan-2020 | 09:01 |  |  | 23-Jan-2020 | 17:00 | NOT WORKING | CHECKED AND FOUND PROBLEM IN TEMPERATURE SENSOR IS NOT WORKING THEN MAKE PROPER CONNECTION NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | HIGH | 7:59:0 | 1
1114 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00008 | SUCTION PUMP PORTABLE, ELECTRONIC | BHUMESH RAUT | 9765929465 | 3293 | 24-Apr-2017 | 09:00 | 26-Apr-2017 | 01:48 | 27-Apr-2017 | 13:38 | NOT WORKING | REPLACE 16A SWITCH AND SUCING TUBE | Closed | NO INFO | MEDIUM | 76:38:0 | 4
1115 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00010 | BP APPARATUS ANEROID | R. K. Dhatrak | 9049704300 | 44131 | 13-Jul-2020 | 08:37 |  |  | 13-Jul-2020 | 10:00 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER, VALVE AND BULB. FIXED THE NEW BLADDER, VAVLE AND BULB. BP APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 1:23:0 | 1
1116 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00010 | BP APPARATUS ANEROID | UMESH RAUT | 9765929465 | 5610 | 24-Jul-2017 | 10:30 |  |  | 26-Jul-2017 | 18:42 | NOT WORKING | REPLACE RUPPER BLADDER AND TUBE | Closed | NO INFO | LOW | 56:12:0 | 3
1117 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00010 | BP APPARATUS ANEROID | B M RAUT | 9765929465 | 17761 | 11-Jul-2018 | 10:15 |  |  | 11-Jul-2018 | 15:45 | NOT WORKING | CHECK AND FOUND BLADDER IS NOT WORKING, REPLACED SAME. NOW B.P. APRATUS WORKING SATISFACTORY. | Closed | NO INFO | LOW | 5:30:0 | 1
1118 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00015 | BP APPARATUS ANEROID | R. K. Dhatrak | 9049704300 | 44130 | 13-Jul-2020 | 08:37 |  |  | 13-Jul-2020 | 10:00 | NOT WORKING | CHECK AND FOUND BULB AND VALVE IS LEAK. CONNECTOR IS ALSO BROKEN. FIXED THE NEW BULB, VALVE AND CONNECTOR. B.P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 1:23:0 | 1
1119 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00015 | BP APPARATUS ANEROID | DHUMESH RAUT | 9765929465 | 17303 | 29-Jun-2018 | 08:00 |  |  | 04-Jul-2018 | 02:35 | NOT WORKING | CHECK AND FOUND BLADDER, RUBBER TUBE, AND CLOTH IS FAULTY. REPLACED SAME , NOW B.P. APRATUS IS WORKING SATISFACTORY. | Closed | NO INFO |  | 114:35:0 | 6
1120 | BHANDARA | PHC PIMPALGAON | MH/R1/BHA/PHC/PLG/00031 | BP APPARATUS ANEROID | R. K. Dhatrak | 9049704300 | 44129 | 13-Jul-2020 | 08:37 |  |  | 13-Jul-2020 | 10:00 | NOT WORKING | CHECK AND FOUND BLADER, BULB AND VALVE IS LEAK. CHANGE THE BLADER, BULB AND VALVE AND NOW MACHINE IS WORKING. | Closed | 17M0244 |  | 1:23:0 | 1
1121 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00001 | BP APPARATUS ANEROID | A G SATKAR | 9420620227 | 15803 | 16-May-2018 | 08:20 |  |  | 21-May-2018 | 13:12 | NOT WORKING | FILLED MERCURY 1 ML . NOW THE MACHINE IS IN WORKING CONDITION | Closed | 54186 | LOW | 124:52:0 | 6
1122 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00001 | BP APPARATUS ANEROID | ALKA SATKAR | 9420620227 | 33294 | 01-Oct-2019 | 09:26 |  |  | 03-Oct-2019 | 10:27 | NOT WORKING | CHEAKED AND FOUND LEAKAGE IN BLADDER, BULB AND VALVE. CHANGE THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P. APPRAUTUS IS WORKING. | Closed | 54186 | LOW | 49:1:0 | 3
1123 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00001 | BP APPARATUS ANEROID | A V BUYAR | 9420620227 | 21085 | 05-Oct-2018 | 08:10 |  |  | 11-Oct-2018 | 10:37 | NOT WORKING | CHECK AND FOUND BLADDER VALVE IS LEAK.CUFF AND BULB IS WORN OUT CHANGED THE CUFF BLADDER VALVE AND BLADDER VALVE AND BULB.BP APPRATUS IS NOW WORKING | Closed | 54186 | LOW | 146:27:0 | 7
1124 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00001 | BP APPARATUS ANEROID | ALKA KATKAR | 9420620227 | 26829 | 15-Apr-2019 | 09:25 |  |  | 16-Apr-2019 | 16:07 | NOT WORKING | CHECK AND FOUND BLADDER, BULB AND VALVE IS LEAK. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 54186 | LOW | 30:42:0 | 2
1125 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00003 | SYRINGE NEEDLE DESTROYER | RAJENDRA FULZDE | 7875735841 | 10269 | 09-Dec-2017 | 08:40 |  |  | 13-Dec-2017 | 14:58 | NOT WORKING | REPLACED NEW POWER WIRE, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 102:18:0 | 5
1126 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00003 | SYRINGE NEEDLE DESTROYER | ALKA SATKAR | 9420620227 | 17330 | 29-Jun-2018 | 09:15 |  |  | 05-Jul-2018 | 14:15 | NOT WORKING | CHECK AND FOUND INTERNAL WIRING IS FAULTY. RECONNECTION DONE. NOW MACHINE IS WORKING SATISFACTORY. | Closed | NO INFO | LOW | 149:0:0 | 7
1127 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00004 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | V R JAMBULKAR | 9420359722 | 5319 | 14-Jul-2017 | 09:30 |  |  | 20-Jul-2017 | 14:30 | NOT WORKING | REPLACED NEW 500 WATTS HEATING COIL AND 16A PLUG TOP, NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 149:0:0 | 7
1128 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00004 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | VISHAKA JAMMBHULKAR | 9420359722 | 9083 | 14-Nov-2017 | 08:30 |  |  | 15-Nov-2017 | 11:04 | NOT WORKING | REPLACED NEW 2KW HEATING COIL, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 26:34:0 | 2
1129 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00004 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | ALAKA SATKAR | 9420620227 | 25613 | 06-Mar-2019 | 09:13 |  |  | 06-Mar-2019 | 16:59 | NOT WORKING | CHECK AND FOUND POWER CORD IS DEFECTIVE. REPAIRED THE POWER CORD. VERTICAL AUTOCLAVE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 7:46:0 | 1
1130 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00004 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | ALKA TATKAR | 9420620227 | 22805 | 10-Dec-2018 | 09:43 |  |  | 15-Dec-2018 | 12:04 | NOT WORKING | CHECK AND FOUND HEATING ELEMENT IS DEFECTIVE CHANGED THE NEW HEATING ELEMENT. AUTOCLAVE IS NOW WORKING | Closed | NO INFO | MEDIUM | 122:21:0 | 6
1131 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00004 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | V R JAMBULKAR | 9420359722 | 3608 | 09-May-2017 | 09:40 |  |  | 11-May-2017 | 14:54 | NOT WORKING | REPLACED THE HEATER COIL | Closed | NO INFO | MEDIUM | 53:14:0 | 3
1132 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00004 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | A V BUYAR | 9420620227 | 21088 | 05-Oct-2018 | 08:10 |  |  | 11-Oct-2018 | 11:40 | NOT WORKING | CHECK AND FOUND HEATING ELEMENT IS DEFECTIVE CHANGED THE NEW HEATING ELEMENT OF 1.5KW AUTOCLAVE IS NOW WORKING | Closed | NO INFO | MEDIUM | 147:30:0 | 7
1133 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00005 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | VISHAKHA JAMBUDKAR | 9420359722 | 12497 | 03-Feb-2018 | 08:46 |  |  | 05-Feb-2018 | 11:14 | NOT WORKING | REPLACED NEW 14 INCHES GASKIT, MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 50:28:0 | 3
1134 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00006 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | ALKA TATKAR | 9420620227 | 22806 | 10-Dec-2018 | 09:43 |  |  | 15-Dec-2018 | 11:25 | NOT WORKING | CHECK AND FOUND POWER CORD IS FAULTY CHANGED THE NEW POWER CORD. NOW AUTOCLAVE IS WORKING | Closed | NO INFO | MEDIUM | 121:42:0 | 6
1135 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00006 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | RAJENDRA FULZDE | 7875735841 | 10268 | 09-Dec-2017 | 08:40 |  |  | 13-Dec-2017 | 09:45 | NOT WORKING | REPLACED NEW GASKIT, MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 97:5:0 | 5
1136 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00006 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | VISHAKHA JAMBUDKAR | 9420359722 | 12498 | 03-Feb-2018 | 09:46 |  |  | 05-Feb-2018 | 10:17 | NOT WORKING | REPLACED NEW 14 INCH RUBBER GASKIT, MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 48:31:0 | 3
1137 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00006 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | A V BUYAR | 9420620227 | 21086 | 05-Oct-2018 | 08:10 |  |  | 11-Oct-2018 | 11:43 | NOT WORKING | CHECK AND FOUND HEATING ELEMENT IS DEFECTIVE CHANGED THE NEW HEATING OF 1.5KW CERAMIC CONNECTOR IS ALSO CHANGED .AOUTOCLAVE IS WORKING NOW | Closed | NO INFO | MEDIUM | 147:33:0 | 7
1138 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00008 | Fumigation Machine | V R JAMBULKAR | 9420359722 | 5320 | 14-Jul-2017 | 09:30 |  |  | 20-Jul-2017 | 11:29 | NOT WORKING | REARRANGED THE ROTATOR IN THAT MOTOR, AND REFIXED THE CAPACITOR CONNECTION NOW MACHINE IS WORKING. | Closed | NO INFO | LOW | 145:59:0 | 7
1139 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00008 | Fumigation Machine | ALKA SATHKAR | 9420620227 | 23703 | 08-Jan-2019 | 08:12 |  |  | 17-Jan-2019 | 10:46 | NOT WORKING | CHECK AND FOUND HOLLOW PLASTIC TUBE FIXED TO THE MOTOR THAT HELPS THE WATER TO BE SUCKED FROM THE BOTTOM IS MISSING. IF THE ABOVE PART IS ARRANGED LATER, THEN IT WILL BE FIXED AND THE MACHINE WILL BE MADE FUNCTIONAL. NOW AT PRESENT THE CALL IS BEING CLOSE TILL FURTHER NOTIFICATIONS OF ARRANGE MISSING PART | Closed | NO INFO | LOW | 218:34:0 | 10
1140 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00008 | Fumigation Machine | VISHAKA JAMMBHULKAR | 9420359722 | 9082 | 14-Nov-2017 | 08:30 |  |  | 18-Nov-2017 | 11:25 | NOT WORKING | DONE REWINDING ON MOTOR AND ALSO REPLACED NEW BEARING AND CAPACITOR INTO AN MACHINE.NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 98:55:0 | 5
1141 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00008 | Fumigation Machine | ALAKA SATKAR | 9020620227 | 25614 | 06-Mar-2019 | 09:13 |  |  | 08-Mar-2019 | 10:25 | NOT WORKING | CHECK AND FOUND MACHINE PLASTIC HOLLOW TUBE IS MISSING AND MOTOR IS BURNT OUT.FIXED THE NEW HOLLOW TUBE AND REWINDING THE MOTOR NOW MACHINE IS WORKING GOOD | Closed | NO INFO | LOW | 49:12:0 | 3
1142 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00008 | Fumigation Machine | V R JAMBULKAR | 9420359722 | 3607 | 09-May-2017 | 09:40 |  |  | 11-May-2017 | 14:35 | NOT WORKING | REPLACED THE POWER CORD WITH PLUG TOP | Closed | NO INFO | LOW | 52:55:0 | 3
1143 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00011 | SUCTION PUMP PORTABLE, ELECTRONIC | ALKA KATKAR | 9420620227 | 26832 | 15-Apr-2019 | 09:25 |  |  | 17-Apr-2019 | 14:42 | NOT WORKING | CHECK AND FOUND LOW SUCTION PRESSURE IS GENERATED. INTERNAL PIPE IS OPENED. FIXED THE INTERNAL PIPING AND OILING IS DONE. NOW, DESIRED SUCTION PRESSURE IS GENERATED.SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 53:17:0 | 3
1144 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00011 | SUCTION PUMP PORTABLE, ELECTRONIC | V R JAMBULKAR | 9420359722 | 5321 | 14-Jul-2017 | 09:30 |  |  | 20-Jul-2017 | 13:29 | NOT WORKING | AFTER CHECKED A MACHINE A MOTOR WINDING GET BURNT COMPLETELY AND IT NEED EXTENSIVE REPAIR AND ALSO BEYOND ECONOMICAL REPAIR.SO WE RECOMMEND THAT MACHINE TO CONDEMNATION. | Closed | NO INFO | MEDIUM | 147:59:0 | 7
1145 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00014 | FETAL DOPPLER | ALKA SATKAR | 9420620227 | 17331 | 29-Jun-2018 | 09:15 |  |  | 05-Jul-2018 | 13:15 | NOT WORKING | CHECK AND FOUND ON /OFF SWITCH IS NOT WORKING, FIX THE PROBLEM, NOW MACHINE IS WORKING PROPERLY. | Closed | NO INFO | MEDIUM | 148:0:0 | 7
1146 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00014 | FETAL DOPPLER | RAJENDRA FULZDE | 7875735841 | 10270 | 09-Dec-2017 | 08:40 |  |  | 13-Dec-2017 | 15:01 | NOT WORKING | SOLDERED THE 12V ADAPTER, NOW MACHINE IS WORKING CONDITION. . | Closed | NO INFO | MEDIUM | 102:21:0 | 5
1147 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00016 | RADIANT WARMER | ALAKA SATKAR | 9420620227 | 25618 | 06-Mar-2019 | 09:13 |  |  | 06-Mar-2019 | 18:08 | NOT WORKING | CHECK AND FOUND DISPLAY IS NOT COMING. FIXED THE CONNECTION THROUGH BOARD PROPERLY.IC'S IS INSERTED AGAIN AND FIXED. NOW, DISPLAY IS PROPERLY WORKING. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 8:55:0 | 1
1148 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00022 | HB METER | ALKA SATKAR | 9420620227 | 27888 | 24-May-2019 | 08:40 |  |  | 26-May-2019 | 16:35 | NOT WORKING | CHECK AND FOUND EO2 ERROR IS DISPLAYED ON SCREEN. CLEANED THE OPTRONIC SENSOR AND ERROR IS GONE. HB METER IS NOW WORKING. | Closed | 1407013289 | MEDIUM | 55:55:0 | 3
1149 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00022 | HB METER | ALKA TATKAR | 9420620227 | 22807 | 10-Dec-2018 | 09:43 |  |  | 15-Dec-2018 | 11:00 | NOT WORKING | CHECK AND FOUND SENSOR IS COVERD WITH DUST AND DUE TO WHICH READING MEASURED IS NOT ACCURATE. CLEANED THE SENSOR. NOW READING MEASURED IS ACCURATE. HA METER IS WORKING | Closed | 1407013289 | MEDIUM | 121:17:0 | 6
1150 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00026 | AUTOCLAVE HP VERTICAL (2 BIN) DIGITAL | ALAKA SATKAR | 9420620227 | 25615 | 06-Mar-2019 | 09:13 |  |  | 06-Mar-2019 | 17:02 | NOT WORKING | CHECK AND FOUND POWER CORD IS DEFECTIVE. REPAIRED THE POWER CORD. VERTICAL AUTOCLAVE IS NOW WORKING. | Closed | NO INFO |  | 7:49:0 | 1
1151 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00026 | AUTOCLAVE HP VERTICAL (2 BIN) DIGITAL | A V BUYAR | 9420620227 | 21087 | 05-Oct-2018 | 08:10 |  |  | 11-Oct-2018 | 11:45 | NOT WORKING | CHECK AND FOUND HEATING ELEMNET IS DEFECTIVE CHANGED THE NEW HEATING OF 2KW. AUTOCLAVE IS NOW WORKING . | Closed | NO INFO |  | 147:35:0 | 7
1152 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00027 | BP APPARATUS ANEROID | A G SATKAR | 9420620227 | 15802 | 16-May-2018 | 08:17 |  |  | 21-May-2018 | 10:09 | NOT WORKING | REPLACED WITH NEW METAL CONNECTOR. NOW THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO |  | 121:52:0 | 6
1153 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00027 | BP APPARATUS ANEROID | ALKA SATKAR | 9420620227 | 27889 | 24-May-2019 | 08:45 |  |  | 28-May-2019 | 15:26 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER, BULB AND VALVE. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 102:41:0 | 5
1154 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00027 | BP APPARATUS ANEROID | ALKA KATKAR | 9420620227 | 26830 | 15-Apr-2019 | 09:25 |  |  | 16-Apr-2019 | 17:11 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF, BULB AND VALVE IS MISSING. FIXED THE NEW BLADDER, CUFF, BULB AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 31:46:0 | 2
1155 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00027 | BP APPARATUS ANEROID | A V BUYAR | 9420620227 | 21084 | 05-Oct-2018 | 08:10 |  |  | 11-Oct-2018 | 11:49 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT.LEAKAGE IN BLADDER AND VALVE BEING DECTECTED .CHANGED THE NEW CUFF VALVED AND BLADDER B.P APPRATUS IS NOW WORKING | Closed | NO INFO |  | 147:39:0 | 7
1156 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00029 | BP APPARATUS ANEROID | ALKA KATKAR | 9420620227 | 26831 | 15-Apr-2019 | 09:25 |  |  | 16-Apr-2019 | 11:09 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK AND CUFF IS WORN OUT. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 25:44:0 | 2
1157 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00029 | BP APPARATUS ANEROID | ALKA SATKAR | 9420620227 | 27890 | 24-May-2019 | 08:45 |  |  | 28-May-2019 | 16:28 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND BULB IS FOUND. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 103:43:0 | 5
1158 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00030 | BP APPARATUS ANEROID | ALKA SATKAR | 9420620227 | 34848 | 08-Nov-2019 | 08:35 |  |  | 09-Nov-2019 | 16:46 | NOT WORKING | CHECK AND FOUND LEAKAGE IN VALVE AND CUFF IS WORN OUT. CHANGE THE DEFECTIVE CUFF AND VALVE WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 32:11:0 | 2
1159 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00032 | DIGITAL BP APPARATUS | ALKA SATKAR | 7588564008 | 33859 | 15-Oct-2019 | 08:15 |  |  | 15-Oct-2019 | 10:09 | NOT WORKING | CHECK AND FOUND DIGITAL B. P. APPRATUS POWER SUPPLY ADAPTER IS TOTALLY DAMAGE. CHANGE WITH NEW POWER ADAPTER. NOW DIGITAL B. P. APPRATUS IS WORKING GOOD | Closed | NO INFO |  | 1:54:0 | 1
1160 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00042 | BP APPARATUS ANEROID | ALKA SATKAR | 9420620227 | 34847 | 08-Nov-2019 | 08:35 |  |  | 09-Nov-2019 | 14:44 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE IN VALVE. CHANGED THE DEFECTIVE VALVE WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 30:9:0 | 2
1161 | BHANDARA | PHC POHARA | MH/R1/BHA/PHC/POR/00047 | DIGITAL BP APPARATUS | ALKA SATKAR | 9420620227 | 35868 | 28-Nov-2019 | 08:28 |  |  | 28-Nov-2019 | 11:40 | NOT WORKING | CHECK ANS FOUND NIBP CUFF IS FAULTY AND HAVING LEAKAGE. POWER ADAPTER IS ALSO MISSING. FIXED THE NEW NIBP CUFF AND NEW POWER ADAPTER IS HANDED OVER. DIGITAL BP APPARATUS IS NOW WORKING. | Closed | 11453 |  | 3:12:0 | 1
1162 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00001 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DAHIWALE | 8275117813 | 35068 | 13-Nov-2019 | 08:34 |  |  | 16-Nov-2019 | 15:34 | NOT WORKING | CHECK THE EQUIPMENT AND FOUND PROBLEM IN HEATING COIL IS BURNED. THEN CHANGED THE NEW HEATING COIL. NOW MACHINE IS WORKING IN GOOD CONDITION. | Closed | NO INFO | MEDIUM | 79:0:0 | 4
1163 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00010 | BP APPARATUS ANEROID | DAHIWALE | 9421947401 | 19342 | 20-Aug-2018 | 08:40 |  |  | 21-Aug-2018 | 09:30 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK.CHANGED THE NEW BLADDER.B.P APPARATUS IS WORKING NOW. | Closed | 58956 | LOW | 24:50:0 | 2
1164 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00010 | BP APPARATUS ANEROID | DAVALE | 8275117813 | 29794 | 11-Jul-2019 | 08:20 |  |  | 14-Jul-2019 | 09:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND BLADER AND BULB IS LEACKED. REPLACE THE NEW BLADDER AND BULB NOW MACHINE IS IN WORKING CONDITION HAND OVER TO THE USES. | Closed | 58956 | LOW | 72:40:0 | 4
1165 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00010 | BP APPARATUS ANEROID | DR KAMKAR | 8378034134 | 2836 | 30-Mar-2017 | 09:00 |  |  | 04-Apr-2017 | 13:45 | NOT WORKING | REPLACED THE BULB CONNECTOR AND BULB | Closed | 58956 | LOW | 124:45:0 | 6
1166 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00011 | ADULT WEIGHING MACHINE | DAHIWALE | 8275117813 | 27130 | 26-Apr-2019 | 08:58 |  |  | 01-May-2019 | 15:26 | NOT WORKING | CHECK AND FOUND POWER ON/OFF SWITCH IS FAULTY. REPAIRED THE POWER ON/OFF SWITCH. ADULT WEIGHING MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 126:28:0 | 6
1167 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00011 | ADULT WEIGHING MACHINE | Dr Mamta Lanjewar | 7262831510 | 42551 | 23-May-2020 | 08:40 |  |  | 23-May-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS MACHINE WAS NOT START . THEN SHORT OUT THE PROBLEM AND RECONNECT THE DISPLAY CABLE NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | LOW | 8:20:0 | 1
1168 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00011 | ADULT WEIGHING MACHINE | DAHIVALE | 9421947401 | 21867 | 03-Nov-2018 | 09:12 |  |  | 20-Nov-2018 | 12:51 | NOT WORKING | CHECK AND FOUND POWER CORD IS DEFECTIVE AND 12V TRANSFORMER IS NOT WORKING. CHANGED THE NEW POWER CORD AND ALSO THE TRANSFORMER. ADULT WEIGHING MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 411:39:0 | 18
1169 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00012 | NEBULIZER ELECTRIC | DAHIVALE | 9421941401 | 19014 | 10-Aug-2018 | 09:20 |  |  | 17-Aug-2018 | 10:05 | NOT WRKING | CHECK AND FOUND ELECTRIC SHOCK BEING EXPERIENCE WHILE OPERATING THE MACHINE CHECK THE PHC PROMISES AND FOUND NO ERTHING IS AVAILABLE .MACHINE IS WORKING WITH NO ERROR | Closed | NGM253222 | MEDIUM | 168:45:0 | 8
1170 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00012 | NEBULIZER ELECTRIC | DAHIWALE | 8275117813 | 21469 | 19-Oct-2018 | 08:17 |  |  | 30-Oct-2018 | 12:49 | NOT WORKING | CHECK AND FOUND MOTOR IS JAMMED. OPENED THE MOTOR, OILING IS DONE FOR FREE ROTATION. ENTIRE CONNECTION IS CHECKED. NOW, MOTOR IS RUNNING NORMALLY. NEBULIZER MACHINE IS NOW WORKING. | Closed | NGM253222 | MEDIUM | 268:32:0 | 12
1171 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00012 | NEBULIZER ELECTRIC | DAHIWALE | 8275117813 | 33546 | 09-Oct-2019 | 08:12 |  |  | 10-Oct-2019 | 14:28 | NOT WORKING | CHECK AND FOUND MOTOR IS NOT WORKING AS ITS INTERNAL WINDING IS SHORT. GAVE THE MOTOR TO LOCAL VENDOR FOR REPAIR. AT PRESENT THIS BREAKDOWN COMPLAINT IS BEING CLOSED. AS SOON AS THE MOTOR WILL BE REPAIRED, AGAIN THE CALL WILL BE RAISED AND MACHINE WILL HE HANDED IN WORKING CONDITION. | Closed | NGM253222 | MEDIUM | 30:16:0 | 2
1172 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00012 | NEBULIZER ELECTRIC | K J DAHIWALE | 8275113813 | 35656 | 23-Nov-2019 | 09:09 |  |  | 23-Nov-2019 | 16:54 | NOT WORKING | CHECK AND FOUND MOTOR WINDING IS SHORT. REPAIRED MOTOR OF WHICH NEW WINDING IS DONE IS FIXED. MOTOR IS WORKING AND DESIRED AMOUNT OF PRESSURE AT THE NOZZLE IS EXPERIENCED. FUMES IS BEING GENERATED. NEBULIZER MACHINE IS NOW WORKING. | Closed | NGM253222 | MEDIUM | 7:45:0 | 1
1173 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00014 | SUCTION PUMP PORTABLE, ELECTRONIC | DAHIVALE | 9421941401 | 19015 | 10-Aug-2018 | 09:20 |  |  | 17-Aug-2018 | 12:10 | NOT WORKING | CHECK AND FOUND ELECTRIC SHOCK BEING EXPERIENCE WHILE OPERATING THE MACHINE CHECK THE PHC PROMISES AND FOUND NO ERTHING IS AVAILABLE .MACHINE IS WORKING WITH NO ERROR | Closed | NO INFO | MEDIUM | 170:50:0 | 8
1174 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00015 | BINOCULAR MICROSCOPE | MUKESH THULSUNGE | 8275398486 | 5124 | 07-Jul-2017 | 10:02 |  |  | 08-Jul-2017 | 11:27 | NOT WORKING | REPLACED NEW 6V 20 WATTS LAMP WITH LAMP HOLDER, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 25:25:0 | 2
1175 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00018 | HB METER | GANGURDE | 9029818904 | 5262 | 12-Jul-2017 | 11:19 |  |  | 13-Jul-2017 | 11:15 | NOT WORKING | DONE CLEANING PROCESS NOW ERROR PROBLEM GET RECTIFIED,MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 23:56:0 | 2
1176 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00019 | Fumigation Machine | DAVALE | 8275117813 | 29796 | 11-Jul-2019 | 08:23 |  |  | 14-Jul-2019 | 10:10 | NOT WORKING | CHECK AND FOUND MOTOR IS BURN I. E. IT'S REWARDING NEED TO BE DONE. REWINDING IS DONE AND. MOTOR IS REPAIRED. MOTOR IS FIND AND CHECKED FUMES ARE NOW GENERATED. FUMIGATION MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 73:47:0 | 4
1177 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00019 | Fumigation Machine | DAHIWALE | 8275117813 | 22951 | 13-Dec-2018 | 09:10 |  |  | 20-Dec-2018 | 13:13 | NOT WORKING | CHECK AND FOUND POWER PLUCK AND CORD IS FAULTY. CHANGED THE NEW POWER PLUCK AND CORD. FUMIGATION MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 172:3:0 | 8
1178 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00019 | Fumigation Machine | DR KAMKAR | 8378034134 | 3023 | 12-Apr-2017 | 09:15 |  |  | 13-Apr-2017 | 13:03 | NOT WORKING | REPLACED THE REWINDING , BEARING AND BEARING BUSH | Closed | NO INFO | LOW | 27:48:0 | 2
1179 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00020 | BP APPARATUS ANEROID | DAHIWALE | 8275117813 | 21470 | 19-Oct-2018 | 08:17 |  |  | 30-Oct-2018 | 11:39 | NOT WORKING | CHECK AND FOUND NIBP CUFF FAULTY. CHANGED THE NEW NIBP CUFF. B. P APPARATUS IS NOW WORKING. | Closed | 58952 | LOW | 267:22:0 | 12
1180 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00020 | BP APPARATUS ANEROID | DAVALE | 8275117813 | 29795 | 11-Jul-2019 | 08:21 |  |  | 14-Jul-2019 | 10:03 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND LOCK OF THE BP APPRAUTUS IS FAULTY. CHANGE THE LOCK NOW MACHINE IS IN WORKING CONDITION HAND OVER TO THE USES. | Closed | 58952 | LOW | 73:42:0 | 4
1181 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00023 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DAHIWALE | 8275117813 | 35069 | 13-Nov-2019 | 08:35 |  |  | 16-Nov-2019 | 16:36 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT AND POWER CORD IS INTERNALLY SHORT. FIXED THE NEW HEATING COIL AND POWER CORD. AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO |  | 80:1:0 | 4
1182 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00024 | BP APPARATUS ANEROID | K J DAHIWALE | 9421947401 | 14129 | 23-Mar-2018 | 09:16 |  |  | 24-Mar-2018 | 14:12 | NOT WORKING | REPLACED NEW BULB WITH CONNECTOR NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 28:56:0 | 2
1183 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00051 | DIGITAL BP APPARATUS | DAHIWALE | 8275117813 | 22950 | 13-Dec-2018 | 09:10 |  |  | 20-Dec-2018 | 12:05 | PROB CUFF REQUIRED | CHECK AND FOUND CUFF IS LEAK. CHANGED THE NEW NIBP CUFF. DIGITAL B. P APPARATUS IS NOW WORKING | Closed | W16120701182 |  | 170:55:0 | 8
1184 | BHANDARA | PHC SANGADI | MH/R1/BHA/PHC/SND/00054 | DIGITAL BP APPARATUS | DAHIWALE | 8275117813 | 27131 | 26-Apr-2019 | 08:58 |  |  | 01-May-2019 | 16:28 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS FAULTY. CHANGED THE OLD DEFECTIVE CUFF WITH THE NEW ONE. DIGITAL B. P APPARATUS IS NOW WORKING. | Closed | W16120701194 |  | 127:30:0 | 6
1185 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00011 | RADIANT WARMER | SP BANJARI | 8830745273 | 32344 | 11-Sep-2019 | 09:05 |  |  | 12-Sep-2019 | 11:12 | NOT WORKING | CHECK AND FOUND BULB IS FUSED . CHANGED THE NEW ONE BULB RADIANT WARMER. MACHINE IS NOW WORKING. | Closed | NO INFO | HIGH | 26:7:0 | 2
1186 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00011 | RADIANT WARMER | S P WANJARE | 9763632461 | 8380 | 26-Oct-2017 | 09:25 |  |  | 28-Oct-2017 | 11:49 | NOT WORKING | REPLACED THE BULB | Closed | NO INFO | HIGH | 50:24:0 | 3
1187 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00015 | BP APPARATUS ANEROID | SP BANJARI | 8830745273 | 32345 | 11-Sep-2019 | 09:08 |  |  | 11-Sep-2019 | 11:26 | NOT WORKING | CHECK AND FOUND LEACKEAGE IN BLADDER IS ALSO WORK OUT CHANGE THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B.P. APPRATUS IS WORKING. | Closed | NO INFO | LOW | 2:18:0 | 1
1188 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00016 | BP APPARATUS ANEROID | S P WANJARE | 9763632461 | 8376 | 26-Oct-2017 | 09:25 |  |  | 28-Oct-2017 | 11:12 | NOT WORKING | REPLACED THE BULB,BLADDER AND BULB CONNECTOR | Closed | NO INFO | LOW | 49:47:0 | 3
1189 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00017 | BP APPARATUS ANEROID | S P WANJARE | 9763632461 | 8377 | 26-Oct-2017 | 09:25 |  |  | 28-Oct-2017 | 11:14 | NOT WORKING | REPLACED THE METAL CONNECTOR | Closed | NO INFO | LOW | 49:49:0 | 3
1190 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00018 | BP APPARATUS ANEROID | S P WANJARE | 9763632461 | 8378 | 26-Oct-2017 | 09:25 |  |  | 28-Oct-2017 | 11:16 | NOT WORKING | REPLACED THE BULB AND BLADDER | Closed | NO INFO | LOW | 49:51:0 | 3
1191 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00018 | BP APPARATUS ANEROID | SP BANJARI | 8830745273 | 32346 | 11-Sep-2019 | 09:08 |  |  | 11-Sep-2019 | 11:48 | NOT WORKING | CHECK AND FOUND LEACKEAGE IN VALVE AND CUFF IS ALSO WORK OUT CHANGE THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P. APPRAUTUS IS NOW WORKING. | Closed | NO INFO | LOW | 2:40:0 | 1
1192 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00019 | BP APPARATUS ANEROID | S P WANJARE | 9763632461 | 8379 | 26-Oct-2017 | 09:25 |  |  | 28-Oct-2017 | 10:19 | NOT WORKING | REPLACED THE BLADDER | Closed | NO INFO | LOW | 48:54:0 | 3
1193 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00019 | BP APPARATUS ANEROID | SP BANJARI | 8830745273 | 32347 | 11-Sep-2019 | 09:08 |  |  | 11-Sep-2019 | 11:56 | NOT WORKING | CHECK AND FOUND LEACKEAGE IN BLADDER IS ALSO WORK OUT. CHANGE THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P. APPRTAUS IS NOW WORKING. | Closed | NO INFO | LOW | 2:48:0 | 1
1194 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00037 | BP APPARATUS ANEROID | WANJARE NURSE | 9763632461 | 19774 | 29-Aug-2018 | 09:00 |  |  | 03-Sep-2018 | 14:00 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK.CHANGED THE NEW BLADDER.B.P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 125:0:0 | 6
1195 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00047 | BP APPARATUS ANEROID | WANJARE NURSE | 9763632461 | 19772 | 29-Aug-2018 | 09:00 |  |  | 03-Sep-2018 | 12:02 | NOT WORKING | CHECK AND FOUND BULB BLADDER,VALVE IS LEAK CHANGED THE ABOVE MENTION PART.NOW MACHINE IS WORKING | Closed | 17M0102 |  | 123:2:0 | 6
1196 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00048 | BP APPARATUS ANEROID | WANJARE NURSE | 9763632461 | 19773 | 29-Aug-2018 | 09:00 |  |  | 03-Sep-2018 | 13:00 | NOT WORKING | CHECK AND FOUND BLADDER AND VALVE LEAKAGE.CUFF IS ALSO WORN OUT.CHANGED THE NEW CUFF,BLADDER,VALVE AND BULB. | Closed | 17M0137 |  | 124:0:0 | 6
1197 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00051 | INFRARED LAMP | S.P.VANGARI | 8830745273 | 34849 | 08-Nov-2019 | 08:40 |  |  | 11-Nov-2019 | 15:45 | NOT WORKING | CHECK AND FOUND IR BULB IS FUSED AND POWER CABLE IS DEFECTIVE. FIXED THE NEW IR BULB AND POWER CABLE. INFRARED LAMP IS NOW WORKING. | Closed | NA |  | 79:5:0 | 4
1198 | BHANDARA | PHC SHAHAPUR | MH/R1/BHA/PHC/SPR/00051 | INFRARED LAMP | SP BANJARI | 8830745273 | 32348 | 11-Sep-2019 | 09:08 |  |  | 13-Sep-2019 | 12:24 | NOT WORKING | CHECK AND FOUND IR BULB IS FUSED. MR IS RAISED AND MATERIAL IS IN TRANSIT. AT PRESENT THIS BREAKDOWN CALL IS CLOSED. AS SOON AS MATERIAL WILL BE RECEIVED, AGAIN THE CALL WILL BE RAISED AND MACHINE WILL BE MADE FUNCTIONAL. | Closed | NA |  | 51:16:0 | 3
1199 | BHANDARA | PHC SARANDI | MH/R1/BHA/PHC/SRD/00002 | BP APPARATUS ANEROID | N N WANJARI | 9539374598 | 35659 | 23-Nov-2019 | 09:12 |  |  | 24-Nov-2019 | 12:00 | NOT WORKING | CHECK THE EQUIPMENT AND FOUND PROBLEM IN BLADDER IS LEAK .CHANGE THE NEW BLADDER NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | LOW | 26:48:0 | 2
1200 | BHANDARA | PHC SARANDI | MH/R1/BHA/PHC/SRD/00014 | BP APPARATUS ANEROID | GANESH INGOTE | 9823105731 | 11655 | 13-Jan-2018 | 09:01 |  |  | 18-Jan-2018 | 12:25 | NOT WORKING | REPLACED NEW BLADDER AND BULB WITH CONNECTOR NOW MACHINE IS WORKING CONDITION. | Closed | 746 | LOW | 123:24:0 | 6
1201 | BHANDARA | PHC SARANDI | MH/R1/BHA/PHC/SRD/00021 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | HINGOLE | 9823105731 | 1670 | 14-Feb-2017 | 09:55 |  |  | 15-Feb-2017 | 15:03 | NOT WORKING | FUSE REPLACED | Closed | 68223 | MEDIUM | 29:8:0 | 2
1202 | BHANDARA | PHC SARANDI | MH/R1/BHA/PHC/SRD/00028 | SUCTION PUMP PORTABLE, ELECTRONIC | GANESH HINGOLE | 9823105731 | 10289 | 11-Dec-2017 | 08:20 |  |  | 11-Dec-2017 | 14:04 | NOT WORKING | RECONNECT THE MOTOR CONNECTION | Closed | NO INFO | MEDIUM | 5:44:0 | 1
1203 | BHANDARA | PHC SAWARLA | MH/R1/BHA/PHC/SRL/00008 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | YOGINI DEVGADE | 9673642434 | 24875 | 11-Feb-2019 | 09:05 |  |  | 14-Feb-2019 | 16:46 | NOT WORKING | CHECK AND FOUND WEIGHT IS NOT MEASURED. LOAD CELL CONNECTOR THROUGH PANEL BOARD IS LOOSENED. FIXED THE CONNECTOR OF LOAD CELL. NOW WEIGHT IS BEING MEASURED. | Closed | NO INFO | MEDIUM | 79:41:0 | 4
1204 | BHANDARA | PHC SAWARLA | MH/R1/BHA/PHC/SRL/00009 | BP APPARATUS ANEROID | YOGINI DEVGADE | 9673642434 | 24876 | 11-Feb-2019 | 09:05 |  |  | 12-Feb-2019 | 08:50 | NOT WORKING | CHECK AND FOUND BLADDER,BULB AND VALVE IS LEAK.CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW.B.P APPARATUS IS NOW WORKING. | Closed | RM0105697 | LOW | 23:45:0 | 2
1205 | BHANDARA | PHC SAWARLA | MH/R1/BHA/PHC/SRL/00010 | BP APPARATUS ANEROID | DR AMRAPALI | 7769023671 | 36019 | 02-Dec-2019 | 08:09 |  |  | 02-Dec-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS BLADDER IS LEAK. THEN CHANGE THE NEW BLADDER NOW MACHINE IS IN WORKING CONDITION. | Closed | 62138 | LOW | 8:51:0 | 1
1206 | BHANDARA | PHC SAWARLA | MH/R1/BHA/PHC/SRL/00010 | BP APPARATUS ANEROID | MINAKSHI NURSE | 8766895540 | 27686 | 17-May-2019 | 08:12 |  |  | 21-May-2019 | 12:10 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS HAVING LEAKAGE. CHANGED THE OLD DEFECTIVE CUFF WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 62138 | LOW | 99:58:0 | 5
1207 | BHANDARA | PHC SAWARLA | MH/R1/BHA/PHC/SRL/00010 | BP APPARATUS ANEROID | YOGINI | 9673642434 | 27095 | 25-Apr-2019 | 08:56 |  |  | 27-Apr-2019 | 09:17 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER, VALVE AND BULB. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 62138 | LOW | 48:21:0 | 3
1208 | BHANDARA | PHC SAWARLA | MH/R1/BHA/PHC/SRL/00011 | BP APPARATUS ANEROID | YOGINI | 9673642434 | 27096 | 25-Apr-2019 | 08:56 |  |  | 27-Apr-2019 | 09:20 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND VALVE. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 98135306 | LOW | 48:24:0 | 3
1209 | BHANDARA | PHC SAWARLA | MH/R1/BHA/PHC/SRL/00011 | BP APPARATUS ANEROID | MINAKSHI NURSE | 8766895540 | 27687 | 17-May-2019 | 08:12 |  |  | 21-May-2019 | 12:05 | NOT WOKRING | CHECK AND FOUND BLADDER, VALVE AND BULB IS LEAK. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 98135306 | LOW | 99:53:0 | 5
1210 | BHANDARA | PHC SAWARLA | MH/R1/BHA/PHC/SRL/00016 | DIGITAL BP APPARATUS | MINAKSHI NURSE | 8766895540 | 32891 | 23-Sep-2019 | 09:03 |  |  | 25-Sep-2019 | 10:36 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS HAVING LEAKAGE AND DUE TO WHICH B.P COULD NOT BE MEASURED. CHANGED THE OLD DEFECTIVE CUFF WITH THE NEW ONE. DIGITAL B. P APPARATUS IS NOW WORKING. | Closed | 133 |  | 49:33:0 | 3
1211 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00001 | BP APPARATUS ANEROID | DR NAVKHARKAR | 9890230763 | 4803 | 24-Jun-2017 | 09:25 |  |  | 28-Jun-2017 | 10:36 | NOT WORKING | REPLACE BLADDER AND BULB METAL CONNECTOR | Closed | RM0105711 | LOW | 97:11:0 | 5
1212 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00001 | BP APPARATUS ANEROID | DR SOLE | 8956750579 | 27707 | 17-May-2019 | 08:50 |  |  | 22-May-2019 | 11:54 | NOT WORKING | CHECK AND FOUND BLADDER AND EXTENSION TUBE IS HAVING LEAKAGE. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B.P APPARATUS IS NOW WORKING. | Closed | RM0105711 | LOW | 123:4:0 | 6
1213 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00002 | BP APPARATUS ANEROID | DR ANCHAL | 8956750579 | 32144 | 05-Sep-2019 | 09:03 |  |  | 05-Sep-2019 | 15:45 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS LEAK. CUFF IS ALSO WORN OUT. CHANGED THE OLD DEFECTIVE PARTS WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | 901361 | LOW | 6:42:0 | 1
1214 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00003 | NEBULIZER ELECTRIC | DHARMIK JANBANDHU | 9326979013 | 23447 | 29-Dec-2018 | 09:08 |  |  | 01-Jan-2019 | 14:12 | STEAM ISSUE | CHECK AND FOUND TUBING IS DEFECTIVE. NEW TUBING IS TO BE ARRANGED FROM HOSPITAL ADMINISTRATION. REST, MACHINE IS WORKING. | Closed | NGM253288 | MEDIUM | 77:4:0 | 4
1215 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00008 | RADIANT WARMER | DR BADOLE | 9421896752 | 39563 | 15-Feb-2020 | 08:38 |  |  | 15-Feb-2020 | 13:38 | NOT WORKING | CHECK AND FOUND SMPS BOARD IS BURNT DUE TO SHORT CIRCUIT AND DISPLAY CONTROLLER BOARD IS DEFECTIVE. FIXED THE NEW SMPS BOARD AND DISPLAY CONTROLLER BOARD.CUT-OFF OF HEATER AT LOW AND HIGH BODY TEMPERATURE IS CHECKED AND FOUND TO BE OK. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 5:0:0 | 1
1216 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00008 | RADIANT WARMER | DR CHOLE | 8956750579 | 32073 | 03-Sep-2019 | 09:20 |  |  | 07-Sep-2019 | 15:42 | NOT WORKING | CHECK AND FOUND SMPS IS BURNT AND IS FAULTY. DEFECTIVE SMPS IS BEING SENT TO MUMBAI PCB LAB FOR REPAIR. AS SOON AS REPAIRED SMPS WOULD BE RECEIVED FROM MUMBAI, IT WILL BE FIXED AND MACHINE WILL BE MADE FUNCTIONAL AGAIN BY RAISING THE COMPLAINT. AT PRESENT, THIS PRESENT BREAKDOWN IS CLOSED. | Closed | NO INFO | HIGH | 102:22:0 | 5
1217 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00009 | ADULT WEIGHING MACHINE | DR NAVKHRKAR | 9890230763 | 4347 | 07-Jun-2017 | 09:20 |  |  | 09-Jun-2017 | 15:41 | NOT WORKING | CALIBRATE THE MACHINE THE ERROR PROBLEM WAS SOLVED | Closed | NO INFO | LOW | 54:21:0 | 3
1218 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00009 | ADULT WEIGHING MACHINE | DR NAVKHARKAR | 9890230763 | 1140 | 24-Jan-2017 | 09:36 |  |  | 28-Jan-2017 | 12:29 | NOT WORKING | REFIXED THE LOAD CELL | Closed | NO INFO | LOW | 98:53:0 | 5
1219 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00013 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR ANCHAL | 8956750579 | 32185 | 06-Sep-2019 | 09:05 |  |  | 09-Sep-2019 | 15:16 | NOT WORKING | CHECK AND FOUND 1 DOOR HANDLE AND 1 DRUM HANDLE IS BROKEN. NEED NEW HANDLES. MR TO BE RAISED.WHEN NEW HANDLES WOULD BE RECEIVED, AGAIN COMPLIANT WILL BE RAISED AND PART WOULD BE FIXED. AT PRESENT, PRESENT BREAKDOWN COMPLAINT IS CLOSED. | Closed | NO INFO | MEDIUM | 78:11:0 | 4
1220 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00013 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR NAVKHAKAR | 9890230763 | 10841 | 22-Dec-2017 | 09:40 |  |  | 22-Dec-2017 | 19:02 | NOT WORKING | REPLACED THE 16A POWER CARD | Closed | NO INFO | MEDIUM | 9:22:0 | 1
1221 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00013 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | HARSHALI SELOKAR | 9404308287 | 33858 | 15-Oct-2019 | 08:12 |  |  | 15-Oct-2019 | 10:15 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT AND ALSO POWER CABLE IS SHORT FIXED THE NEW HEATING COIL OF 2KW AND ALSO THE POWER CABLE. VERTICAL AUTOCLAVE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 2:3:0 | 1
1222 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00014 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR ANCHAL | 8956750579 | 32186 | 06-Sep-2019 | 09:05 |  |  | 09-Sep-2019 | 14:10 | NOT WORKING | CHECK AND FOUND POWER CORD IS MISSING. FIXED THE NEW POWER CORD THROUGH THE COIL. AUTOCLAVE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 77:5:0 | 4
1223 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00014 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR NAVKHAKAR | 9890230763 | 10842 | 22-Dec-2017 | 09:40 |  |  | 22-Dec-2017 | 19:09 | NOT WORKING | REPLACED THE POWER CARD | Closed | NO INFO | MEDIUM | 9:29:0 | 1
1224 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00014 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | ROHINI NURSE | 8485078189 | 26652 | 09-Apr-2019 | 08:15 |  |  | 12-Apr-2019 | 17:04 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT AND ALSO POWER CABLE IS MISSING. FIXED THE NEW HEATING COIL OF 2KW AND ALSO THE POWER CABLE. VERTICAL AUTOCLAVE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 80:49:0 | 4
1225 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00020 | HB METER | DHARMIK JANBANDHU | 9326979013 | 40483 | 05-Mar-2020 | 08:08 |  |  | 06-Mar-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS ERROR 02 IS SHOWING ON SCREEN. THEN CLEAN THE SENSOR NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | 1407013288 | MEDIUM | 27:52:0 | 2
1226 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00020 | HB METER | DR ANCHAL | 8956750579 | 32184 | 06-Sep-2019 | 09:08 |  |  | 09-Sep-2019 | 15:06 | NOT WORKING | CHECK AND FOUND POWER ADAPTER IS NOT WORKING. NEW POWER ADAPTER OF RATING 12V,700MA IS REQUIRED. MR TO BE RAISED. WHEN NEW POWER ADAPTER WOULD BE RECEIVED, AGAIN COMPLIANT WILL BE RAISED AND HB METER WOULD BE MADE FUNCTIONAL. AT, PRESENT, PRESENT BREAKDOWN COMPLIANT IS CLOSED. | Closed | 1407013288 | MEDIUM | 77:58:0 | 4
1227 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00020 | HB METER | DILIP LANDGEWAR | 8087244696 | 5398 | 17-Jul-2017 | 10:45 | 18-Jul-2017 | 13:18 | 19-Jul-2017 | 13:18 | NOT WORKING | CLEAN THE LENSE AND REMOVED BLOOD LOCKAGE | Closed | 1407013288 | MEDIUM | 50:33:0 | 3
1228 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00022 | BP APPARATUS ANEROID | DR PADOLE | 9421896752 | 39030 | 05-Feb-2020 | 09:09 |  |  | 06-Feb-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BULB IS DAMAGE . THEN CHANGE THE NEW BULB NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 31:51:0 | 2
1229 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00022 | BP APPARATUS ANEROID | DR HARSHALI | 9404308287 | 38623 | 28-Jan-2020 | 08:45 |  |  | 28-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER AND VALVE CHANGE THE NEW BLADDER AND VALVE NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 8:15:0 | 1
1230 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00022 | BP APPARATUS ANEROID | DR SOLE | 8956750579 | 27709 | 17-May-2019 | 08:50 |  |  | 22-May-2019 | 11:51 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS HAVING LEAKAGE. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 123:1:0 | 6
1231 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00028 | LOW COST GLUCOMETER | DR NAVHARKAR | 9890230763 | 12206 | 25-Jan-2018 | 09:28 |  |  | 25-Jan-2018 | 17:24 | NOT WORKING | CLEAN THE BLOOD SENSING PART WITH IN USE OF SWAP | Closed | NO INFO |  | 7:56:0 | 1
1232 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00033 | BP APPARATUS ANEROID | DHARMIK JANBANDHU | 9326979013 | 23446 | 29-Dec-2018 | 09:08 |  |  | 01-Jan-2019 | 15:13 | AIR LIKAGE ISSUE | CHECK AND FOUND B. P APPARATUS BLADDER IS LEAK. CHANGED THE NEW BLADDER. B.P APPARATUS IS NOW WORKING. | Closed | NA |  | 78:5:0 | 4
1233 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00034 | BP APPARATUS ANEROID | DR SOLE | 8956750579 | 27708 | 17-May-2019 | 08:50 |  |  | 22-May-2019 | 11:58 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER, BULB AND VALVE. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NA |  | 123:8:0 | 6
1234 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00037 | DIGITAL BP APPARATUS | DR KAMBLE | 9764499775 | 32920 | 23-Sep-2019 | 09:17 |  |  | 25-Sep-2019 | 10:45 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS FAULTY AND DUE WHICH BP COULD NOT BE MEASURED. CHANGED THE DEFECTIVE CUFF WITH THE NEW ONE. DIGITAL BP APPARATUS IS NOW WORKING. | Closed | W16120702344 |  | 49:28:0 | 3
1235 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00037 | DIGITAL BP APPARATUS | DR CHOLE | 8956750579 | 28750 | 18-Jun-2019 | 09:21 |  |  | 22-Jun-2019 | 16:00 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS LEAK AND DUE TO WHICH ERROR IS DISPLAYED ON SCREEN. CHANGED THE OLD CUFF WITH THE NEW ONE. DIGITAL B. P APPARATUS IS NOW WORKING. | Closed | W16120702344 |  | 102:39:0 | 5
1236 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00037 | DIGITAL BP APPARATUS | DHARMIC JANBANDU | 9326979013 | 23406 | 28-Dec-2018 | 08:04 |  |  | 01-Jan-2019 | 12:10 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING. | Closed | W16120702344 |  | 100:6:0 | 5
1237 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00037 | DIGITAL BP APPARATUS | M H NEVARE | 9764373346 | 27530 | 13-May-2019 | 09:01 |  |  | 16-May-2019 | 10:52 | NOT WORKING | CHECK AND FOUND DIGITAL B. P APPARATUS IS WORKING. | Closed | W16120702344 |  | 73:51:0 | 4
1238 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00037 | DIGITAL BP APPARATUS | DR HARSHALI SELOKAR | 9404308287 | 33455 | 05-Oct-2019 | 09:17 |  |  | 07-Oct-2019 | 14:43 | NOT WORKING | CHECK AND FOUND POWER ADAPTER IS NOT WORKING. REPLACED THE DEFECTIVE POWER ADAPTER WITH THE NEW ONE. DIGITAL B. P APPARATUS IS NOW WORKING. | Closed | W16120702344 |  | 53:26:0 | 3
1239 | BHANDARA | PHC SALEBHATA | MH/R1/BHA/PHC/STL/00037 | DIGITAL BP APPARATUS | HARSHALI SELOKAR | 9404308287 | 33857 | 15-Oct-2019 | 08:12 |  |  | 15-Oct-2019 | 10:21 | NOT WORKING | CHECK AND FOUND DIGITAL B. P. APPRATUS POWER SUPPLY ADAPTER IS TOTALLY DAMAGE. CHANGE WITH NEW POWER ADAPTER. NOW DIGITAL B. P. APPRATUS IS WORKING GOOD | Closed | W16120702344 |  | 2:9:0 | 1
1240 | BHANDARA | PHC VIRSHI | MH/R1/BHA/PHC/VRS/00001 | BP APPARATUS ANEROID | TEMBUREKAR NURSE | 9607524122 | 29921 | 13-Jul-2019 | 08:50 |  |  | 16-Jul-2019 | 13:43 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT AND BLADDER IS LEAK. CHANGED THE NEW CUFF AND BLADDER. B. P APPARATUS IS NOW WORKING. | Closed | 28914 | LOW | 76:53:0 | 4
1241 | BHANDARA | PHC VIRSHI | MH/R1/BHA/PHC/VRS/00010 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | TEMBUREKAR NURSE | 9607524122 | 29919 | 13-Jul-2019 | 08:50 |  |  | 16-Jul-2019 | 12:35 | NOT WORKING | CHECK AND FOUND ELEMENT IS BURNT OUT. POWER CORD IS ALSO DEFECTIVE. CHANGED THE OLD DEFECTIVE ELEMENT AND POWER CORD WITH THE NEW ONE. STERILIZER MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 75:45:0 | 4
1242 | BHANDARA | PHC VIRSHI | MH/R1/BHA/PHC/VRS/00011 | BP APPARATUS ANEROID | TEMBUREKAR NURSE | 9607524122 | 29920 | 13-Jul-2019 | 08:50 |  |  | 16-Jul-2019 | 13:46 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT. VALVE IS ALSO LEAK. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | RM105690 | LOW | 76:56:0 | 4
1243 | BHANDARA | PHC VIRSHI | MH/R1/BHA/PHC/VRS/00013 | RADIANT WARMER | B M RAUT | 9765929465 | 33720 | 11-Oct-2019 | 09:11 |  |  | 12-Oct-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN POWER CABLE AND POWER BOARD. NOW MACHINE IS IN GOOD WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | HIGH | 26:49:0 | 2
1244 | BHANDARA | PHC VIRSHI | MH/R1/BHA/PHC/VRS/00021 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR SUBHODH | 9823810860 | 11023 | 27-Dec-2017 | 09:30 |  |  | 27-Dec-2017 | 19:25 | NOT WORKING | REPLACED THE POWER CARD AND COIL | Closed | NO INFO | MEDIUM | 9:55:0 | 1
1245 | BHANDARA | PHC VIRSHI | MH/R1/BHA/PHC/VRS/00022 | Fumigation Machine | TEMBUREKAR NURSE | 9607524122 | 29918 | 13-Jul-2019 | 08:50 |  |  | 16-Jul-2019 | 16:50 | NOT WORKING | CHECK AND FOUND POWER CORD IS DEFECTIVE. REPAIRED THE POWER CORD. FUMIGATION MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 80:0:0 | 4
1246 | BHANDARA | PHC VIRSHI | MH/R1/BHA/PHC/VRS/00022 | Fumigation Machine | DR SUBHODH | 9823810860 | 11024 | 27-Dec-2017 | 09:30 |  |  | 27-Dec-2017 | 19:21 | NOT WORKING | REPLACED THE MOTOR BEARING AND CAPACITOR | Closed | NO INFO | LOW | 9:51:0 | 1
1247 | BHANDARA | PHC VIRSHI | MH/R1/BHA/PHC/VRS/00024 | DEEP FREEZER | DR SUBODH | 9823810860 | 1513 | 08-Feb-2017 | 09:16 |  |  | 07-Mar-2017 | 14:32 | NOT WORKING | WAITING FOR SPARES PARTS | Closed | NO INFO | MEDIUM | 653:16:0 | 28
1248 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00001 | BP APPARATUS ANEROID | DR ANSARI | 9923527997 | 8579 | 31-Oct-2017 | 08:10 |  |  | 03-Nov-2017 | 11:56 | NOT WORKING | RECONNECTED TUBING CONNECTION | Closed | 98135319 | LOW | 75:46:0 | 4
1249 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00001 | BP APPARATUS ANEROID | DURGA BHANDARE | 9763895196 | 26075 | 20-Mar-2019 | 08:40 |  |  | 21-Mar-2019 | 10:44 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS LEAK. CUFF IS ALSO WORN OUT. CHANGED THE NEW BLADDER, BULB AND CUFF. B. P APPARATUS IS NOW WORKING. | Closed | 98135319 | LOW | 26:4:0 | 2
1250 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00004 | BP APPARATUS ANEROID | KALPANA M PARADHI | 9822359271 | 22872 | 12-Dec-2018 | 08:36 |  |  | 18-Dec-2018 | 11:02 | NOT WORKING | CHECK AND FOUND CUFF BLADDER BULB AND VALVE IS MISSING. FIXED THE NEW CUFF BLADDER VALVE AND BULB BP APPRATUS IS NOW WORKING | Closed | NO INFO | LOW | 146:26:0 | 7
1251 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00007 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | DURGA BHANDARE | 9763895196 | 26077 | 20-Mar-2019 | 08:40 |  |  | 21-Mar-2019 | 21:42 | NOT WORKING | CHECK AND FOUND MACHINE IS NOT TURNING ON. POWER SUPPLY FAILURE IS THERE. FIXED THE POWER SUPPLY ISSUE. ADULT WEIGHING MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 37:2:0 | 2
1252 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00009 | SUCTION PUMP PORTABLE, ELECTRONIC | Kalpana Pardi | 9822359271 | 43361 | 23-Jun-2020 | 08:35 |  |  | 23-Jun-2020 | 10:45 | NOT WORKING | CHECK AND FOUND POWER CABLE IS DEFECTIVE. FIXED THE NEW POWER CABLE AND POWER SUPPLY FAILURE ISSUE IS RESOLVED. DESIRED SUCTION PRESSURE IS GENERATED. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 2:10:0 | 1
1253 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00010 | RADIANT WARMER | KALPANA | 9822359271 | 42802 | 02-Jun-2020 | 08:11 |  |  | 03-Jun-2020 | 11:30 | NOT WORKING | CHECK AND FOUND OBSERVING BULB IS FUSED AND BULB HOLDER IS BROKEN. FIXED THE NEW LED BULB AND BULB HOLDER. CUT-OFF OF HEATER IS CHECKED AT LOW AND HIGH BODY TEMPERATURE AND ITS OK. RADIANT WARMER IS NOW WORKING AND ALSO THE OBSERVING LIGHT BULB. | Closed | NO INFO | HIGH | 27:19:0 | 2
1254 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00012 | SUCTION PUMP PORTABLE, ELECTRONIC | DR ANSARI | 9923527997 | 13359 | 01-Mar-2018 | 08:36 |  |  | 03-Mar-2018 | 11:30 | NOT WORKING | REARRANGED THE VACUUM PRESSURE TUBING,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 50:54:0 | 3
1255 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00012 | SUCTION PUMP PORTABLE, ELECTRONIC | Kalpana | 9822359271 | 44011 | 10-Jul-2020 | 08:15 |  |  | 11-Jul-2020 | 10:30 | NOT WORKING | CHECK AND FOUND POWER CABLE IS SHORT AND DUE TO WHICH MACHINE COULD NOT BE TURNED 'ON'. FIXED THE NEW POWER CABLE OF 16A. NOW, MACHINE IS TURNED 'ON' AND DESIRED SUCTION PRESSURE IS GENERATED. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 26:15:0 | 2
1256 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00012 | SUCTION PUMP PORTABLE, ELECTRONIC | KALPANA NURSE | 9822359271 | 17269 | 28-Jun-2018 | 08:45 |  |  | 04-Jul-2018 | 15:45 | NOT WORKING | CHECK THE MACHINE PROBLEM, FOUND LEAKAGE IN TUBING . REPLACED TUBING NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 151:0:0 | 7
1257 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00015 | ADULT WEIGHING MACHINE | Kalpana | 9822359271 | 44012 | 10-Jul-2020 | 08:15 |  |  | 11-Jul-2020 | 11:45 | NOT WORKING | CHECK AND FOUND POWER 'ON/OFF' SWITCH IS FAULTY. FIXED THE NEW POWER 'ON/OFF' SWITCH OF 6A. POWER SUPPLY FAILURE ISSUE IS RESOLVED. ADULT WEIGHING MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 27:30:0 | 2
1258 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00015 | ADULT WEIGHING MACHINE | DR ANSARI | 9923527997 | 13360 | 01-Mar-2018 | 08:38 |  |  | 03-Mar-2018 | 11:20 | NOT WORKING | CALIBRATION PROCESS DONE ON THE MACHINE ,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 50:42:0 | 3
1259 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00015 | ADULT WEIGHING MACHINE | Kalpana Pardi | 9822359271 | 42839 | 03-Jun-2020 | 08:20 |  |  | 03-Jun-2020 | 11:45 | N0T WORKING | CHECK AND FOUND 9V POWER SUPPLY BOARD IS NOT WORKING & ALSO THE POWER 'ON/OFF' SWITCH AND DUE TO WHICH MACHINE IS NOT TURNING ON. FIXED THE NEW 9V POWER SUPPLY BOARD & 2 PIN POWER 'ON/OFF' SWITCH. POWER SUPPLY FAILURE ISSUE IS RESOLVED. ADULT WEIGHING MACHINE IS NOW WORKING. | Closed | NO INFO | LOW | 3:25:0 | 1
1260 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00018 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | DR ANSARI | 9923527997 | 8581 | 31-Oct-2017 | 08:10 |  |  | 03-Nov-2017 | 11:59 | NOT WORKING | REPLACE THE POWER CORD | Closed | NO INFO | MEDIUM | 75:49:0 | 4
1261 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00019 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | DR ANSARI | 9923527997 | 8582 | 31-Oct-2017 | 08:10 |  |  | 03-Nov-2017 | 11:01 | NOT WORKING | REPLACED THE POWER CORD | Closed | NO INFO | MEDIUM | 74:51:0 | 4
1262 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00021 | CENTRIFUGE (BENCH TOP) | MADHURI | 8623941157 | 34499 | 31-Oct-2019 | 09:33 |  |  | 01-Nov-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN POWER CONNECTION AND POWER CABLE IS FAULTY. CHANGE THE NEW POWER CABLE AND THE MAKE POWER CONNECTION. NOW MACHINE IS IN GOOD WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 31:27:0 | 2
1263 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00022 | SYRINGE NEEDLE DESTROYER | DR AG SAMRIT | 8551951555 | 36020 | 02-Dec-2019 | 08:09 |  |  | 02-Dec-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS POWER CABLE LOOSE CONNECTION. MAKE CONNECTION PROPERLY NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | LOW | 8:51:0 | 1
1264 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00022 | SYRINGE NEEDLE DESTROYER | MADHURI | 8623941157 | 35059 | 13-Nov-2019 | 08:10 |  |  | 14-Nov-2019 | 12:00 | NOT WORKING | CHECK THE EQUIPMENT AND FOUND PROBLEM IN FUSE IS BURNED.CHANGE THE NEW FUSE NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER | Closed | NO INFO | LOW | 27:50:0 | 2
1265 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00024 | BP APPARATUS ANEROID | DR ANSARI | 9923527997 | 8580 | 31-Oct-2017 | 08:10 |  |  | 03-Nov-2017 | 11:57 | NOT WORKING | REPLACED THE BLADDER AND BULB CONNECTOR | Closed | NO INFO | LOW | 75:47:0 | 4
1266 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00024 | BP APPARATUS ANEROID | D P SHENDE | 9403926261 | 13091 | 21-Feb-2018 | 09:20 |  |  | 22-Feb-2018 | 11:17 | NOT WORKING | REPLACED THE BULB AND CUFF | Closed | NO INFO | LOW | 25:57:0 | 2
1267 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00024 | BP APPARATUS ANEROID | KALPANA PARADHI | 9822359271 | 17865 | 13-Jul-2018 | 08:04 |  |  | 14-Jul-2018 | 02:30 | NOT WORKING | CHECK THE MACHINE CHANGED THE B.P. BULB , NOW MACHINE WORKING FINE AND HANDOVER TO USER IN WORKING CONDITION. | Closed | NO INFO | LOW | 18:26:0 | 2
1268 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00026 | BP APPARATUS ANEROID | KALPANA PARADHI | 9822359271 | 17866 | 13-Jul-2018 | 08:10 |  |  | 14-Jul-2018 | 02:30 | NOT WORKING | CHECK THE MACHINE PROBLEM REPORTED IN B.P. CUFF OF MACHINE, REPLACED IT , NOW MACHINE WORKING FINE. | Closed | NO INFO |  | 18:20:0 | 2
1269 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00026 | BP APPARATUS ANEROID | D P SHENDE | 9403926261 | 13090 | 21-Feb-2018 | 09:20 |  |  | 22-Feb-2018 | 11:21 | NOT WORKING | REPLACED THE CUFF | Closed | NO INFO |  | 26:1:0 | 2
1270 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00034 | BP APPARATUS ANEROID | DURGA BHANDARE | 9763895196 | 26076 | 20-Mar-2019 | 08:40 |  |  | 21-Mar-2019 | 10:46 | NOT WORKING | CHECK AN FOUND BLADDER, BULB, VALVE AND CUFF IS MISSING. FIXED THE NEW BLADDER, VALVE, CUFF AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | 98135326 |  | 26:6:0 | 2
1271 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00034 | BP APPARATUS ANEROID | DR PRASHANT | 9665441517 | 39088 | 06-Feb-2020 | 08:23 |  |  | 06-Feb-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER AND EXTENSION TUBE IS DAMAGE. THEN CHANGE THE NEW BLADDER AND EXTENSION TUBE. NOW MACHINE IS IN WORKING CONDITION. | Closed | 98135326 |  | 8:37:0 | 1
1272 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00034 | BP APPARATUS ANEROID | DR MESHRAM | 9423673674 | 26838 | 16-Apr-2019 | 08:00 |  |  | 20-Apr-2019 | 16:23 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN AND MERCURY IS SPILL OUT. ALSO VALVE AND BULB IS HAVING LEAKAGE. NEW GLASS TUBE IS FIXED AND MERCURY IS FILLED. ALSO NEW VALVE AND BULB IS INSERTED. B. P APPARATUS IS NOW WORKING. | Closed | 98135326 |  | 104:23:0 | 5
1273 | BHANDARA | PHC WARATHI | MH/R1/BHA/PHC/WRT/00036 | Shadowless Lamp Stand Type | DR PRASHANT | 9665441517 | 39087 | 06-Feb-2020 | 08:23 |  |  | 06-Feb-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS LAMP CONNECTION IS BURNED. THEN MAKE NEW CONNECTION AND DONE PPM. NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 8:37:0 | 1
1274 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00003 | NEBULIZER ELECTRIC | MANISHA NURSE | 9119503209 | 22256 | 20-Nov-2018 | 09:04 |  |  | 24-Nov-2018 | 11:00 | NOT WORKING | CHECK AND FOUND MASK IS DEFECTIVE. AS IT IS CONSUMABLE ITEM SO IT WON'T BE PROVIDED BY COMPANY.HOSPITAL ADMINISTRATION HAS TO ARRANGE IT.AT PRESENT IT IS NOT IN OUR SCOPE .REST OF THE MACHINE IS FUNCTIONAL | Closed | 17174 | MEDIUM | 97:56:0 | 5
1275 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00004 | SUCTION PUMP PORTABLE, ELECTRONIC | MANISHA | 9119503209 | 24569 | 02-Feb-2019 | 08:10 |  |  | 07-Feb-2019 | 14:22 | NOT WORKING | CHECK AND FOUND CONNECTING PIPE THROUGH JAR IS LOOSE. FIXED THE PIPE. PROPER SUCTION PRESSURE IS GENERATED NOW. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 126:12:0 | 6
1276 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00007 | PHOTOTHERAPHY UNIT | MANISHA | 9119503209 | 24570 | 02-Feb-2019 | 08:10 |  |  | 07-Feb-2019 | 11:19 | NOT WORKING | CHECK AND FOUND 4 TUBE LIGHT IS FUSED. CHANGED THE OLD WITH THE NEW TUBELIGHT.PHOTOTHERAPHY UNIT IS NOW WORKING. | Closed | OE0508 | HIGH | 123:9:0 | 6
1277 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00008 | RADIANT WARMER | MANISHA PATIL | 9119503209 | 20358 | 14-Sep-2018 | 09:15 |  |  | 20-Sep-2018 | 11:01 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE ISSUE. RESOLVED THE POWER SUPPLY ISSUE.RADIANT WARMER IS WORKING | Closed | NO INFO | HIGH | 145:46:0 | 7
1278 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00008 | RADIANT WARMER | MANISHA PATALE | 9119503209 | 36505 | 13-Dec-2019 | 09:05 |  |  | 14-Dec-2019 | 17:13 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE ISSUE. FEMALE CABLE CONNECTOR IS FAULTY AS FUSE CONNECTION IS OPEN. FIXED THE FUSE CONNECTION OF FEMALE CABLE CONNECTOR. POWER SUPPY FAILURE RESOLVED. CUT-OFF OF HEATER IS CHECKED. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 32:8:0 | 2
1279 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00008 | RADIANT WARMER | MANISHA | 9119503209 | 24143 | 21-Jan-2019 | 09:20 |  |  | 22-Jan-2019 | 13:33 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE ISSUE. REPAIRED THE POWER SUPPLY ISSUE. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 28:13:0 | 2
1280 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00008 | RADIANT WARMER | DR RAUT | 9637607377 | 13319 | 28-Feb-2018 | 09:12 |  |  | 03-Mar-2018 | 11:57 | NOT WORKING | REPLACED NEW POWER CORD SOCKET,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | HIGH | 74:45:0 | 4
1281 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00010 | CARDIOTOCOGRAPHY MACHINE | DR RAUT | 9637607377 | 13320 | 28-Feb-2018 | 09:12 |  |  | 03-Mar-2018 | 13:03 | NOT WORKING | SOLDERED THE FHR PROBE,NOW HEART RATE AND SOUND IS SHOWS CORRECT MEASURES. | Closed | JA3C100229 | MEDIUM | 75:51:0 | 4
1282 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00010 | CARDIOTOCOGRAPHY MACHINE | MANISHA | 9119503209 | 22607 | 04-Dec-2018 | 08:20 |  |  | 08-Dec-2018 | 11:37 | NOT WORKING | CHECK AND FOUND CARDIOTOCOGRAPHY MACHINE POWER ADAPTER IS BORN CHANGED THE POWER SUPPLY ADAPTER NOW MACHINE IS WORKING | Closed | JA3C100229 | MEDIUM | 99:17:0 | 5
1283 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00010 | CARDIOTOCOGRAPHY MACHINE | MANISHA PATHALE | 9119503209 | 21408 | 17-Oct-2018 | 08:00 |  |  | 20-Oct-2018 | 10:00 | NOT WORKING | CHECK AND FOUND CODDLE GRAPH FEMALE CONNECTOR IS DETACHED AND WENT INSIDE THE MACHINE. FIXED THE FEMALE CONNECTOR OF THE CODDLE GRAPH. CARDIOTOCOGRAPHY MACHINE IS NOW WORKING. | Closed | JA3C100229 | MEDIUM | 74:0:0 | 4
1284 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00010 | CARDIOTOCOGRAPHY MACHINE | MANISHA NURSE | 9119503209 | 25468 | 28-Feb-2019 | 09:01 |  |  | 04-Mar-2019 | 15:42 | NOT WORKING | CHECK AND FOUND CABLE CONNECTOR FROM CONTROLLER BOARD TO DISPLAY BOARD IS LOOSE AND DUE TO WHICH MACHINE WAS NOT TURNING ON. FIXED THE CABLE. CLEANED THE CONTROLLER BOARD. NOW MACHINE IS WORKING. | Closed | JA3C100229 | MEDIUM | 102:41:0 | 5
1285 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00010 | CARDIOTOCOGRAPHY MACHINE | MANISHA NURSE | 9119503209 | 22260 | 20-Nov-2018 | 09:04 |  |  | 25-Nov-2018 | 11:21 | NOT WORKING | CHECK AND FOUND CHARGER OF CARDIOGRAPHY MACHINE IS BURNT.MEDICAL OFFICER HAS ARRANGED THE CHARGER FROM THE VENDOR. NOW MACHINE IS WORKING | Closed | JA3C100229 | MEDIUM | 122:17:0 | 6
1286 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00011 | SHADOWLESS LAMP CEILING TYPE MINOR | MANISHA PATHALE | 9119503209 | 21405 | 17-Oct-2018 | 08:00 |  |  | 20-Oct-2018 | 13:00 | NOT WORKING | CHECK AND FOUND ALL THE HALOGEN BULB IS FUSED.CHANGED THE NEW HALOGEN LIGHT OF 24V,70W RATING OF 4 IN QUANTITY. SHADOW LESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 77:0:0 | 4
1287 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00021 | BINOCULAR MICROSCOPE | MANISHA NURSE | 9119503209 | 30051 | 17-Jul-2019 | 08:06 |  |  | 20-Jul-2019 | 14:15 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING. | Closed | NO INFO | MEDIUM | 78:9:0 | 4
1288 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00027 | AUTOCLAVE HP VERTICAL (2 BIN) DIGITAL | MANISHA | 9119503409 | 27255 | 02-May-2019 | 09:25 |  |  | 03-May-2019 | 14:43 | NOT WORKING | CHECK AND FOUND POWER CABLE IS MISSING. FIXED THE NEW POWER CABLE. AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 29:18:0 | 2
1289 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00028 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | MANISHA PATALE | 9119503209 | 26023 | 19-Mar-2019 | 08:16 |  |  | 23-Mar-2019 | 17:30 | NOT WORKING | CHECK AND FOUND HEATING ELEMENT IS BURNT AND POWER CORD IS MISSING. FIXED THE NEW HEATING ELEMENT AND POWER CORD. AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 105:14:0 | 5
1290 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00030 | BP APPARATUS ANEROID | MANISHA PATHALE | 9119503209 | 21406 | 17-Oct-2018 | 08:00 |  |  | 20-Oct-2018 | 12:00 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK AND CUFF IS WORN OUT. BULB AND VALVE IS MISSING. CHANGED THE NEW BLADDER AND VALVE ALONG WITH CUFF AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | 64042 | LOW | 76:0:0 | 4
1291 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00030 | BP APPARATUS ANEROID | Varsha Kasankar | 8830572267 | 41721 | 18-Apr-2020 | 08:50 |  |  | 18-Apr-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN GLASS TUBE AND BULB IS DAMAGE THEN CHANGE THE NEW GLASS TUBE AND BULB. NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | 64042 | LOW | 8:10:0 | 1
1292 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00037 | ANESTHESIA MACHINE (BOYLES APPARATUS) | MANISHA PATALE | 9119503209 | 36506 | 13-Dec-2019 | 09:05 |  |  | 14-Dec-2019 | 16:09 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING. USER ONLY DEMANDED HANDS ON TRAINING AND IT HAS BEEN GIVEN. | Closed | HT12G1956 | HIGH | 31:4:0 | 2
1293 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00044 | PULSE OXIMETER LINE POWERED | MANISHA | 9119503209 | 24571 | 02-Feb-2019 | 08:10 |  |  | 07-Feb-2019 | 15:20 | NOT WORKING | CHECK AND FOUND MACHINE IS OK AND ALSO THE SPO2 PROBE. | Closed | 174819 | MEDIUM | 127:10:0 | 6
1294 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00045 | Mobile X-ray machine (HF) | MANISHA | 9119503209 | 27409 | 08-May-2019 | 08:49 |  |  | 14-May-2019 | 11:15 | NOT WORKING | CHECK AND FOUND X-RAY TUBE HEAD WAS FAULTY. REPAIRED THE SAME, MACHINE TESTED AND HAND OVER TO THE USER IN PROPER WORKING ORDER. | Closed | CM/L-8208670 | MEDIUM | 146:26:0 | 7
1295 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00045 | Mobile X-ray machine (HF) | MANISHA NURSE | 9119503209 | 30052 | 17-Jul-2019 | 08:09 |  |  | 20-Jul-2019 | 14:30 | NOT WORKING | CHECK AND FOUND OIL LEAKAGE IS THERE FROM BELLOWS. TIGHTENED THE SCREWS OF THE BELLOWS AND FIXED IT. CLEANED THE COLLIMETER CUVETTE AND REMOVED EXCESS OIL. X-RAY IS BEING TAKEN AND FOUND FILM OBTAINED IS PROPER. X-RAY MACHINE IS NOW WORKING. | Closed | CM/L-8208670 | MEDIUM | 78:21:0 | 4
1296 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00045 | Mobile X-ray machine (HF) | MANISHA | 9119503209 | 29114 | 26-Jun-2019 | 09:07 |  |  | 01-Jul-2019 | 16:56 | NOT WORKING | CHECK AND FOUND OIL LEAKAGE FROM BELLOW NUT WHICH WAS STORED OIL. REMOVED ALL STORED OIL. NOW MACHINE CHECKED AND FOUND OK. | Closed | CM/L-8208670 | MEDIUM | 127:49:0 | 6
1297 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00045 | Mobile X-ray machine (HF) | MANISHA PATIL | 9119503209 | 20359 | 14-Sep-2018 | 08:30 |  |  | 20-Sep-2018 | 11:41 | NOT WORKING | CHECK AND FOUND FILAMENT HEAD IS HINGE IS TIGHT AND CANNOT BE ADJUST.ADJUST AND FIXED THE HINGE.OILING AT HEAD IS DONE.X-RAY MACHINE IS WORKING | Closed | CM/L-8208670 | MEDIUM | 147:11:0 | 7
1298 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00045 | Mobile X-ray machine (HF) | MANISHA | 9119503209 | 28377 | 10-Jun-2019 | 08:22 |  |  | 12-Jun-2019 | 16:40 | NOT WORKING | CHECK AND FOUND IMAGE QUALITY WAS VERY POOR.FOUND 10 OHM, 10 WATT RESISTOR BURNT. RECTIFIED THE SAME PROBLEM.NOW MACHINE TESTED AND HAND OVER TO USER IN PROPER WORKING CONDITION. | Closed | CM/L-8208670 | MEDIUM | 56:18:0 | 3
1299 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00046 | BP APPARATUS ANEROID | MANISHA PATHALE | 9119503209 | 21407 | 17-Oct-2018 | 08:00 |  |  | 20-Oct-2018 | 11:00 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS LEAK. CUFF IS WORN OUT. CHANGED THE BULB, VALVE, BLADDER AND CUFF. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 75:0:0 | 4
1300 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00048 | BP APPARATUS ANEROID | MANISHA | 9119503209 | 24144 | 21-Jan-2019 | 09:20 |  |  | 22-Jan-2019 | 12:30 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN. FIXED THE NEW GLASS TUBE. B.P APPARATUS IS NOW WORKING. | Closed | 1509060 | LOW | 27:10:0 | 2
1301 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00061 | BP APPARATUS ANEROID | MANISHA PATLE | 9119503206 | 23843 | 11-Jan-2019 | 09:08 |  |  | 14-Jan-2019 | 17:03 | MARCUARY DAMAGE | CHECK AND FOUND GLASS TUBE IS BROKEN AND MERCURY IS SPILLED. FIXED THE NEW GLASS TUBE. REFILLED THE MERCURY. NOW, B. P APPARATUS IS WORKING. | Closed | NO INFO |  | 79:55:0 | 4
1302 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00061 | BP APPARATUS ANEROID | MANISH PATALE | 9119503209 | 40048 | 25-Feb-2020 | 09:11 |  |  | 27-Feb-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS GLASS TUBE IS DAMAGE AND VALVE IS FAULTY THEN CHANGE THE NEW GLASS TUBE AND VALVE NOW MACHINE IS IN GOOD WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO |  | 50:49:0 | 3
1303 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00061 | BP APPARATUS ANEROID | MANISHA NURSE | 9119503209 | 22259 | 20-Nov-2018 | 09:04 |  |  | 24-Nov-2018 | 11:02 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN WITH SPILL OF MERCURY TUBE IS INSERTED AND MERCURY IS REFILLED .B. P APPARATUS IS NOW WORKING | Closed | NO INFO |  | 97:58:0 | 5
1304 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00063 | AUTOCLAVE HP VERTICAL (2 BIN) DIGITAL | MANISHA PATALE | 9119503209 | 36507 | 13-Dec-2019 | 09:05 |  |  | 14-Dec-2019 | 15:07 | NOT WORKING | CHECK AND FOUND U-TYPE (DOUBLE), 2KW COIL IS NOT PRESENT. MR TO BE RAISED FOR NEW COIL. AS SOON AS NEW COIL WILL BE RECEIVED, IT WILL BE FIXED AND MACHINE WILL BE MADE FUNCTIONAL. AT PRESENT, THIS BREAKDOWN COMPLAINT IS CLOSED. | Closed | NO INFO |  | 30:2:0 | 2
1305 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00063 | AUTOCLAVE HP VERTICAL (2 BIN) DIGITAL | MANISHA PATALE | 9119503209 | 26024 | 19-Mar-2019 | 08:18 |  |  | 23-Mar-2019 | 16:27 | NOT WORKING | CHECK AND FOUND HEATING ELEMENT AND POWER CORD IS MISSING. FIXED THE NEW HEATING ELEMENT AND POWER CORD. AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO |  | 104:9:0 | 5
1306 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00070 | BP APPARATUS ANEROID | MANISHA PATLE | 9119503209 | 36950 | 23-Dec-2019 | 08:00 |  |  | 23-Dec-2019 | 16:51 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT. FIXED THE NEW CUFF. BP APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 8:51:0 | 1
1307 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00084 | BP APPARATUS ANEROID | MANISHA NURSE | 9119503209 | 22257 | 20-Nov-2018 | 09:04 |  |  | 24-Nov-2018 | 11:03 | NOT WORKING | CHECK AND FOUND CUFF BLADDER,BULB AND VALVE IS MISSING GLASS TUBE IS ALSO BROKEN WITH SPILL OF MERCURY . NEW CUFF BLADDER BULB,VALVE AND GLASS TUBE IS INSERTED AND FIXED MERCURY IS REFILLED B. P APPARATUS IS NOW WORKING | Closed | NO INFO |  | 97:59:0 | 5
1308 | BHANDARA | RH ADYAL | MH/R1/BHA/RH/ADL/00085 | BP APPARATUS ANEROID | MANISHA NURSE | 9119503209 | 22258 | 20-Nov-2018 | 09:04 |  |  | 24-Nov-2018 | 11:04 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN WITH SPILL OF MERCURY. NEW GLASS TUBE IS FIXED AND MERCURY IS REFILLED B. P APPARATUS IS NOW WORKING | Closed | NO INFO |  | 98:0:0 | 5
1309 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00005 | BINOCULAR MICROSCOPE | TEJASWINI BANSODE |  | 16838 | 18-Jun-2018 | 08:40 |  |  | 18-Jun-2018 | 15:40 | NOT WORKING | INSPECT AND FOUND DIGIVOLT NOT PROPER SET ON GREEN LINE. SET AND ADJUST SAME. SUGGEST CUSTOMER TO INSTALL STABILIZER TO AVOID VOLTAGE FLUCTUATIONS. | Closed | NO INFO | MEDIUM | 7:0:0 | 1
1310 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00005 | BINOCULAR MICROSCOPE | PRASHANT JAMDHADE | 9405378212 | 10801 | 21-Dec-2017 | 08:40 |  |  | 22-Dec-2017 | 23:39 | NOT WORKING | REPLACED THE ROTOR TUBE WINDING | Closed | NO INFO | MEDIUM | 38:59:0 | 2
1311 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00018 | ELECTRIC BABY WEIGHING SCALE (DIGITAL) | PRASHANT | 9405378212 | 39131 | 06-Feb-2020 | 09:30 |  |  | 07-Feb-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BABY TRAY STRIP IS DAMAGE THEN TRAY ESTABLISHED PROPERLY AND CONNECT NEW SCREW NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 26:30:0 | 2
1312 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00019 | CARDIOTOCOGRAPHY MACHINE | SUSHMA | 7875782719 | 36621 | 17-Dec-2019 | 08:11 |  |  | 18-Dec-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND BATTERY IS DISCHARGE NEED TO CONNECT NEW BATTERY FOR BACK UP. MACHINE IS IN WORKING CONDITION. | Closed | JAC100228 | MEDIUM | 27:49:0 | 2
1313 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00020 | SHADOWLESS LAMP STAND TYPE | TIJARA NURSE | 9637306671 | 32874 | 21-Sep-2019 | 09:01 |  |  | 23-Sep-2019 | 11:29 | NOT WORKING | CHECK ANF FOUND POWER SUPPLY ISSUE IS THERE AS SMPS IS NOT WORKING. REPAIRED THE SMPS AND POWER SUPPLY ISSUE IS RESOLVED. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 50:28:0 | 3
1314 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00022 | RADIANT WARMER | DIPA DESHMUKH | 8390689372 | 28351 | 08-Jun-2019 | 08:02 |  |  | 11-Jun-2019 | 15:00 | NOT WORKING | CHECK AND FOUND TEMPERATURE IS NOT SET AND DUE TO WHICH HEATER DOESN'T TURNED ON. SET THE TEMPERATURE AT NORMAL BODY TEMPERATURE AND CUT OFF OF HEATER IS CHECKED AT HIGH AND LOW BODY TEMPERATURE. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 78:58:0 | 4
1315 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00023 | RADIANT WARMER | USHA B DHURBE | 7181260061 | 4108 | 30-May-2017 | 10:45 |  |  | 03-Jun-2017 | 10:35 | NOT WORKING | REPLACED BULB AND FUSE | Closed | NO INFO | HIGH | 95:50:0 | 5
1316 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00029 | SUCTION PUMP PORTABLE, ELECTRONIC | SUSHILKUMAR WALDE | 9823332723 | 7264 | 18-Sep-2017 | 10:03 |  |  | 21-Sep-2017 | 13:30 | NOT WORKING | REPLACED NEW 6A PLUG TOP,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 75:27:0 | 4
1317 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00031 | CENTRIFUGE (BENCH TOP) | PRASANT | 7743803725 | 35946 | 29-Nov-2019 | 08:30 |  |  | 30-Nov-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS SAMPLE TUBE HOLDER IS MISSING NEED TO CONNECT NEW SAMPLE HOLDER. MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 27:30:0 | 2
1318 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00045 | Mobile X-ray machine (HF) | Jayshree Bansod | 9309852047 | 42438 | 19-May-2020 | 09:50 |  |  | 20-May-2020 | 12:00 | NOT WORKING | CHECKED THE MACHINE AND FOUND PROBLEM IS SHADOW PATCHES IS COMING DURING X-RAY . THEN SCREW IN CALORIMETER IS COME IN FRONT OF THE RAYS THEN ADJUST THAT SCREW. NOW MACHINE IS IN WORKING CONDITION. | Closed | 2B035 | MEDIUM | 26:10:0 | 2
1319 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00045 | Mobile X-ray machine (HF) | DR SUNIL RANGARI | 9403418839 | 33279 | 01-Oct-2019 | 09:05 |  |  | 01-Oct-2019 | 17:20 | NOT WORKING | CHECKED AND FOUND DIGIVOLT NOT WORKING AND IMAGE QUALITY PROBLEM. RECTIFIED THE BOTH. NOW MACHINE WORKING FINE. | Closed | 2B035 | MEDIUM | 8:15:0 | 1
1320 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00045 | Mobile X-ray machine (HF) | BANSOD | 9309852047 | 32641 | 18-Sep-2019 | 08:08 |  |  | 21-Sep-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN EXPLOSER IS NOT WORKING AND ALSO NEED TO PREPARE DARK ROOM PROPERLY. NEED APRONE TO TECHNICIAN FOR TAKING X -RAY. | Closed | 2B035 | MEDIUM | 75:52:0 | 4
1321 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00054 | BP APPARATUS ANEROID | KAMINI DHALKE | 9637606677 | 13868 | 16-Mar-2018 | 09:01 |  |  | 17-Mar-2018 | 12:02 | NOT WORKING | REPLACED NEW BP HEAD STAND HOLDER AND BLADDER, NOW MACHINE IS WORKING CONDITION. . | Closed | NO INFO |  | 27:1:0 | 2
1322 | BHANDARA | RH LAKHANDUR | MH/R1/BHA/RH/LKD/00054 | BP APPARATUS ANEROID | PRASHANT JAMDADE | 9405378212 | 22215 | 19-Nov-2018 | 09:07 |  |  | 26-Nov-2018 | 12:18 | NOT WORKING | CHECK AND FOUND BP APPARATUS STANDING SCREW WAS NOT FIXED. FIXED THE STAND SCREW AND PRESSURE REALISED VALVED CHANGED NOW MACHINE IS WORKING | Closed | NO INFO |  | 171:11:0 | 8
1323 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00002 | SUCTION PUMP PORTABLE, ELECTRONIC | KOKATE NURSE | 8999119170 | 25630 | 07-Mar-2019 | 09:03 |  |  | 09-Mar-2019 | 16:48 | ON OFF BOTTON ISSUE | CHECK AND FOUND POWER ON/OFF SWITCH FAULTY. FIXED THE NEW POWER ON/OFF SWITCH. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 55:45:0 | 3
1324 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00004 | BP APPARATUS ANEROID | MANJUSHA DHANDE | 8806767340 | 5898 | 04-Aug-2017 | 11:25 |  |  | 05-Aug-2017 | 17:33 | NOT WORKING | REPLACED THE BLADDER | Closed | 16342 | LOW | 30:8:0 | 2
1325 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00004 | BP APPARATUS ANEROID | M M DHANDE | 8806767346 | 12911 | 15-Feb-2018 | 08:47 |  |  | 15-Feb-2018 | 22:04 | NOT WORKING | REPLACED THE BULB | Closed | 16342 | LOW | 13:17:0 | 1
1326 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00005 | CENTRIFUGE (BENCH TOP) | ASHISH | 8308931315 | 40162 | 27-Feb-2020 | 08:16 |  |  | 28-Feb-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS BLOOD SAMPLE TUBE HOLDER IS MISSING THEN CONNECT THE NEW BLOOD SAMPLE TUBE HOLDER NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | NO INFO | MEDIUM | 27:44:0 | 2
1327 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00007 | BINOCULAR MICROSCOPE | ASHISH | 8308931315 | 38612 | 28-Jan-2020 | 08:28 |  |  | 29-Jan-2020 | 12:56 | NOT WORKING | CHECK AND FOUND OIL IMMERSION LENS OF 100X/1.25 IS DEFECTIVE AND DUE TO WHICH THERE IS PROBLEM OF HAZINESS WHILE VIEWING THE SLIDES. FIXED THE NEW OIL IMMERSION LENS. BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | 140368026 | MEDIUM | 28:28:0 | 2
1328 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00007 | BINOCULAR MICROSCOPE | ASHISH | 8308931315VVVV | 41149 | 20-Mar-2020 | 09:14 |  |  | 21-Mar-2020 | 11:09 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. FIXED THE NEW HALOGEN BULB OF 6A,20W. MICROSCOPE IS NOW WORKING. | Closed | 140368026 | MEDIUM | 25:55:0 | 2
1329 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00008 | 3 PART HAEMATOLOGY ANALYZER | P N BHUTE | 9421808987 | 5799 | 01-Aug-2017 | 11:05 |  |  | 02-Aug-2017 | 19:17 | NOT WORKING | ENTER THE REAGENT NUMBER THEN WORKING PROPERLY | Closed | NO INFO | MEDIUM | 32:12:0 | 2
1330 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00008 | 3 PART HAEMATOLOGY ANALYZER | BHUTE | 9421808987 | 3200 | 19-Apr-2017 | 10:45 |  |  | 22-Apr-2017 | 02:30 | NOT WORKING | CLEANED RBC AND WBC CHAMBERS APERTURE,NOW PLT VALUE IS NORMAL MACHINE HAND OVER WORKING CONDITION. | Closed | NO INFO | MEDIUM | 63:45:0 | 4
1331 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00010 | BINOCULAR MICROSCOPE | SANJAY DHURAVE | 9637923746 | 24443 | 30-Jan-2019 | 08:35 |  |  | 30-Jan-2019 | 15:26 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. CHANGED THE NEW HALOGEN BULB OF RATING 6V,10 W. BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | CM/L 9010250 | MEDIUM | 6:51:0 | 1
1332 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00022 | Mobile X-ray machine (HF) | YOGENDRA | 9423687866 | 4638 | 20-Jun-2017 | 09:30 | 23-Jun-2017 | 13:35 | 23-Jun-2017 | 13:35 | NOT WORKING | RESOLDER THE RELAY CABLES | Closed | 4B053 | MEDIUM | 76:5:0 | 4
1333 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00022 | Mobile X-ray machine (HF) | SUNDE | 9423687866 | 5106 | 07-Jul-2017 | 09:07 | 08-Jul-2017 | 10:53 | 11-Jul-2017 | 13:58 | NOT WORKING | SUCCESSFULLY COMPLETE VOLTAGE CALIBRATION | Closed | 4B053 | MEDIUM | 100:51:0 | 5
1334 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00022 | Mobile X-ray machine (HF) | Y L PUNDE | 9423687866 | 14968 | 17-Apr-2018 | 09:00 |  |  | 17-Apr-2018 | 14:29 | NOT WORKING | SET THE DIGIVOLTS AS PER THE SPECIFICATION,NOW MACHINE IS WORKING CONDITION. | Closed | 4B053 | MEDIUM | 5:29:0 | 1
1335 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00022 | Mobile X-ray machine (HF) | Y L PUNDE | 9423687866 | 34114 | 19-Oct-2019 | 09:06 |  |  | 19-Oct-2019 | 17:49 | NOT WORKING | CHECKED AND FOUND IMAGE QUALITY PROBLEM. DONE KV/MA CALIBRATION. BUT ON 100 MA TUBE NOT WORKING PROPERLY. SUGGEST TO USER TAKE ON X-RAY 80 MA. MACHINE WORKING PROPERLY. | Closed | 4B053 | MEDIUM | 8:43:0 | 1
1336 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00023 | DIGITAL BP APPARATUS | PRATIMA SONEWANE | 9766868045 | 24541 | 01-Feb-2019 | 09:09 |  |  | 04-Feb-2019 | 11:45 | REQUIRED PROB & NOT PROPER WORKING | CHECK AND FOUND NIBP CUFF IS NOT PRESENT. FIXED THE NEW NIBP CUFF. B.P APPARATUS(DIGITAL) IS NOW WORKING. | Closed | NO INFO | LOW | 74:36:0 | 4
1337 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00027 | CENTRIFUGE (BENCH TOP) | RUPALI WAGHMARE | 9765433044 | 5752 | 31-Jul-2017 | 09:05 | 01-Aug-2017 | 14:24 | 04-Aug-2017 | 12:41 | NOT WORKING | REWINDING THE MOTOR AND REPLACE CABLE | Closed | NO INFO | MEDIUM | 99:36:0 | 5
1338 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00030 | RADIANT WARMER | P S BHUTE | 9421808987 | 10159 | 07-Dec-2017 | 08:30 |  |  | 08-Dec-2017 | 15:34 | NOT WORKING | RECONNECT THE SWITCH PROPERLY | Closed | NO INFO | HIGH | 31:4:0 | 2
1339 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00030 | RADIANT WARMER | ASHWINI KOKATE | 8007916079 | 21969 | 12-Nov-2018 | 08:00 |  |  | 16-Nov-2018 | 16:54 | NOT WORKING | CHECK AND FOUND SMPS BOARD ISSUE AS NO POWER SUPPLY IS GENERATED. REPAIRED THE SMPS BOARD WITH PROPERLY CONNECTING THE LINE AND PHASE. POWER SUPPLY ISSUE IS RESOLVED. RADIANT WARMER IS WORKING. | Closed | NO INFO | HIGH | 104:54:0 | 5
1340 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00033 | FETAL DOPPLER | ASHWINI KOKATE | 8007916079 | 21973 | 12-Nov-2018 | 08:02 |  |  | 16-Nov-2018 | 17:06 | NOT WORKING | CHECK AND FOUND NO SOUND OF HEART BEAT IS HEARD. MUTE BUTTON IS TURNED ON. TURNED 'OFF' THE MUTE BUTTON. NOW CLEARLY HEART BEAT IS HEARD. MEASURED THE HEART BEAT WITH PROPERLY SOUND BEING HEARD. FETAL DOPPLER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 105:4:0 | 5
1341 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00042 | AUTOCLAVE HP HORIZONTAL | PRITI MANKAR | 8446142584 | 34065 | 18-Oct-2019 | 09:10 |  |  | 18-Oct-2019 | 15:44 | NOT WORKING | CHECK AND FOUND THERE IS LEAKAGE OF WATER THROUGH COIL HOLE. FIXED THE COIL TIGHTLY AND REFIXED THE CONNECTION. CHECKED THE MACHINE BY DOING ONE ROUND AUTOCLAVE OF DRUMS AND NO LEAKAGE IS FOUND FURTHER. AUTOCLAVE MACHINE IS NOW WORKING. | Closed | 453 | MEDIUM | 6:34:0 | 1
1342 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00042 | AUTOCLAVE HP HORIZONTAL | PRITI | 8446142584 | 38354 | 23-Jan-2020 | 09:25 |  |  | 25-Jan-2020 | 14:43 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT AND DUE TO WHICH WATER COULD NOT BE BOILED. CHANGED THE OLD DEFECTIVE COIL WITH THE NEW ONE. NOW, WATER IS BOILED AND PROPER PROCESS OF AUTOCLAVE IS BEING DONE. HORIZONTAL AUTOCLAVE IS NOW WORKING. | Closed | 453 | MEDIUM | 53:18:0 | 3
1343 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00042 | AUTOCLAVE HP HORIZONTAL | PRAVIN BHUTE | 9421808987 | 3687 | 12-May-2017 | 10:45 | 13-May-2017 | 14:00 | 16-May-2017 | 02:30 | NOT WORKING | AFTER REPLACED 6KW COIL NOW HEATING DONE PROPERLY,MACHINE HANDOVER GOOD WORKING CONDITION | Closed | 453 | MEDIUM | 87:45:0 | 5
1344 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00042 | AUTOCLAVE HP HORIZONTAL | PREETI MANKAR | 9446142584 | 25972 | 18-Mar-2019 | 09:04 |  |  | 23-Mar-2019 | 15:25 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT OUT. CHANGED THE NEW HEATING COIL OF 9KW.AUTOCLAVE MACHINE IS NOW WORKING. | Closed | 453 | MEDIUM | 126:21:0 | 6
1345 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00042 | AUTOCLAVE HP HORIZONTAL | KOKATE | 8999119970 | 31825 | 26-Aug-2019 | 09:10 |  |  | 28-Aug-2019 | 15:35 | NOT WORKING | CHECK AND FOUND WATER LEAKAGE ISSUE AND HEATER IS NOT TURNING 'ON'. WATER ISSUE PROBLEM IS BEING RESOLVED. OUT OF 3 PHASE FROM MAIN SUPPLY BOARD, ONE PHASE IS NOT COMING AND DUE TO WHICH CONTACTOR SWITCH IS NOT WORKING AND HEATER IS NOT TURNING 'ON'. DIRECTED THE HOSPITAL ADMINISTRATION TO REPAIR 3-PHASE ISSUE FROM MAIN SUPPLY BOARD BY ELECTRICIAN. | Closed | 453 | MEDIUM | 54:25:0 | 3
1346 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00042 | AUTOCLAVE HP HORIZONTAL | PRITI MANKAR | 8446142584 | 24762 | 07-Feb-2019 | 08:38 |  |  | 13-Feb-2019 | 15:40 | NOT WORKING | CHECK AND FOUND HEATING ELEMENT IS BURNT. CHANGED THE NEW HEATING ELEMENT OF 9KW.HORIZONTAL AUTOCLAVE IS NOW WORKING. | Closed | 453 | MEDIUM | 151:2:0 | 7
1347 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00042 | AUTOCLAVE HP HORIZONTAL | PREETI | 8446142584 | 40630 | 07-Mar-2020 | 09:21 |  |  | 09-Mar-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS HEATING COIL IS BURNED DUE TO USER FAULT NEED TO GIVEN TRAINING TO USER AT DH BHANDARA. MACHINE IS IN NOT WORKING CONDITION AND CALL WILL BE CLOSED. | Closed | 453 | MEDIUM | 50:39:0 | 3
1348 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00042 | AUTOCLAVE HP HORIZONTAL | D P MANKAR | 8446142584 | 40251 | 29-Feb-2020 | 09:02 |  |  | 02-Mar-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN HEATING COIL IS BURNED DUE TO LOW WATER LEVEL USED BY ATTENDANT. THEN CHANGE THE NEW HEATING COIL . NOW MACHINE IS IN WORKING CONDITION. NEED TO MAINTAIN PROPER WATER LEVEL AT WATER TANK BY ATTENDANT. | Closed | 453 | MEDIUM | 50:58:0 | 3
1349 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00048 | BP APPARATUS ANEROID | KOKATE NURSE | 8999119170 | 24305 | 25-Jan-2019 | 09:04 |  |  | 30-Jan-2019 | 14:29 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS LEAK. CUFF IS ALSO WORN OUT. FIXED THE NEW BLADDER, VALVE, BULB AND CUFF. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 125:25:0 | 6
1350 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00049 | BP APPARATUS ANEROID | KOKATE NURSE | 8999119170 | 24306 | 25-Jan-2019 | 09:04 |  |  | 30-Jan-2019 | 13:32 | NOT WORKING | CHECK AND FOUND NO BLADDER, CUFF, VALVE AND BULB IS PRESENT. GLASS TUBE IS ALSO BROKEN. FIXED THE NEW GLASS TUBE, BLADDER, CUFF, VALVE AND BULB. MERCURY IS ALSO REFILLED. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 124:28:0 | 6
1351 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00058 | SUCTION PUMP PORTABLE, ELECTRONIC | P S BHUTE | 9421808987 | 10160 | 07-Dec-2017 | 08:30 |  |  | 08-Dec-2017 | 15:29 | NOT WORKING | CLEAN THE DUST PARTICLES | Closed | NO INFO |  | 30:59:0 | 2
1352 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00058 | SUCTION PUMP PORTABLE, ELECTRONIC | KOKATE NURSE | 8999119170 | 25669 | 08-Mar-2019 | 09:40 |  |  | 09-Mar-2019 | 18:50 | NOT WORKING | CHECK AND FOUND NO SUCTION PRESSURE IS GENERATED. FILLED THE OIL. NOW PROPER SUCTION PRESSURE IS GENERATED. SUCTION MACHINE IS NOW WORKING. | Closed | NO INFO |  | 33:10:0 | 2
1353 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00058 | SUCTION PUMP PORTABLE, ELECTRONIC | P S BHUTE | 9421808987 | 12871 | 14-Feb-2018 | 09:25 |  |  | 15-Feb-2018 | 22:01 | NOT WORKING | REFILL THE 2T OIL | Closed | NO INFO |  | 36:36:0 | 2
1354 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00069 | FETAL DOPPLER | KOKATE | 8999119970 | 31826 | 26-Aug-2019 | 09:10 |  |  | 30-Aug-2019 | 16:35 | NOT WORKING | CHECK AND FOUND DOPPLER PROBE IS FAULTY AND DUE TO WHICH HEART BEAT COULD NOT BE MEASURED AND HEARD. CHANGED THE OLD DEFECTIVE PROBE WITH THE NEW ONE. FETAL DOPPLER IS NOW WORKING. | Closed | NA |  | 103:25:0 | 5
1355 | BHANDARA | RH LAKHANI | MH/R1/BHA/RH/LKN/00069 | FETAL DOPPLER | ASHWINI KOKATE | 8007916079 | 21972 | 12-Nov-2018 | 08:02 |  |  | 19-Nov-2018 | 15:48 | NOT WORKING | CHECK AND FOUND DOPPLER PROBE FAULTY. REPAIRED THE DOPPLER PROBE BY CHANGING THE NEW PROBE. FOETAL DOPPLER IS NOW WORKING. | Closed | NA |  | 175:46:0 | 8
1356 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00002 | BP APPARATUS ANEROID | UDDHAV KOLI | 7276562246 | 39520 | 14-Feb-2020 | 08:55 |  |  | 14-Feb-2020 | 13:47 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS HAVING LEAKAGE. FIXED THE NEW BLADDER AND BULB. BP APPARATUS IS NOW WORKING. | Closed | 254010 | LOW | 4:52:0 | 1
1357 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00002 | BP APPARATUS ANEROID | PATWAR NURSE | 8412877590 | 5932 | 05-Aug-2017 | 10:41 |  |  | 07-Aug-2017 | 13:30 | NOT WORKING | REPLACED NEW BLADDER AND EXTENSION TUBE,NOW MACHINE IS WORKING CONDITION | Closed | 254010 | LOW | 50:49:0 | 3
1358 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00005 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | PATWAR NURSE | 8412877590 | 5933 | 05-Aug-2017 | 10:41 |  |  | 07-Aug-2017 | 13:20 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB WITH CONNECTOR,NOW MACHINE IS WORKING CONDITION | Closed | 1090 | LOW | 50:39:0 | 3
1359 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00006 | BP APPARATUS ANEROID | Uddhav Koli | 9518927268 | 43264 | 19-Jun-2020 | 08:42 |  |  | 19-Jun-2020 | 11:00 | NOT WORKING | CHECK AND FOUND BLADDER & BULB IS HAVING LEAKAGE. FIXED THE NEW BLADDER AND BULB. BP APPARATUS IS NOW WORKING. | Closed | RM1000167 | LOW | 2:18:0 | 1
1360 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00006 | BP APPARATUS ANEROID | UDDHAV KOLI | 7276562246 | 39519 | 14-Feb-2020 | 08:50 |  |  | 14-Feb-2020 | 14:50 | NOT WORKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS HAVING LEAKAGE. FIXED THE NEW BLADDER, VALVE AND BULB. BP APPARATUS IS NOW WORKING. | Closed | RM1000167 | LOW | 6:0:0 | 1
1361 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00006 | BP APPARATUS ANEROID | DR SHILPA | 8530110822 | 34453 | 30-Oct-2019 | 08:10 |  |  | 30-Oct-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER AND BULB. CHANGE NEW BLADDER AND BULB NOW MACHINE IS IN GOOD WORKING CONDITION HAND OVER TO USER. | Closed | RM1000167 | LOW | 8:50:0 | 1
1362 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00006 | BP APPARATUS ANEROID | PATWAR NURSE | 8412877590 | 5934 | 05-Aug-2017 | 10:45 |  |  | 07-Aug-2017 | 01:35 | NOT WORKING | REPLACED NEW BLADDER, CUFF AND BULB WITH CONNECTOR. NOW MACHINE IS WORKING CONDITION | Closed | RM1000167 | LOW | 38:50:0 | 3
1363 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00016 | RADIANT WARMER | UDHAV KOLI | 7276562246 | 26178 | 25-Mar-2019 | 08:00 |  |  | 01-Apr-2019 | 17:22 | NOT WORKING | CHECK AND FOUND BULB IS FUSED AND ON/OFF SWITCH FAULTY. CHANGED ALL THE SWITCH AND BULB. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 177:22:0 | 8
1364 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00016 | RADIANT WARMER | Uddhav Koli | 9518927268 | 44215 | 15-Jul-2020 | 08:45 |  |  | 15-Jul-2020 | 14:45 | NOT WORKING | CHECK AND FOUND OUT OF 6 SPOT LIGHT 4 IS NOT WORKING AS ITS 2 PIN 'ON/OFF' SWITCH IS FAULTY.FIXED THE NEW 2 PIN 'ON/OFF' SWITCHES. ALL THE ENTIRE 6 'ON/OFF' SWITCH IS REPLACED WITH THE NEW ONE. ALL THE SPOT LIGHT IS NOW WORKING. RADIANT WARMER IS NOW FUNCTIONAL. | Closed | NO INFO | HIGH | 6:0:0 | 1
1365 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00017 | FETAL DOPPLER | Uddhav Koli | 9518927268 | 43263 | 19-Jun-2020 | 08:41 |  |  | 19-Jun-2020 | 10:30 | NOT WORKING | PROBLEM REPORTED OF LOW VOLUME OF HEART SOUND AND INACCURATE READING. CHECK AND FOUND DOPPLER PROBE IS DEFECTIVE. CHANGED THE OLD PROBE WITH THE NEW ONE. NOW,VOLUME IS AUDIBLE WITH ACCEPTABLE READING AS CONFIRMED WHILE ANALYSING READINGS OF 3-4 PATIENT. FETAL DOPPLER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 1:49:0 | 1
1366 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00017 | FETAL DOPPLER | UDHAV KOLI | 7276562246 | 26179 | 25-Mar-2019 | 08:00 |  |  | 01-Apr-2019 | 14:53 | PROB ISSUE | CHECK AND FOUND PROBE IS FAULTY. CHANGED THE OLD WITH THE NEW PROBE. NOW FETAL DOPPLER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 174:53:0 | 8
1367 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00017 | FETAL DOPPLER | ROSHANI GHODAPE | 9503305717 | 22145 | 16-Nov-2018 | 09:07 |  |  | 17-Nov-2018 | 14:26 | NOT WORKING | CHECK AND FOUND DOPPLER PROBE FAULTY. REPAIRED THE DOPPLER PROBE. FOETAL DOPPLER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 29:19:0 | 2
1368 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00024 | 3 PART HAEMATOLOGY ANALYZER | DR N V GAIDANE | 8805242608 | 3936 | 22-May-2017 | 10:22 |  |  | 26-May-2017 | 12:15 | NOT WORKING | CLEANING PROCESS IS DONE | Closed | 201ESOH03482 | MEDIUM | 97:53:0 | 5
1369 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00024 | 3 PART HAEMATOLOGY ANALYZER | MADHURI | 8806717976 | 9205 | 16-Nov-2017 | 09:55 |  |  | 20-Nov-2017 | 14:29 | NOT WORKING | REPLACED NEW WBC CHAMBER WITH APERTURE,AND ALSO WE REQUESTED TO PURCHASE NEW SOLUTIONS FORM USER. | Closed | 201ESOH03482 | MEDIUM | 100:34:0 | 5
1370 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00026 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | MADHURI NARAD | 8806717976 | 10813 | 21-Dec-2017 | 09:10 |  |  | 26-Dec-2017 | 11:10 | NOT WORKING | REPLACED THE LAMP ASSEMBLY BOARD | Closed | NO INFO | MEDIUM | 122:0:0 | 6
1371 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00038 | SHADOWLESS LAMP CEILING TYPE MINOR | Uddhav Koli | 9518927268 | 43687 | 30-Jun-2020 | 09:10 |  |  | 30-Jun-2020 | 11:45 | NOT WORKING | CHECK AND FOUND WIRING FROM POWER SUPPLY BOARD TO SHADOWLESS LAMP IS SHORT. REWIRING FROM MAIN SUPPLY BOARD TO SHADOWLESS LAMP IS DONE. NEW BULB IS FIXED ALONG WITH THE HOLDER. SHADOWLESS LAMP CEILING TYPE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 2:35:0 | 1
1372 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00045 | PULSE OXIMETER LINE POWERED | PATWAR NURSE | 8412877590 | 7729 | 03-Oct-2017 | 09:00 |  |  | 06-Oct-2017 | 11:30 | NOT WORKING | REPLACED NEW SPO2 PROBE,NOW MACHINE IS WORKING CONDITION | Closed | NO INFO |  | 74:30:0 | 4
1373 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00051 | NEBULIZER ELECTRIC | UDDHAV KOLI | 7276562246 | 27198 | 30-Apr-2019 | 08:35 |  |  | 03-May-2019 | 16:00 | NOT WORKING | CHECK AND FOUND MOTOR IS JAMMED. OILING OF MOTOR IS DONE. NOW, MOTOR IS ROTATING NORMALLY AND ALSO GENERATING FUMES.NEBULIZER MACHINE IS NOW WORKING. | Closed | NO INFO |  | 79:25:0 | 4
1374 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00055 | ECG UNIT 3 CHANNEL PORTABLE | Uddhav Koli | 9518927268 | 43709 | 30-Jun-2020 | 09:44 |  |  | 30-Jun-2020 | 13:45 | NOT WORKING | CHECK AND FOUND ECG LEAD CABLE IS RAT BITTEN AND IS FAULTY. CLAMP IS BROKEN AND CHEST BULB IS ALSO RAT BITTEN AND IS DEGRADED. FIXED THE NEW 10 LEAD ECG CABLE, ECG RUBBER BULB WITH METAL ELECTRODES AND CLAMP. ECG IS BEING TAKEN OF PATIENT AND GRAPH OBTAINED IS ACCEPTABLE. ECG MACHINE IS NOW WORKING. | Closed | NO INFO |  | 4:1:0 | 1
1375 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00056 | ECG UNIT 3 CHANNEL PORTABLE | Uddhav Koli | 9518927268 | 43265 | 19-Jun-2020 | 08:41 |  |  | 19-Jun-2020 | 11:30 | NOT WORKING | PROBLEM REPORTED OF JAMMED GRAPH SLOT WITH PRINTING ERROR. OPENED THE GRAPH SLOT. CLEANED IT. HANDS-ON-TRAINING IS BEING GIVEN TO STAFF NURSES ON HOW TO USE THE MACHINE, HOW TO CHANGE GRAPH PAPER WHEN IT'S EMPTY AND HOW TO CLEAN THE ECG BULB AND CLAMP AFTER EVERY USE. ECG IS BEING TAKEN OF PATIENT AND GRAPH IS OBTAINED WITHOUT ANY ERROR. ECG MACHINE IS NOW WORKING. | Closed | NO INFO |  | 2:49:0 | 1
1376 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00064 | BP APPARATUS ANEROID | Uddhav Koli | 9518927268 | 43266 | 19-Jun-2020 | 08:41 |  |  | 19-Jun-2020 | 13:30 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BULB AND VALVE. CUFF IS ALSO WORN OUT. FIXED THE NEW BULB, VALVE AND CUFF. BP APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 4:49:0 | 1
1377 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00064 | BP APPARATUS ANEROID | UDDHAV KOLI | 7276562246 | 39518 | 14-Feb-2020 | 08:50 |  |  | 14-Feb-2020 | 10:42 | NOT WOEKING | CHECK AND FOUND BLADDER, VALVE AND BULB IS HAVING LEAKAGE. CUFF IS WORN OUT. FIXED THE NEW BLADDER, VALVE, BULB AND CUFF. BP APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 1:52:0 | 1
1378 | BHANDARA | RH MOHADI | MH/R1/BHA/RH/MOD/00083 | XRAY ABOVE 500mA (HF) | CHINTAMANI | 9403360423 | 37988 | 16-Jan-2020 | 09:04 |  |  | 26-Feb-2020 | 10:00 | NOT WORKING | AFTER CHECKED AND FOUND MACHINE WAS SHOWS INVERTER FAULT ERROR. FOUND HV TANK FALTY. REPLACE THE SAME WITH NEW ONE. MACHINE TESTED AND HANDOVER USER IN WORKING CONDITION. | Closed | 2K16120201-BX/HF |  | 984:56:0 | 42
1379 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00001 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | DR JAKKAL | 8888863689 | 12884 | 15-Feb-2018 | 08:16 |  |  | 17-Feb-2018 | 11:08 | NOT WORKING | CLEANING PROCESS IS DONE | Closed | NO INFO | MEDIUM | 50:52:0 | 3
1380 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00001 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | SAPNA SISTER | 9405986554 | 17670 | 09-Jul-2018 | 08:40 |  |  | 29-Oct-2018 | 11:17 | NOT WORKING | RW IS BEING PREPARED. NOW THE CALL IS BEING CLOSED. AFTER RECEIVING THE SPARES AGAIN THE CALL WILL BE REOPENED AND SPARES WILL BE FIXED. | Closed | NO INFO | MEDIUM | 2690:37:0 | 113
1381 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | MR NARENDRA | 9763233601 | 1793 | 21-Feb-2017 | 11:00 |  |  | 24-Feb-2017 | 12:30 | NOT WORKING | DONE PREVENTIVE MAINTENANCE AND CLEANING PROCESS,NOW MACHINE IS WORKING FINE | Closed | 801604080E | MEDIUM | 73:30:0 | 4
1382 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 38333 | 23-Jan-2020 | 08:56 |  |  | 23-Jan-2020 | 17:00 | NOT WORKING | CHECKED AND FOUND MACHINE SHOWING INACCURATE READINGS THEN DONE PPM AND CALIBRATION IS SUCCESSFUL NOW MACHINE IS WORKING IN GOOD CONDITION. | Closed | 801604080E | MEDIUM | 8:4:0 | 1
1383 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | KHADE NURSE | 8983299433 | 17097 | 23-Jun-2018 | 09:37 |  |  | 25-Jun-2018 | 17:30 | NOT WORKING | CHECK AND FOUND ONE VALVE IS NOT WORKING. CLEAN THE VALVE, NOW MACHINE WORKING SATISFACTORY. | Closed | 801604080E | MEDIUM | 55:53:0 | 3
1384 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 26019 | 19-Mar-2019 | 08:08 |  |  | 25-Mar-2019 | 11:30 | NOT WORKING | CHECK AND FOUND VALVE NO.1 & MAIN BORAD OF MACHINE IS NOT WORKING PROPERLY. NEED TO CHANGE ABOVE BOTH PART, NOW CHANGE THE SAME PART WITH NEW ONE NOW MACHINE IS WORKING GOOD. | Closed | 801604080E | MEDIUM | 147:22:0 | 7
1385 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 27274 | 03-May-2019 | 09:41 |  |  | 07-May-2019 | 08:19 | NOT WORKING | CHECK AND FOUND ASPIRATION PROBE IS NOT WORKING PROPERLY .DO THE LUBRICANT OF ROTOR BELT AND NOW MACHINE IS WORKING GOOD | Closed | 801604080E | MEDIUM | 94:38:0 | 5
1386 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | KHADE NURSE | 9561393453 | 41009 | 17-Mar-2020 | 09:17 |  |  | 18-Mar-2020 | 15:33 | NOT WORKING | CHECK AND FOUND HIGH WBC AND PLATELETE COUNT ON EVERY SAMPLE TEST. DONE THE BASIC CLEANING AND VALVE NUMBER 11 IS BEING CLEANED. NOW, SATISFACTORY TEST RESULT IS BEING OBTAINED. 3-PART HAEMATOLOGY ANALYZER IS NOW WORKING. | Closed | 801604080E | MEDIUM | 30:16:0 | 2
1387 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | RAHUL | 9021601692 | 16588 | 11-Jun-2018 | 09:16 |  |  | 16-Jun-2018 | 14:16 | NOT WORKING | ELECTRODE WIRE IS NOT CONNECTED. RESOLDERED THE WIRE . | Closed | 801604080E | MEDIUM | 125:0:0 | 6
1388 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | MULMULE | 9422144593 | 25846 | 14-Mar-2019 | 08:15 |  |  | 16-Mar-2019 | 12:04 | NOT WORKING | CHECK AND FOUND MACHINE VALVE NO. 6 IS NOT WORKING NEED TO CHANGE THE VALVE. NOW CHANGE THE VALVE NO. 6 MACHINE IS WORKING GOOD | Closed | 801604080E | MEDIUM | 51:49:0 | 3
1389 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | KHADE NURSE | 8983299433 | 26947 | 20-Apr-2019 | 09:15 |  |  | 26-Apr-2019 | 10:03 | NOT WORKING | CHECK AND FOUND MACHINE SOME PARAMETER IS NOT SHOWING LIKE NEUTROPHILS. GIVE IT CLEANING AND BACK FLUSH . NOW MACHINE IS WORKING GOOD | Closed | 801604080E | MEDIUM | 144:48:0 | 7
1390 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | MULMULE A L | 9422144593 | 36386 | 10-Dec-2019 | 09:17 |  |  | 10-Dec-2019 | 16:59 | NOT WORKING | CHECK AND FOUND RT HEMAT SOFTWARE IS CORRUPTED. INSTALLED THE NEW RT HEMAT SOFTWARE. CHECK THE READING OF THE SAMPLE AND ITS SATISFACTORY. 3-PART HAEMATOLOGY ANALYZER IS NOW WORKING. | Closed | 801604080E | MEDIUM | 7:42:0 | 1
1391 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | ARUN MULMULE | 9422144593 | 40861 | 13-Mar-2020 | 09:05 |  |  | 13-Mar-2020 | 14:55 | NOT WORKING | CHECK AND FOUND THERE IS HIGH PLATELETE COUNT AND RBC COUNT DURING EVERY TEST. VALVE NUMBER 10 IS NOT WORKING AND CLEANED THE SAME. PROBLEM RECTIFIED AND RESULT OBTAINED IS DESIRABLE WITH NORMAL READING. 3-PART HAEMATOLOGY ANALZER IS NOW WORKING. | Closed | 801604080E | MEDIUM | 5:50:0 | 1
1392 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | RAHUL MATALE | 9021601692 | 19405 | 21-Aug-2018 | 08:35 |  |  | 28-Aug-2018 | 11:22 | NOT WORKING | CHECK AND FOUND 3PART HEMOTOLOGY ANALYSER 7.NO.CHEMBER IS NOT WORKING.REPLACE SAME WITH NEW ONE.NOW MACHINE IS WORKING | Closed | 801604080E | MEDIUM | 170:47:0 | 8
1393 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | RAHUL MATALE | 9021601692 | 22081 | 15-Nov-2018 | 08:08 |  |  | 21-Nov-2018 | 09:10 | NOT WORKING | CHECK AND FOUND MACHINE DRAINING IS NOT WORKING. DRAIN THE MACHINE PIPE.NOW MACHINE IS WORKING PROPERLY | Closed | 801604080E | MEDIUM | 145:2:0 | 7
1394 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | RAHUL MATLE | 9021601692 | 25127 | 18-Feb-2019 | 08:27 |  |  | 22-Feb-2019 | 11:57 | NOT WORKING | CHECK AND FOUND IN MACHINE CPU BOARD IS DEFECTIVE. CHANGE THE SAME WITH NEW ONE, NOW MACHINE IS WORKING | Closed | 801604080E | MEDIUM | 99:30:0 | 5
1395 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | KHADE SISTER | 8983299433 | 18257 | 21-Jul-2018 | 09:02 |  |  | 26-Jul-2018 | 17:30 | NOT WORKING | CHECK AND FOUND HAVE ISSUES WITH W.B.C. PARAMETERS. SOLVE THE PROBLEM. NOW MACHINE IS IN WORKING CONDITION. | Closed | 801604080E | MEDIUM | 128:28:0 | 6
1396 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00002 | 3 PART HAEMATOLOGY ANALYZER | GURUMULA | 9422144593 | 30204 | 20-Jul-2019 | 09:13 |  |  | 22-Jul-2019 | 11:46 | NOT WORKING | CHECK AND FOUND ALL READING ARE NOT SATISFACTORY. BACK FLUSH, DRAIN CHAMBER, HIGH VOLTAGE PULSE AND CLEAN CHAMBERS PROCESS HAS BEEN DONE FOR CLEANING. NOW READING OBTAINED ARE SATISFACTORY. 3-PART HAEMATOLOGY ANALYZER IS NOW WORKING. | Closed | 801604080E | MEDIUM | 50:33:0 | 3
1397 | BHANDARA | RH PALANDUR | MH/R1/BHA/RH/PLR/00018 | BP APPARATUS ANEROID | SUREKHA GAWAI | 7378714577 | 9542 | 22-Nov-2017 | 09:47 |  |  | 22-Nov-2017 | 16:09 | NOT WORKING | REPLACED THE MERCURY GLASS TUBE | Closed | 26814 | LOW | 6:22:0 | 1
1398 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00003 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | GEETA | 8605509772 | 34693 | 05-Nov-2019 | 09:10 |  |  | 09-Nov-2019 | 15:29 | NOT WORKING | CHECK AND FOUND HEATING COIL IS BURNT AND POWER CORD IS INTERNALLY SHORT. FIXED THE NEW HEATING COIL OF 2KW AND POWER CORD OF 240V~,16A.AUTOCLAVE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 102:19:0 | 5
1399 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00003 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | GEETA | 8605509722 | 37431 | 02-Jan-2020 | 08:45 |  |  | 04-Jan-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS EARTHING COMMING TO THE MACHINE.SOLVE THE PROBLEM NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 51:15:0 | 3
1400 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00003 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | Geeta Gajbe | 86055 09772 | 43371 | 23-Jun-2020 | 09:10 |  |  | 23-Jun-2020 | 09:30 | NOT WORKING | CHECK AND FOUND MACHINE HEATING COIL IS SHORT. NEW ONE REPLACE THE COIL AND FIX THE CONNECTION .NOW MACHINE IS WORKING AND CHEAK WITH INCHARGE SISTER BROTHER AND SISTER . | Closed | NO INFO | MEDIUM | 0:20:0 | 1
1401 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00003 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | GEETA GAJBE | 8605509772 | 35511 | 21-Nov-2019 | 08:30 |  |  | 21-Nov-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS STEAM IS LEAK FORM VALVE. SHORT OUT THIS PROBLEM NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 8:30:0 | 1
1402 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00003 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | GEETA | 8605509772 | 40395 | 03-Mar-2020 | 09:32 |  |  | 04-Mar-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS AIR VALVE IS LEAKED THEN MAKE PROPER CONNECTION.NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | NO INFO | MEDIUM | 26:28:0 | 2
1403 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00003 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | GEETA | 8605509722 | 42701 | 29-May-2020 | 08:50 |  |  | 30-May-2020 | 11:01 | NOT WORKING | CHECK AND FOUND HEATER COIL IS BURNT AND POWER CORD IS FAULTY. FIXED THE NEW HEATER COIL AND POWER CORD OF RATING 16A. AUTOCLAVE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 26:11:0 | 2
1404 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00004 | Fumigation Machine | GEETA GAJABE | 8605509772 | 38560 | 27-Jan-2020 | 09:05 |  |  | 28-Jan-2020 | 17:00 | NOT WORKING | CHECKED AND FOUND PROBLEM IN POWER BOARD IS NOT WORKING THEN MACHINE START FROM ANOTHER POWER BOARD. MACHINE IS IN WORKING CONDITION | Closed | NO INFO | LOW | 31:55:0 | 2
1405 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00007 | ANESTHESIA MACHINE (BOYLES APPARATUS) | GEETA GAJBE | 8605509772 | 20821 | 27-Sep-2018 | 08:22 |  |  | 03-Oct-2018 | 12:30 | NOT WORKING | PROBLEM REPORTED OF OXYGEN LEAKAGE. CHECK AND FOUND LEAKAGE IS FROM GAS CYLINDER. MACHINE IS WORKING NORMALLY WITHOUT ANY LEAKAGE. GAS CYLINDER IS REQUIRED TO BE CHANGED BY HOSPITAL ADMINISTRATION. ANESTHESIA MACHINE IS WORKING. | Closed | HT12G1958 | HIGH | 148:8:0 | 7
1406 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00008 | MONITOR | JYOTI | 9075693166 | 38339 | 23-Jan-2020 | 09:07 |  |  | 23-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN FUSE IS BURNED THEN CONNECT THE NEW 2A FUSE NOW MACHINE IS IN WORKING CONDITION HAND TO USER. | Closed | NO INFO | MEDIUM | 7:53:0 | 1
1407 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00008 | MONITOR | GITA GAJABE | 8605509772 | 16708 | 14-Jun-2018 | 09:35 |  |  | 18-Jun-2018 | 12:31 | NOT WORKING | REPLACED NIBP CUFF | Closed | NO INFO | MEDIUM | 98:56:0 | 5
1408 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00008 | MONITOR | GEETA GAJBE | 8605509772 | 20824 | 27-Sep-2018 | 08:22 |  |  | 02-Oct-2018 | 12:00 | NOT WORKING | CHECK AND FOUND BUTTON TYPE ON/OFF SWITCH FAULTY. CHANGED THE NEW ON/OFF SWITCH. MULTIPARAMONITOR IS NOW WORKING. | Closed | NO INFO | MEDIUM | 123:38:0 | 6
1409 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00009 | SHADOWLESS LAMP STAND TYPE | GEETA GAJABE | 8605509772 | 38561 | 27-Jan-2020 | 09:05 |  |  | 28-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN STABILIZER IS NOT WORKING THEN SHORT OUT THE PROBLEM OUT OF TWO BULB ONE BULB IS NOT WORKING DUE TO DAMAGE HOLDER. MR IS SEND CALL RAISED AGAIN WHEN SPARE IS RECEIVE . MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 31:55:0 | 2
1410 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00017 | CARDIOTOCOGRAPHY MACHINE | GAJABE NURSE | 8605509772 | 27173 | 29-Apr-2019 | 12:20 |  |  | 03-May-2019 | 15:32 | NOT WORKING | CHECK AND FOUND DISPLAY SCREEN IS BROKEN AND PROBE IS NOT WORKING. NEW DISPLAY SCREEN IS FIXED AND PROBE IS REPAIRED. CARDIOTOCOGRAPHY MACHINE IS NOW WORKING. | Closed | JA3C100226 | MEDIUM | 99:12:0 | 5
1411 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00017 | CARDIOTOCOGRAPHY MACHINE | GEETA | 8605509722 | 42700 | 29-May-2020 | 08:50 |  |  | 30-May-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN PROB, INTERNAL CONNECTION OF THE PROBE WIRE ARE TOUCH AND 3 WIRE ARE CUTED. THEN SHORT OUT THAT PROBLEM NOW MACHINE IS IN WORKING CONDITION | Closed | JA3C100226 | MEDIUM | 32:10:0 | 2
1412 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00028 | 3 PART HAEMATOLOGY ANALYZER | VINOD NAGDEVE | 8600183515 | 2885 | 01-Apr-2017 | 09:25 | 03-Apr-2017 | 12:30 | 01-May-2017 | 12:30 | NOT WORKING | REPLACED NEW PM KIT AND WBC,RBC APERTURE NOW MACHINE IS WORKING FINE | Closed | 201ESOH03479 | MEDIUM | 723:5:0 | 31
1413 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00029 | CENTRIFUGE (BENCH TOP) | VINOD | 8600183815 | 36277 | 07-Dec-2019 | 08:10 |  |  | 08-Dec-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN CARBON BRUSH IS DAMAGE THEN CONNECT THE NEW CARBON BRUSH NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 27:50:0 | 2
1414 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00029 | CENTRIFUGE (BENCH TOP) | VINOD | 8600183815 | 34916 | 09-Nov-2019 | 08:55 |  |  | 09-Nov-2019 | 15:12 | NOT WORKING | CHECK AND FOUND SPEED REGULATOR IS NOT WORKING AND CARBON BRUSH IS LOOSELY FIXED. FIXED THE NEW SPEED REGULATOR AND CARBON BRUSH IS TIGHTLY FIXED. ARMATURE OF THE MOTOR IS CLEANED. CENTRIFUGE MACHINE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 6:17:0 | 1
1415 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00035 | SHADOWLESS LAMP CEILING TYPE MINOR | GEETA | 8605509772 | 37886 | 13-Jan-2020 | 08:46 |  |  | 14-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN FUSE IS BURNED AND POWER PLUG IS SHORT CIRCUITED. THEN CHANGE THE NEW FUSE AND MAKE PROPER CONNECTION OF POWER PLUG. NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 32:14:0 | 2
1416 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00037 | PULSE OXIMETER LINE POWERED | GEETA GAZABE | 8605509772 | 24055 | 18-Jan-2019 | 08:16 |  |  | 19-Jan-2019 | 10:21 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS RAT BITTEN AND IT IS COMPLETELY DAMAGED AND DESTROYED. NEW SPO2 PROBE IS FIXED. PULSE OXIMETER IS NOW WORKING. | Closed | 175929 | MEDIUM | 26:5:0 | 2
1417 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00040 | BP APPARATUS ANEROID | GAJABE NURSE | 8605509772 | 27175 | 29-Apr-2019 | 13:00 |  |  | 03-May-2019 | 15:18 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK. GLASS TUBE IS ALSO BROKEN. FIXED THE NEW BLADDER AND GLASS TUBE. MERCURY IS ALSO REFILLED. B. P APPARATUS IS NOW WORKING. | Closed | 98135303 | LOW | 98:18:0 | 5
1418 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00040 | BP APPARATUS ANEROID | GITA GAJABE | 8605509772 | 16710 | 14-Jun-2018 | 09:35 |  |  | 15-Jun-2018 | 12:44 | NOT WORKING | REPLACED THE BLADDER | Closed | 98135303 | LOW | 27:9:0 | 2
1419 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00040 | BP APPARATUS ANEROID | NANDAGAVLI | 9420618966 | 36618 | 17-Dec-2019 | 08:09 |  |  | 17-Dec-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER, VALVE AND BULB THEN CHANGE THE NEW BLADDER, VALVE AND BULB NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | 98135303 | LOW | 8:51:0 | 1
1420 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00042 | BINOCULAR MICROSCOPE | GAJABE NURSE | 8605509772 | 27174 | 29-Apr-2019 | 12:30 |  |  | 03-May-2019 | 15:05 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. CHANGED THE OLD DEFECTIVE BULB WITH THE NEW BULB OF RATING 6V,20W.BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 98:35:0 | 5
1421 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00045 | Mobile X-ray machine (HF) | KUSHAL DHINGE | 9767136138 | 23176 | 19-Dec-2018 | 09:02 |  |  | 22-Dec-2018 | 15:03 | NOT WORKING | INSPECT AND FOUND ON SWITCH PROBLEM. RECTIFY THE SAME PROBLEM. ALSO DONE PROPER CALIBRATION OF MACHINE. NOW MACHINE WORKING PROPERLY. | Closed | 8G054 | MEDIUM | 78:1:0 | 4
1422 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00045 | Mobile X-ray machine (HF) | KUSHAL | 9767136138 | 38929 | 03-Feb-2020 | 08:42 |  |  | 07-Feb-2020 | 14:04 | NOT WORKING | RECTIFIED THE COLLIMATOR PROBLEM. NOW WORKING PROPERLY. | Closed | 8G054 | MEDIUM | 101:22:0 | 5
1423 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00045 | Mobile X-ray machine (HF) | ANKUSH | 9595431529 | 1243 | 28-Jan-2017 | 10:30 | 30-Jan-2017 | 10:45 | 01-Feb-2017 | 12:25 | NOT WORKING | REPAIRED THE TIMER CARD NOW EXPOSURE IS DONE,NOW MACHINE IS WORKING CONDITION | Closed | 8G054 | MEDIUM | 97:55:0 | 5
1424 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00045 | Mobile X-ray machine (HF) | KUSHAL DHENGE | 9767136138 | 30494 | 27-Jul-2019 | 09:00 |  |  | 29-Jul-2019 | 16:00 | NOT WORKING | CHECK AND FOUND DIGIVOLT CORD NOT WORKING. REPAIRED THE SAME. MACHINE TESTED AND HANDOVER TO USER IN PROPER WORKING ORDER. | Closed | 8G054 | MEDIUM | 55:0:0 | 3
1425 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00045 | Mobile X-ray machine (HF) | KUSHAL DHENGE | 9767136138 | 17748 | 11-Jul-2018 | 09:31 |  |  | 13-Jul-2018 | 14:30 | NOT WORKING | INSPECT AND FOUND THE EARTHING CURRENT BURLY TABLE DIGIVOLT TIMER CARD. RECTIFIED ALL THE PROBLEM. MACHINE TESTED AND HAND OVER TO TECHNICIAN IN PROPER WORKING ORDER | Closed | 8G054 | MEDIUM | 52:59:0 | 3
1426 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00050 | BP APPARATUS ANEROID | GEETA SHRIRAME | 8605509772 | 11524 | 10-Jan-2018 | 09:35 |  |  | 11-Jan-2018 | 14:45 | NOT WORKING | REPLACED NEW BLADDER AND EXTENSION TUBE, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 29:10:0 | 2
1427 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00050 | BP APPARATUS ANEROID | GITA GAJABE | 8605509772 | 16709 | 14-Jun-2018 | 09:35 |  |  | 15-Jun-2018 | 12:47 | NOT WORKING | REPLACED THE BLADDER AND METAL CONNECTOR | Closed | NO INFO |  | 27:12:0 | 2
1428 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00051 | BP APPARATUS ANEROID | GAJABE NURSE | 8605509772 | 27176 | 29-Apr-2019 | 13:10 |  |  | 03-May-2019 | 15:11 | NOT WORKING | CHECK AND FOUND BLADDER, BULB, VALVE AND CUFF IS MISSING. FIXED THE NEW CUFF, BULB, BLADDER AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 98:1:0 | 5
1429 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00051 | BP APPARATUS ANEROID | GEETA GAJBE | 8605509772 | 21397 | 16-Oct-2018 | 08:45 |  |  | 16-Oct-2018 | 16:25 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND VALVE. CUFF IS ALSO WORN OUT. CHANGED THE NEW BLADDER,CUFF AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 7:40:0 | 1
1430 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00051 | BP APPARATUS ANEROID | GEETA GAJABE | 8605509772 | 38448 | 24-Jan-2020 | 09:10 |  |  | 25-Jan-2020 | 17:00 | NOT WORKING | CHECKED AND FOUND GLASS TUBE IS DAMAGE THEN CHANGE THE NEW GLASS TUBE AND FEEL MERCURY . NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 31:50:0 | 2
1431 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00051 | BP APPARATUS ANEROID | GEETA GAJBE | 8605509772 | 35512 | 21-Nov-2019 | 08:30 |  |  | 25-Nov-2019 | 15:18 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN.FIXED THE NEW GLASS TUBE.REFILLED THE MERCURY. B.P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 102:48:0 | 5
1432 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00052 | BP APPARATUS ANEROID | GEETA GAJBE | 8605509772 | 21398 | 16-Oct-2018 | 08:45 |  |  | 16-Oct-2018 | 16:45 | NOT WORKING | CHECK AND FOUND BLADDER, CUFF, VALVE AND BULB IS MISSING. MERCURY IS ALSO NOT PRESENT. FIXED THE BLADDER, CUFF, VALVE AND BULB. MERCURY IS REFILLED.B.P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 8:0:0 | 1
1433 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00052 | BP APPARATUS ANEROID | G M NANDAGAHI | 9420618966 | 35301 | 18-Nov-2019 | 08:45 |  |  | 18-Nov-2019 | 16:21 | NOT WORKING | CHECK AND FOUND BLADDER, BULB AND VALVE IS HAVING LEAKAGE. FIXED THE NEW BLADDER, BULB AND VALVE. B P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 7:36:0 | 1
1434 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00052 | BP APPARATUS ANEROID | GEETA GAJABE | 8605509772 | 38449 | 24-Jan-2020 | 09:10 |  |  | 25-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS GLASS TUBE IS DAMAGE. NEED TO CHANGE GLASS TUBE. MR IS TO BE SEND . WHEN SPARE IS RECEIVE THEN RAISED CALL AGAIN. | Closed | NO INFO |  | 31:50:0 | 2
1435 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00052 | BP APPARATUS ANEROID | GEETA GAJBE | 8605509772 | 20822 | 27-Sep-2018 | 08:22 |  |  | 02-Oct-2018 | 14:15 | NOT WORKING | CHECK AND FOUND CUFF IS WORN OUT AND BLADDER, BULB AND VALVE IS LEAK. CHANGED THE NEW CUFF, BLADDER, BULB AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 125:53:0 | 6
1436 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00052 | BP APPARATUS ANEROID | GEETA GAJBE | 8605509772 | 35513 | 21-Nov-2019 | 08:30 |  |  | 25-Nov-2019 | 15:20 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN. FIXED THE NEW GLASS TUBE. REFILLED THE MERCURY. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 102:50:0 | 5
1437 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00053 | BP APPARATUS ANEROID | GEETA GAJBE | 8605509772 | 35514 | 21-Nov-2019 | 08:30 |  |  | 25-Nov-2019 | 17:23 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN. FIXED THE NEW GLASS TUBE. REFILLED THE MERCURY. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 104:53:0 | 5
1438 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00053 | BP APPARATUS ANEROID | GEETA GAJBE | 8605509772 | 20823 | 27-Sep-2018 | 08:22 |  |  | 02-Oct-2018 | 13:16 | NOT WORKING | CHECK AND FOUND CUFF, BLADDER IS WORN OUT. LEAKAGE IN VALVE IS BEING FOUND OUT. CHANGED THE NEW CUFF, BLADDER AND VALVE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 124:54:0 | 6
1439 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00054 | BP APPARATUS ANEROID | GAJABE NURSE | 8605509772 | 27177 | 29-Apr-2019 | 13:25 |  |  | 03-May-2019 | 14:57 | NOT WORKING | CHECK AND FOUND BLADDER IS WORN OUT. VALVE AND BULB IS ALSO HAVING LEAKAGE. CHANGED THE NEW BLADDER, CUFF, VALVE AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 97:32:0 | 5
1440 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00054 | BP APPARATUS ANEROID | GEETA GAJABE | 8605509772 | 38450 | 24-Jan-2020 | 09:10 |  |  | 25-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS GLASS TUBE IS DAMAGE . NEED TO CHANGE GLASS TUBE . MR IS TO BE SEND. CALL RAISED AGAIN WHEN SPARE IS RECEIVE. | Closed | NO INFO |  | 31:50:0 | 2
1441 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00054 | BP APPARATUS ANEROID | GEETA GAJBE | 8605509772 | 35515 | 21-Nov-2019 | 08:30 |  |  | 25-Nov-2019 | 17:26 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN.FIXED THE NEW GLASS TUBE. MERCURY REFILLING IS DONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 104:56:0 | 5
1442 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00054 | BP APPARATUS ANEROID | GEETA SHRIRAME | 8605509772 | 11525 | 10-Jan-2018 | 09:35 |  |  | 11-Jan-2018 | 11:44 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB WITH CONNECTOR NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 26:9:0 | 2
1443 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00055 | BP APPARATUS ANEROID | GEETA | 8605509772 | 34694 | 05-Nov-2019 | 09:15 |  |  | 09-Nov-2019 | 16:37 | NOT WORKING | CHECK AND FOUND LEAKAGE IN BLADDER AND VALVE. CHANGED THE DEFECTIVE BLADDER AND VALVE WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 103:22:0 | 5
1444 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00055 | BP APPARATUS ANEROID | GEETA GAJBE | 8605509772 | 35516 | 21-Nov-2019 | 08:30 |  |  | 25-Nov-2019 | 17:28 | NOT WORKING | CHECK AND FOUND GLASS TUBE IS BROKEN. BLADDER, VALVE AND BULB IS HAVING LEAKAGE. FIXED THE GLASS TUBE, BLADDER, VALVE AND BULB. REFILLING OF MERCURY IS DONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 104:58:0 | 5
1445 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00055 | BP APPARATUS ANEROID | GEETA | 8605509772 | 40394 | 03-Mar-2020 | 08:32 |  |  | 04-Mar-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN MERCURY BOX AIR IS NOT PASSED THEN SHORT OUT THAT PROBLEM NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | NO INFO |  | 27:28:0 | 2
1446 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00055 | BP APPARATUS ANEROID | GEETA SHRIRAME | 8605509772 | 11526 | 10-Jan-2018 | 09:35 |  |  | 11-Jan-2018 | 11:45 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB WITH CONNECTOR. NOW MACHINE IS WORKING CONDITION | Closed | NO INFO |  | 26:10:0 | 2
1447 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00067 | MONITOR | GEETA GAJBE | 8605509772 | 20825 | 27-Sep-2018 | 08:22 |  |  | 29-Oct-2018 | 16:35 | NOT WORKING | CHECK AND FOUND IT IS MULTIPARAMONITOR. SPO2 PROBE, ECG CABLE, POWER ADAPTER, NIBP CUFF ALONG WITH TUBING IS MISSING. IT IS PROPOSED FOR RW.NOW, CALL IS BEING CLOSED. AS SOON AS THE SPARES WOULD BE RECEIVED, AGAIN CALL WOULD BE RAISED AND MACHINE WOULD BE MADE FUNCTIONAL. | Closed | NO INFO |  | 776:13:0 | 33
1448 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00069 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | VINOD SIR | 8600183815 | 15591 | 07-May-2018 | 08:45 |  |  | 09-May-2018 | 11:01 | NOT WORKING | CLEANED THE CELL AND THE MACHINE IS IN WORKING CONDITION BUT NO REAGENTS | Closed | AT1450112AGD |  | 50:16:0 | 3
1449 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00070 | COLORIMETER | VINOD SIR | 8600183815 | 15592 | 07-May-2018 | 08:45 |  |  | 09-May-2018 | 11:31 | NOT WORKING | THE REAGENT NOT AVAILABLE. THE MACHINE IS IN WORKING CONDITION | Closed | NO INFO |  | 50:46:0 | 3
1450 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00071 | ELECTROPHORESIS MACHINE | VINOD SIR | 8600183815 | 15590 | 07-May-2018 | 08:45 |  |  | 09-May-2018 | 11:14 | NOT WORKING | THE MACHINE IS IN WORKING CONDITION BUT NO REAGENTS | Closed | LA690/164/VD0032 |  | 50:29:0 | 3
1451 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00079 | BLOOD BANK REFRIGERATOR 2-8°C | GEETA | 8605509772 | 34502 | 31-Oct-2019 | 09:36 |  |  | 01-Nov-2019 | 10:28 | NOT WORKING | CHECK AND FOUND COMPRESSOR IS INTERNALLY SHORT AND DUE TO WHICH THERE IS TRIP THROUGH ON/OFF SWITCH OF BBR MACHINE. CHANGED THE DEFECTIVE COMPRESSOR WITH THE NEW ONE. GAS CHARGING IS DONE. AT SET TEMPERATURE, COMPRESSOR TRIPS. PROPER COOLING IS ATTAINED. BBR MACHINE IS NOW WORKING. | Closed | NO INFO |  | 24:52:0 | 2
1452 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00083 | DENTAL CHAIR UNIT | RAJKUMAR MUNDLE | 9545188890 | 33996 | 17-Oct-2019 | 09:08 |  |  | 17-Oct-2019 | 17:00 | NOT WORKING | CHECKED THE MACHINE AND FOUND PROBLEM IN WATER LEAKAGE FROM TUBING . CHANGE THE TUBING AND VALVE NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO |  | 7:52:0 | 1
1453 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00083 | DENTAL CHAIR UNIT | RAJKUMAR | 9423112454 | 29683 | 09-Jul-2019 | 08:17 |  |  | 09-Jul-2019 | 15:06 | NOT WORKING | CHECK AND FOUND AIR ROTOR AND WATER LEAKAGE PROBLEM. ALSO FOUND FOOT PADDLE PROBLEM. FIXED THE AIR ROTOR AND FOOT PADDLE ISSUE. WATER LEAKAGE PROBLEM IS ALSO SOLVED. DENTAL CHAIR UNIT IS NOW WORKING. | Closed | NO INFO |  | 6:49:0 | 1
1454 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00084 | DENTAL SCALAR | DR RAJKUMAR A MUNDLE | 9423112454 | 32070 | 03-Sep-2019 | 09:09 |  |  | 06-Sep-2019 | 11:16 | NOT WORKING | CHECK AND FOUND WATER IS NOT COMING OUT FROM SCALER TIPS OF ULTRASONIC SCALER.PROBLEM FOUND IS OF SOLENOID VALVE. CHANGED THE DEFECTIVE VALVE WITH THE NEW ONE. NOW, WATER IS COMING OUT FROM SCALER TIPS. DENTAL SCALER MACHINE IS NOW WORKING. | Closed | K1050798B |  | 74:7:0 | 4
1455 | BHANDARA | RH PAUNI | MH/R1/BHA/RH/PRI/00088 | BP APPARATUS ANEROID | GEETA GAJABE | 8605509772 | 38451 | 24-Jan-2020 | 09:10 |  |  | 25-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS GLASS TUBE IS DAMAGE. THEN CHANGE THE NEW GLASS TUBE AND FEEL MERCURY. NOW MACHINE IS IN WORKING CONDITION. | Closed | 70263 |  | 31:50:0 | 2
1456 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00004 | CARDIOTOCOGRAPHY MACHINE | ASHOK RATHOD | 8255225508 | 13614 | 28-Feb-2018 | 11:15 |  |  | 02-Mar-2018 | 12:31 | NOT WORKING | SOLDERED WITH NEW MICRO SWITCH.NOW MACHINE IS WORKING CONDITION | Closed | JA3C100227 | MEDIUM | 49:16:0 | 3
1457 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00015 | AUTOCLAVE HP HORIZONTAL | ASHOK RATHOD | 9049161199 | 9445 | 21-Nov-2017 | 09:15 |  |  | 24-Nov-2017 | 14:55 | NOT WORKING | REPLACED NEW 17 INCHES GASKIT INTO AN MACHINE,PRESSURE LEAKAGE PROBLEM GET RECTIFIED. | Closed | 785 | MEDIUM | 77:40:0 | 4
1458 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00016 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | ASHOK RATHOD | 9049161199 | 9446 | 21-Nov-2017 | 09:15 |  |  | 24-Nov-2017 | 14:54 | NOT WORKING | REPLACED NEW 14 INCH RUBBER GASKET,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 77:39:0 | 4
1459 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00017 | AUTOCLAVE HP VERTICAL (1 BIN) DIGITAL | ASHOK RATHOD | 9049161199 | 9447 | 21-Nov-2017 | 09:15 |  |  | 24-Nov-2017 | 15:00 | NOT WORING | REPLACED NEW 14 INCH RUBBER GASKET,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 77:45:0 | 4
1460 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00026 | BP APPARATUS ANEROID | ASHOK RATHOD | 9049161199 | 9448 | 21-Nov-2017 | 09:15 |  |  | 24-Nov-2017 | 13:17 | NOT WORKING | REPLACED NEW BLADDER, CUFF . NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 76:2:0 | 4
1461 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00034 | 3 PART HAEMATOLOGY ANALYZER | HARISH RATHOD | 9423646911 | 5057 | 05-Jul-2017 | 11:00 | 06-Jul-2017 | 22:00 | 07-Jul-2017 | 12:30 | NOT WORKING | AFTER CHECKED A MACHINE A VALVES ARE NOT FUNCTIONING WELL,SO WE REPLACED 2 THREE WAY VALVE AND 2 TWO WAY VALVE A VALVE NUMBER ARE 1,6,7,9 AND ALSO RECONNECTING TUBING CONNECTION. NOW MACHINE IS WORKING GOOD CONDITION AND READING SHOWING PROPERLY. | Closed | 801509053E | MEDIUM | 49:30:0 | 3
1462 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00034 | 3 PART HAEMATOLOGY ANALYZER | LANJEWAR MADAM | 7709231577 | 5887 | 04-Aug-2017 | 10:51 |  |  | 09-Aug-2017 | 15:15 | NOT WORKING | REPLACED WBC CHAMBER,TWO WAY VALVE AND DONE CALIBRATION PROCESS,NOW MACHINE IS WORKING CONDITION. | Closed | 801509053E | MEDIUM | 124:24:0 | 6
1463 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00034 | 3 PART HAEMATOLOGY ANALYZER | KHADE SISTER | 8983299433 | 18256 | 21-Jul-2018 | 08:57 |  |  | 26-Jul-2018 | 10:05 | NOT WORKING | CHECK AND FOUND 3 PART HEMATOLOGY ANALYSER MACHINE PART RECEIVED AFTER 3 MONTH | Closed | 801509053E | MEDIUM | 121:8:0 | 6
1464 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00034 | 3 PART HAEMATOLOGY ANALYZER | DR JAGGAL | 8888863689 | 15616 | 07-May-2018 | 09:50 |  |  | 11-May-2018 | 13:20 | NOT WORKING | ACTUALLY THE MACHINE IS IN DH BHANDARA. VALVE 10,4 WERE NOT WORKING. CLEANED THE RESPECTIVE VALVE AND THE CHAMBERS TOO . NOW THE MACHINE IS IN WORKING CONDITION | Closed | 801509053E | MEDIUM | 99:30:0 | 5
1465 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00046 | BP APPARATUS ANEROID | ASHOK RATHOD | 9049161199 | 9449 | 21-Nov-2017 | 09:15 |  |  | 24-Nov-2017 | 12:20 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB.NOW MACHINE IS WORKING CONDITION. | Closed | 36340 | LOW | 75:5:0 | 4
1466 | BHANDARA | RH SIHORA | MH/R1/BHA/RH/SOR/00084 | INFRARED LAMP | NEHA | 7709781231 | 31158 | 12-Aug-2019 | 09:07 |  |  | 14-Aug-2019 | 10:38 | NOT WORKING | CHECKED THE INFRA RED LAMP FOUND PROBLEM LAMP IS FUSE CHANGE THE LAMP NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO |  | 49:31:0 | 3
1467 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00002 | NEBULIZER ELECTRIC | D R BHAKARE | 9421376086 | 24845 | 10-Feb-2019 | 09:02 |  |  | 12-Feb-2019 | 18:04 | NOT WORKING | CHECK AND FOUND MOTOR IS JAMMED. OILING OF MOTOR IS DONE. NOW, NEBULIZER MACHINE IS WORKING. | Closed | NO INFO | MEDIUM | 57:2:0 | 3
1468 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00002 | NEBULIZER ELECTRIC | KAVITA SAKHARE | 8668965342 | 37740 | 10-Jan-2020 | 08:05 |  |  | 10-Jan-2020 | 11:30 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN ON/OFF SWITCH IS FAULTY. THEN CONNECT THE NEW ON/OFF SWITCH NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 3:25:0 | 1
1469 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00002 | NEBULIZER ELECTRIC | SONKUSARE | 7218853925 | 15546 | 04-May-2018 | 08:22 |  |  | 05-May-2018 | 16:39 | NOT WORKING | PLACED PROPERLY WIRING NOW MACHINE IS WORKING OK | Closed | NO INFO | MEDIUM | 32:17:0 | 2
1470 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00005 | SYRINGE NEEDLE DESTROYER | KAVITA SAKHARE | 9175022209 | 3999 | 24-May-2017 | 10:30 |  |  | 26-May-2017 | 02:23 | NOT WORKING | CUTTED PART CLEANING AND GIVE THE GUIDANCE | Closed | NO INFO | LOW | 39:53:0 | 3
1471 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00008 | BP APPARATUS ANEROID | KAVITA | 8668965342 | 35957 | 29-Nov-2019 | 08:50 |  |  | 30-Nov-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BATTERY CELL CONNECTOR AND CUFF. CHANGE THE NEW CUFF NOW MACHINE IS IN WORKING CONDITION | Closed | NO INFO | LOW | 32:10:0 | 2
1472 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00008 | BP APPARATUS ANEROID | SATHAWANE | 8007557285 | 12840 | 13-Feb-2018 | 10:30 |  |  | 14-Feb-2018 | 21:54 | NOT WORKING | RECONNECT THE TUBE CONNECTION AND REARRANGE THE BATTERY | Closed | NO INFO | LOW | 35:24:0 | 2
1473 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00008 | BP APPARATUS ANEROID | KAVITA SAKHARE | 8668965342 | 37739 | 10-Jan-2020 | 08:05 |  |  | 10-Jan-2020 | 11:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN VALVE IS LEAK THEN CHANGE THE NEW VALVE NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | NO INFO | LOW | 2:55:0 | 1
1474 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00009 | BP APPARATUS ANEROID | DAVGADE NURSE | 9673642434 | 22734 | 08-Dec-2018 | 09:02 |  |  | 20-Dec-2018 | 04:20 | marqury not coming on up | CHECK AND FOUND CUFF, BLADDER AND VALVE IS MISSING. FIXED THE NEW CUFF, BLADDER & VALVE . B. P APPRATUS IS NOW WORKING | Closed | NO INFO | LOW | 283:18:0 | 13
1475 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00009 | BP APPARATUS ANEROID | SATGAWANE | 8007557285 | 11115 | 29-Dec-2017 | 09:40 |  |  | 31-Dec-2017 | 18:49 | NOT WORKING | REPLACED BLADDER,METAL VALVE, BULB | Closed | NO INFO | LOW | 57:9:0 | 3
1476 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00010 | O2 HUMIDIFIER | KAVITA SAKHARE | 9175022209 | 4223 | 02-Jun-2017 | 09:45 |  |  | 05-Jun-2017 | 11:50 | NOT WORKING | REPLACED 60W/230V LAMP | Closed | WH-12K10-2 | HIGH | 74:5:0 | 4
1477 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00012 | TRANSPORT INCUBATOR | KAVITA SAKHARE | 9175022209 | 4224 | 02-Jun-2017 | 09:45 |  |  | 06-Jun-2017 | 12:55 | NOT WORKING | REPLACED TEMPERATURE PROBE | Closed | TPB12K1002010/179 | HIGH | 99:10:0 | 5
1478 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00022 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | LANJE | 9422131350 | 30157 | 19-Jul-2019 | 09:31 |  |  | 22-Jul-2019 | 02:30 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN POWER SUPPLY AND TOUCH PAD NEED POST THIS MACHINE TO NAGPUR OFFICE FOR REPAIRING. AFTER REPAIRING MACHINE HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 64:59:0 | 4
1479 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00023 | BLOOD BANK REFRIGERATOR 2-8°C | N S LANJE | 9422131350 | 4962 | 03-Jul-2017 | 09:00 | 06-Jul-2017 | 12:35 | 05-Jul-2017 | 14:31 | NOT WORKING | REPLACE THE CABILLARY AND FILLED GAS. | Closed | 507255 | MEDIUM | 53:31:0 | 3
1480 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00023 | BLOOD BANK REFRIGERATOR 2-8°C | D R MHATRE | 9421376086 | 24729 | 06-Feb-2019 | 09:23 |  |  | 09-Feb-2019 | 16:04 | NOT WORKING | CHECK AND FOUND SUB ZERO SENSOR IS NOT PROPERLY FIXED AND DUE TO WHICH TEMPERATURE IS NOT CONTROLLED. FIXED THE SUB ZERO SENSOR AND TEMPERATUTE CONTROLLER IS CHECKED. BBR IS NOW WORKING. | Closed | 507255 | MEDIUM | 78:41:0 | 4
1481 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00023 | BLOOD BANK REFRIGERATOR 2-8°C | D R BHAKARE | 9421376086 | 29235 | 28-Jun-2019 | 09:15 |  |  | 01-Jul-2019 | 14:57 | NOT WORKING | CHECK AND FOUND TEMPERATURE CONTINUOSLY USE TO CHANGE FROM LOW TO HIGH AND VICE VERSA IN CERTAIN NUMBERS OF HOURS. PROBLEM IS DUE TO HUMIDITY CREATED DUE TO RAINING AND INSIDE BODY OF BBR CANNOT BALANCE THE TEMPERATUTE WITH INTERRUPTION OF HUMIDITY. DIRECTED TO OFF THE MACHINE TO 8-12 HOURS WITH ALL DOORS OPEN. DONE THE SAME AND NOW TEMPERATURE IS MAINTAINED NORMALLY. BBR IS NOW WORKING. | Closed | 507255 | MEDIUM | 77:42:0 | 4
1482 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00023 | BLOOD BANK REFRIGERATOR 2-8°C | BHAKARE | 9421376086 | 4099 | 30-May-2017 | 10:10 | 02-Jun-2017 | 10:37 | 03-Jun-2017 | 14:46 | NOT WORKING | CHANGED THE PT SENSOR NOW PROPER TEMPERATURE APPEAR | Closed | 507255 | MEDIUM | 100:36:0 | 5
1483 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00023 | BLOOD BANK REFRIGERATOR 2-8°C | PAKHARE | 9421376086 | 14127 | 23-Mar-2018 | 09:15 |  |  | 03-May-2018 | 15:15 | NOT WORKING | REPLACED COMPRESSOR AS SAME, THEN WE START INSTRUMENT. INSTRUMENT WORKING FINE WITHOUT ANY FALIOUR. | Closed | 507255 | MEDIUM | 990:0:0 | 42
1484 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00029 | 3 PART HAEMATOLOGY ANALYZER | T P DESHPANDE | 9881159854 | 35840 | 27-Nov-2019 | 09:00 |  |  | 27-Nov-2019 | 15:01 | NOT WORKING | CHECK AND FOUND REAGENT IS NOT AVAILABLE BUT MACHINE IS TOTALLY WORKING. SUGGESTED USER TO ARRANGE THE REAGENT. | Closed | 108ESOH02520 | MEDIUM | 6:1:0 | 1
1485 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00036 | DENTAL CHAIR UNIT | DR MISHRA | 8087074183 | 18252 | 21-Jul-2018 | 08:41 |  |  | 26-Jul-2018 | 17:30 | NOT WORKING | CHECK AND FOUND BORD AND TRANSFORMER IS MISSING. DID REQUEST FOR QUOTATION. TILL MACHINE IS IN DOWN CONDITION. PURPOSE FOR R.W. | Closed | NO INFO | LOW | 128:49:0 | 6
1486 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00038 | OPERATION TABLE HYDRAULIC (MAJOR) | REENA | 8669573944 | 32333 | 11-Sep-2019 | 08:38 |  |  | 12-Sep-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN ADJUSTING WALL. ADJUST THE WALL FROM SURFACE. NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 27:22:0 | 2
1487 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00039 | SHADOWLESS LAMP STAND TYPE | RINA DHUPARE | 8669573944 | 30429 | 25-Jul-2019 | 08:20 |  |  | 26-Jul-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN LAMP IS BURNED. CONNECT THE NEW LAMP NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 27:40:0 | 2
1488 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00040 | SHADOWLESS LAMP STAND TYPE | LINAT SONKUSARE | 7218853925 | 11426 | 09-Jan-2018 | 09:00 |  |  | 10-Jan-2018 | 15:14 | NOT WORKING | REPLACED THE 12V/50W, 24V /50W LAMP | Closed | NO INFO | MEDIUM | 30:14:0 | 2
1489 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00040 | SHADOWLESS LAMP STAND TYPE | LINATA SONKUSARE | 7218853925 | 20635 | 24-Sep-2018 | 08:12 |  |  | 27-Sep-2018 | 15:00 | NOT WORKING | CHECK AND FOUND ALL THE HALOGEN LIGHT OF RATING (12 V,50W) ATTACHED IN A DOME IS FUSED.CHANGED ALL THE NEW HALOGEN BULB(12V,50W) OF A DOME.DESIRED ILLUMINATION IS ACHIEVED.SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 78:48:0 | 4
1490 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00042 | ANESTHESIA MACHINE (BOYLES APPARATUS) | RINA DUPARE | 8669573944 | 38459 | 24-Jan-2020 | 09:14 |  |  | 25-Jan-2020 | 14:59 | NOT WORKING | CHECK AND FOUND NITROUS OXIDE PASSAGE IS BLOCKED AND DUE TO WHICH NITROUS OXIDE COULD NOT BE PASSED ALONG WITH OXYGEN TO OUTLET PASSAGE. ALSO LOW OXYGEN ALARM SYSTEM IS NOT WORKING. FIXED THE NITROUS OXIDE BLOCKAGE THROUGH VALVE AND NEW INTERNAL BATTERY OF 9V IS FIXED TO MAKE ALARM SYSTEM FUNCTIONAL. BOTH OXYGEN AND NITROUS OXIDE PASSAGE IS WORKING. ANESTHESIA MACHINE IS NOW WORKING. | Closed | HT12L1826 | HIGH | 29:45:0 | 2
1491 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00042 | ANESTHESIA MACHINE (BOYLES APPARATUS) | PRAVINA | 9021293547 | 29892 | 13-Jul-2019 | 08:15 |  |  | 16-Jul-2019 | 05:05 | NOT WORKING | CHECKED THE MACHINE AND CLEAN ANESTHESIA MACHINE NOW MACHINE IS IN GOOD WORKING CONDITION HAND OVER TO USER. | Closed | HT12L1826 | HIGH | 68:50:0 | 4
1492 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00044 | MONITOR | LINATA SONKUSARE | 7218853925 | 14073 | 22-Mar-2018 | 08:15 |  |  | 24-Mar-2018 | 12:27 | NOT WORKING | REPLACED NEW SPO2 PROBE INTO A MACHINE NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 52:12:0 | 3
1493 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00044 | MONITOR | SONKUSRE NURSE | 7218853925 | 11114 | 29-Dec-2017 | 09:40 |  |  | 31-Dec-2017 | 18:55 | NOT WORKING | REPLACED THE NIBP CUFF, ECG 3 LEAD CABLE, | Closed | NO INFO | MEDIUM | 57:15:0 | 3
1494 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00044 | MONITOR | RINA DHUPARE | 8669573944 | 31025 | 08-Aug-2019 | 09:05 |  |  | 10-Aug-2019 | 15:35 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS DEFECTIVE. CHANGED THE OLD DEFECTIVE PROBE WITH THE NEW ONE. MONITOR IS NOW WORKING. | Closed | NO INFO | MEDIUM | 54:30:0 | 3
1495 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00044 | MONITOR | LINATA SONKUR | 7218853925 | 20471 | 18-Sep-2018 | 09:30 |  |  | 19-Sep-2018 | 05:00 | NOT WORKING | CHECK AND FOUND MACHINE IS TOTALLY FUNCTIONAL WITH NIBP BUTTON WORKING NORMALLY B.P MESURED IS ALSO SATISFACTORY.MONITOR IS WORKING | Closed | NO INFO | MEDIUM | 19:30:0 | 2
1496 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00044 | MONITOR | REENA NURSE | 8669573944 | 33576 | 09-Oct-2019 | 09:05 |  |  | 09-Oct-2019 | 15:32 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS NOT WORKING.CHANGED THE DEFECTIVE SPO2 PROBE(6 PIN, DUAL LOCK) WITH THE NEW ONE. NOW MACHINE IS IN WORKING CONDITION AND HANDED OVER TO USER. | Closed | NO INFO | MEDIUM | 6:27:0 | 1
1497 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00044 | MONITOR | L R SONKUSARE | 7218853925 | 19279 | 18-Aug-2018 | 09:07 |  |  | 20-Aug-2018 | 14:12 | NOT WORKING | CHECK AND FOUND POWER CABLE IS DEFECTIVE .CHANGED THE NEW POWER CORD. MULTIPARAMONITOR IS NOW WORKING | Closed | NO INFO | MEDIUM | 53:5:0 | 3
1498 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00044 | MONITOR | PRAVINA | 9021293547 | 29890 | 13-Jul-2019 | 08:13 | 16-Jul-2019 | 11:53 | 20-Jul-2019 | 02:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN SPO2 PROB AND CUFF. CHANGE THE NEW PROB AND CUFF NOW MACHINE IS IN GOOD WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 161:47:0 | 8
1499 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00044 | MONITOR | L R SONKUSARE | 7251153925 | 21485 | 19-Oct-2018 | 08:33 |  |  | 23-Oct-2018 | 11:06 | NOT WORKING | CHECK AND FOUND MONITOR PULSE IS SHOWING VERY HIGH SET THE NORMAL MODE NOW MACHINE IS WORKING GOOD | Closed | NO INFO | MEDIUM | 98:33:0 | 5
1500 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00045 | BP APPARATUS ANEROID | DATTA BHAKARE | 9421376086 | 15996 | 22-May-2018 | 09:00 |  |  | 26-May-2018 | 01:29 | NOT WORKING | CHECK AND FOUND CHARGER IS FAULTY. CHARGER IS NOT UNDER SCOPE OF SUPPLY. | Closed | NO INFO | LOW | 88:29:0 | 5
1501 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00045 | BP APPARATUS ANEROID | LINATA SONKUSARE | 7218853925 | 14074 | 22-Mar-2018 | 08:15 |  |  | 24-Mar-2018 | 12:22 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB. NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | LOW | 52:7:0 | 3
1502 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00047 | SHADOWLESS LAMP STAND TYPE | L R SONKUSARE | 7251153925 | 21486 | 19-Oct-2018 | 08:33 |  |  | 30-Oct-2018 | 15:56 | NOT WORKING | CHECK AND FOUND HINGE OF DOME IS LOOSE. TIGHTENED THE HINGE. DOME IS FIXED. SHADOW LESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 271:23:0 | 12
1503 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00047 | SHADOWLESS LAMP STAND TYPE | REENA | 8669573944 | 40851 | 13-Mar-2020 | 08:30 |  |  | 13-Mar-2020 | 10:00 | NOT WORKING | CHECK AND FOUND SMPS IS SHORT. CHANGE THE OLD SMPS NEW ONE AND FIX WITH CONNECTION. LAMP IS WORKING AT PRESENT ME AND INCHARG SISTER. | Closed | NO INFO | MEDIUM | 1:30:0 | 1
1504 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00047 | SHADOWLESS LAMP STAND TYPE | LINATA SONKUR | 7218853925 | 20470 | 18-Sep-2018 | 09:30 |  |  | 19-Sep-2018 | 05:01 | NOT WORKING | CHECK AND FOUND HINGE OF DOME IS LOOSE AND CONNOT BE FIXED AT ONE PLACE.TIGHTENED THE HINGE OF THE DOME.SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 19:31:0 | 2
1505 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00048 | SUCTION PUMP PORTABLE, ELECTRONIC | LINATA SONKUR | 7218853925 | 20468 | 18-Sep-2018 | 09:30 |  |  | 19-Sep-2018 | 05:02 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE ISSUE .REQUIRED THE POWER SUPPLY PROBLEM.SUCTION MACHINE IS NOW WORKING | Closed | NO INFO | MEDIUM | 19:32:0 | 2
1506 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00048 | SUCTION PUMP PORTABLE, ELECTRONIC | SONKUSARE NURSE | 7218853925 | 11888 | 18-Jan-2018 | 10:12 |  |  | 19-Jan-2018 | 12:09 | NOT WORKING | REPLACED NEW SUCTION TUBES, NOW MACHINE IS WORKING CONDITION. . | Closed | NO INFO | MEDIUM | 25:57:0 | 2
1507 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00050 | SHADOWLESS LAMP STAND TYPE | PRAVINA | 9021293547 | 29893 | 13-Jul-2019 | 08:16 | 16-Jul-2019 | 12:55 | 18-Jul-2019 | 02:05 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM OF ALL LAMP IS BURNED CHANGE ALL THE LAMP NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 113:49:0 | 6
1508 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00050 | SHADOWLESS LAMP STAND TYPE | DHUPARE NURSE | 8669573944 | 38822 | 01-Feb-2020 | 09:00 |  |  | 03-Feb-2020 | 13:57 | NOT WORKING | CHECK AND FOUND 3 BULB IS FUSED OF RATING 12V,55W.FIXED THE NEW HALOGEN BULB OF 3 IN QUANTITY. SHADOWLESS LAMP STAND TYPE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 52:57:0 | 3
1509 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00050 | SHADOWLESS LAMP STAND TYPE | LR SONKUSARE | 7218853925 | 21731 | 30-Oct-2018 | 08:30 |  |  | 03-Nov-2018 | 11:14 | NOT WORKING | CHECK AND FOUND SHADOW LESS LAMP WAS DEFECTIVE .CHANGE THE SAME NOW MACHINEIS WORKING | Closed | NO INFO | MEDIUM | 98:44:0 | 5
1510 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00050 | SHADOWLESS LAMP STAND TYPE | RINA DHUPARE | 8669573944 | 30427 | 25-Jul-2019 | 08:16 | 26-Jul-2019 | 12:00 | 27-Jul-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN AIR FAN AND LAMP ARE BURNED DUE TO SHORT CIRCUITED NEED TO CHANGE AIR FAN AND LAMP. | Closed | NO INFO | MEDIUM | 51:44:0 | 3
1511 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00050 | SHADOWLESS LAMP STAND TYPE | L R SONKUTRE | 7218853925 | 23182 | 19-Dec-2018 | 09:15 |  |  | 01-Jan-2019 | 13:32 | NOT WORKING | CHECK AND FOUND BOTH THE HALOGEN BULB IS FUSED. CHANGED THE NEW HALOGEN BULB OF RATING 24V,150W.NOW,SHADOWLESS LAMP IS WORKING. | Closed | NO INFO | MEDIUM | 316:17:0 | 14
1512 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00050 | SHADOWLESS LAMP STAND TYPE | SONKUSARE NURSE | 7218853925 | 11889 | 18-Jan-2018 | 10:12 |  |  | 22-Jan-2018 | 11:08 | NOT WORKING | REPLACED NEW 12V 55WATTS OF HALOGEN LAMP INTO AN MACHINE AND REPAIRED 12V TRANSFORMER ,NOW MACHINE IS WORKING CONDITION. . | Closed | NO INFO | MEDIUM | 96:56:0 | 5
1513 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00050 | SHADOWLESS LAMP STAND TYPE | RINA DHUPARE | 8669573944 | 31024 | 08-Aug-2019 | 09:00 |  |  | 10-Aug-2019 | 03:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM OF ALL 6 BULB IS BURNT. CHANGE THE ALL BULB AND NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 42:0:0 | 3
1514 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00050 | SHADOWLESS LAMP STAND TYPE | DHUPARE | 8669573944 | 38536 | 27-Jan-2020 | 08:22 |  |  | 27-Jan-2020 | 15:31 | NOT WORKING | CHECK AND FOUND OUT OF 6 HALOGEN BULB IN A DOME, 3 IS NOT WORKING AS IT IS FUSED. FIXED THE NEW HALOGEN BULB OF 12V,55W OF 3 IN QUANTITY. ALL THE BULB IN A DOME IS WORKING. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 7:9:0 | 1
1515 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00051 | SHADOWLESS LAMP STAND TYPE | SONKUSARE NURSE | 7218853925 | 11890 | 18-Jan-2018 | 10:12 |  |  | 19-Jan-2018 | 12:19 | NOT WORKING | REPLACED NEW 24V 150WATTS HALOGEN LAMP. | Closed | NO INFO | MEDIUM | 26:7:0 | 2
1516 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00051 | SHADOWLESS LAMP STAND TYPE | DR MESHRAM | 9422760602 | 13247 | 26-Feb-2018 | 09:29 |  |  | 26-Feb-2018 | 19:14 | NOT WORKING | REPLACED THE 24V, 150W HALOGEN LAMP | Closed | NO INFO | MEDIUM | 9:45:0 | 1
1517 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00051 | SHADOWLESS LAMP STAND TYPE | D R BHAKARE | 9421376086 | 29234 | 28-Jun-2019 | 09:15 |  |  | 29-Jun-2019 | 14:01 | NOT WORKING | DUE TO ONGOING CONSTRUCTION WORK ALL THE WIRE CONNECTION IS OPENED AND SHADOWLESS LAMP IN UNINSTALLED. REINSTALLATION PROCESS IS NOT TO BE DONE BY US. ONCE THE VOLTAGE STABLIZER IS INSTALLED AND ALL THE WIRE CONNECTION IS FIXED THEN ONLY MACHINE WILL BE CHECKED. UNNECESSARY CALL IS BEING RAISED. | Closed | NO INFO | MEDIUM | 28:46:0 | 2
1518 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00051 | SHADOWLESS LAMP STAND TYPE | LR SONKUSARE | 7218853925 | 21732 | 30-Oct-2018 | 08:31 |  |  | 03-Nov-2018 | 11:54 | NOT WORKING | CHECK AND FOUND POWER SUPPLY CONNECTION WAS NOT FIXED. FIXED THE POWER SUPPLY CONNECTION PROPERLY NOW MACHINE IS WORKING GOOD | Closed | NO INFO | MEDIUM | 99:23:0 | 5
1519 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00051 | SHADOWLESS LAMP STAND TYPE | DHUPARE | 8669573944 | 38537 | 27-Jan-2020 | 08:32 |  |  | 27-Jan-2020 | 17:38 | NOT WORKING | CHECK AND FOUND OUT OF 2 DOME, LIGHT OF 1 DOME IS NOT WORKING. FIXED THE LOOSE CONNECTION OF INPUT WIRING THROUGH VOLTAGE STABLIZER. LIGHT OF BOTH THE DOME IS WORKING. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 9:6:0 | 1
1520 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00051 | SHADOWLESS LAMP STAND TYPE | RINA DHUPARE | 8669573944 | 30428 | 25-Jul-2019 | 08:17 | 26-Jul-2019 | 12:00 | 27-Jul-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN LAMP IS BURNED. CHANGE LAMP NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER | Closed | NO INFO | MEDIUM | 51:43:0 | 3
1521 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00051 | SHADOWLESS LAMP STAND TYPE | DHUPARE NURSE | 8669573944 | 38823 | 01-Feb-2020 | 09:00 |  |  | 03-Feb-2020 | 15:10 | NOT WORKING | CHECK AND FOUND VOLTAGE STABLISER TRIPS DUE TO SHORT OF INTERNAL WIRING OF SHADOWLESS LAMP. EXTERNAL REWIRING IS DONE. NEW HALOGEN BULB OF RATING 24V,150W OF 2 IN QUANTITY IS FIXED. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 54:10:0 | 3
1522 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00052 | PULSE OXIMETER LINE POWERED | SONKUSARE NURSE | 7218853925 | 8313 | 25-Oct-2017 | 10:40 |  |  | 27-Oct-2017 | 12:36 | NOT WORKING | REPLACED NEW NELCOR SPO2 PROBE, NOW MACHINE IS WORKING CONDITION | Closed | NO INFO | MEDIUM | 49:56:0 | 3
1523 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00052 | PULSE OXIMETER LINE POWERED | REENA | 8669573944 | 32332 | 11-Sep-2019 | 08:38 |  |  | 13-Sep-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN SPO2 PROB IS FAULTY. CHANGE THE NEW SPO2 PROB NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 51:22:0 | 3
1524 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00052 | PULSE OXIMETER LINE POWERED | REENA NURSE | 8669573944 | 33575 | 09-Oct-2019 | 09:05 |  |  | 09-Oct-2019 | 16:34 | NOT WORKING | CHECK AND FOUND SPO2 PROBE IS NOT WORKING. CHANGED THE OLD DEFECTIVE SPO2 PROBE WITH THE NEW ONE(9 PIN, D-TYPE). NOW MACHINE IS IN WORKING CONDITION AND HANDED OVER TO USER | Closed | NO INFO | MEDIUM | 7:29:0 | 1
1525 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00054 | BP APPARATUS ANEROID | SOKALI MESHRAM | 9422760602 | 13215 | 24-Feb-2018 | 08:42 |  |  | 24-Feb-2018 | 15:15 | NOT WORKING | REPLACED THE BLADDER, METAL VALVE, BULB, TUBE | Closed | NO INFO | LOW | 6:33:0 | 1
1526 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00055 | ANESTHESIA MACHINE (BOYLES APPARATUS) | PRAVINA | 9021293547 | 29891 | 13-Jul-2019 | 08:14 |  |  | 16-Jul-2019 | 05:00 | NOT WORKING | CHECKED THE MACHINE FOUND PROBLEM IN O2 VALVE IS LEAKAGE, REPAIR THAT LEAKAGE NOW MACHINE IS IN WORKING CONDITION HAND OVER TO THE USER. | Closed | NO INFO | HIGH | 68:46:0 | 4
1527 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00055 | ANESTHESIA MACHINE (BOYLES APPARATUS) | SONKUSARE NURSE | 7218853925 | 8314 | 25-Oct-2017 | 10:40 |  |  | 27-Oct-2017 | 15:38 | NOT WORKING | REPLACED NEW WASHER INTO AN O2 CYLINDER UNIT, MACHINE IS WORKING CONDITION. | Closed | NO INFO | HIGH | 52:58:0 | 3
1528 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00055 | ANESTHESIA MACHINE (BOYLES APPARATUS) | RINA DUPARE | 8669573944 | 38460 | 24-Jan-2020 | 09:14 |  |  | 25-Jan-2020 | 17:16 | NOT WORKING | CHECK AND FOUND INTERNAL TUBINGS IS HAVING LEAKAGE AND LOW OXYGEN ALARM SYSTEM IS NOW WORKING. FIXED THE NEW INTERNAL TUBINGS AND NEW INTERNAL BATTERY OF 9V FOR LOW OXYGEN ALARM SYSTEM. OXYGEN FLOW PASSAGE IS FIXED BUT NITROUS OXIDE IS PENDING AS CAST IRON PNEUMATIC PU ELBOW SPARE IS REQUIRED AS IT FOUND BROKEN. MACHINE WILL BE FULLY REPAIRED AS SOON AS PART WILL BE RECEIVED. | Closed | NO INFO | HIGH | 32:2:0 | 2
1529 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00056 | SUCTION PUMP PORTABLE, ELECTRONIC | LINATA SONKUR | 7218853925 | 20469 | 18-Sep-2018 | 09:30 |  |  | 19-Sep-2018 | 05:01 | NOT WORKING | CHECK AND FOUND POWER SUPPLY ISSUE FIXED THE POWER SUPPLY ISSUE.SUCTION MACHINE IS NOW WORKING | Closed | NO INFO | MEDIUM | 19:31:0 | 2
1530 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00057 | SUCTION PUMP PORTABLE, ELECTRONIC | D R BHAKARE | 9421376086 | 23329 | 24-Dec-2018 | 09:25 |  |  | 27-Dec-2018 | 13:15 | NOT WORKING | CHECK AND FOUND SUCTION MACHINE TUBE IS LEAK. FIXED THE TUBE. SUCTION PRESSURE IS GENERATED. NOW, MACHINE IS WORKING. | Closed | NO INFO | MEDIUM | 75:50:0 | 4
1531 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00061 | BP APPARATUS ANEROID | KAVITA | 8668965342 | 35953 | 29-Nov-2019 | 08:45 |  |  | 30-Nov-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER AND BATTERY CELL. CONNECT THE NEW BATTERY CELL AND BLADDER NOW MACHINE IS IN WORKING CONDITION. | Closed | 2631 | LOW | 32:15:0 | 2
1532 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00061 | BP APPARATUS ANEROID | KAVITA SAKHARE | 8668965342 | 37738 | 10-Jan-2020 | 08:05 |  |  | 10-Jan-2020 | 11:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN EXTENSION TUBE IS LEAK. THEN CHANGE THE NEW EXTENSION TUBE NOW MACHINE IS IN WORKING CONDITION. | Closed | 2631 | LOW | 2:55:0 | 1
1533 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00061 | BP APPARATUS ANEROID | SATHAWANE | 8007557285 | 12841 | 13-Feb-2018 | 10:30 |  |  | 15-Feb-2018 | 21:59 | NOT WORKING | REARRANGE THE BATTERY AND RECONNECT THE CONNECTOR | Closed | 2631 | LOW | 59:29:0 | 3
1534 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00062 | BP APPARATUS ANEROID | PRASHANT RAMTEKE | 9404312964 | 23596 | 04-Jan-2019 | 09:02 |  |  | 01-Feb-2019 | 10:45 | NOT WORKING | CHECK AND FOUND CONTROLLER BOARD IS DEFECTIVE. CONTROLLER BOARD COULD NOT BE REPAIRED AS ITS PURCHASE COST IS HIGH AND IS EQUIVALENT TO MACHINE COST. SO, MACHINE IS IN PROCESS FOR CONDEMNATION. | Closed | NO INFO | LOW | 673:43:0 | 29
1535 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00062 | BP APPARATUS ANEROID | PRASHANT RAMTEKE | 94043122964 | 17212 | 27-Jun-2018 | 08:20 |  |  | 02-Jul-2018 | 15:45 | NOT WORKING | CHECK THE MACHINE PROBLEM REPORTED B.P. BLADDER AND CLOTH AND BULB. REPLACED ALL THE PARTS WITH NEW ONE. NOW B.P. APRATUS WORKING FINE. | Closed | NO INFO | LOW | 127:25:0 | 6
1536 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00064 | BP APPARATUS ANEROID | DATTATARY BHAKRE | 9421376086 | 15879 | 18-May-2018 | 08:05 |  |  | 21-May-2018 | 02:30 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS FAULTY. REPLACED SAME AND PROVIDE NEW ONE. NOW B. P APRATUS IS WORKING FINE. | Closed | NO INFO | LOW | 66:25:0 | 4
1537 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | O R SINDHRAM | 8806600795 | 37135 | 26-Dec-2019 | 09:18 |  |  | 26-Dec-2019 | 10:26 | NOT WORKING | CHECKED AND FOUND INCOMING SUPPLY AND RELAY PROBLEM. RELAY CHANGE, RECTIFIED THE BOTH. NOW MACHINE WORKING PROPERLY. | Closed | SC-L/15853 | MEDIUM | 1:8:0 | 1
1538 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | OR SHINDARAM | 8806600795 | 36171 | 05-Dec-2019 | 08:16 |  |  | 06-Dec-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN EXPOSER IS NOT WORKING BECAUSE OF RELAY IS FAULTY. NEED TO CHANGE THE RELAY MACHINE IS NOT WORKING CONDITION. | Closed | SC-L/15853 | MEDIUM | 27:44:0 | 2
1539 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | O R SHINDRAM | 9403058530 | 23818 | 10-Jan-2019 | 05:00 |  |  | 10-Jan-2019 | 06:00 | NOT WORKING | check the machine found 20 awp mains fuse wdo blown up replaced with the same.also found 25 ma and 50wa station not working repairedthe veriable resister machine tasted and handover to user | Closed | SC-L/15853 | MEDIUM | 1:0:0 | 1
1540 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | O R SHINDRAM | 8806600795 | 25451 | 27-Feb-2019 | 09:11 |  |  | 01-Mar-2019 | 16:32 | NOT WORKING | CHECK AND FOUND INTERNAL FUSE 20 AMPERE WAS BLOWN UP. REPLACED WITH SAME. NOW MACHINE TESTED AND HANDED TO TECHNICIAN IN PROPER WORKING CONDITION. | Closed | SC-L/15853 | MEDIUM | 55:21:0 | 3
1541 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | O R SHINDRAM | 8806600795 | 24137 | 21-Jan-2019 | 09:16 |  |  | 27-Jan-2019 | 10:03 | NOT WORKING | CHECK AND FOUND EXTERNAL TUBE IS MACHINE CABLE IS FAULTY. NEED TOBE TO BE REPLACED FOR PROPER WIRING OF MACHINE REPLACED CABLE WITH NEW ONE & EXPOSURE SWITCH NOW MACHINE IS WORKING | Closed | SC-L/15853 | MEDIUM | 144:47:0 | 7
1542 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | O R SINDHGRAM | 8806600795 | 26374 | 29-Mar-2019 | 09:02 |  |  | 05-Apr-2019 | 04:05 | NOT WORKING | CHECK AND FOUND 24V DC REALY WAS NOT WORKING . REPLACE THE SAME MACHINE TESTED AND HANDED OVER TO THE MACHINE NOW MACHINE IS WORKING SATISFACTORY | Closed | SC-L/15853 | MEDIUM | 163:3:0 | 8
1543 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | O R SHINDRAM | 9403058530 | 11413 | 06-Jan-2018 | 09:43 |  |  | 08-Jan-2018 | 22:34 | NOT WORKING | RECONNECT THE ROTOR TUBE AND CONTROL PANEL CONNECTION | Closed | SC-L/15853 | MEDIUM | 60:51:0 | 3
1544 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | OR SINDRAM | 8806600795 | 38787 | 31-Jan-2020 | 08:30 |  |  | 31-Jan-2020 | 15:30 | NOT WORKING | CHECK AND FOUND X-RAY WAS NOT EXPOSED. RECONNECT ALL CONNECTORS PROPERLY. RECTIFIED THE LOOSE CONNECTION PROBLEM. NOW MACHINE IS WORKING PROPERLY. | Closed | SC-L/15853 | MEDIUM | 7:0:0 | 1
1545 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | OM PRAKASH | 8806600795 | 39702 | 18-Feb-2020 | 09:19 |  |  | 19-Feb-2020 | 16:30 | NOT WORKING | CHECKED AND FOUND FILAMENT WIRE (F1 & F2) ON TUBE HEAD NOT WORKING. RECTIFIED THE SAME. MACHINE TESTED AND HANDOVER TO USER IN PROPER WORKING ORDER. | Closed | SC-L/15853 | MEDIUM | 31:11:0 | 2
1546 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | OM PRAKASH | 8806600795 | 33882 | 15-Oct-2019 | 08:45 |  |  | 19-Oct-2019 | 14:30 | NOT WORKING | CHECKED AND FOUND INCOMING SUPPLY AND TIMER CARD PROBLEM. RECTIFIED THE BOTH. NOW MACHINE WORKING PROPERLY. | Closed | SC-L/15853 | MEDIUM | 101:45:0 | 5
1547 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | RAMCHANDRA | 8806600795 | 35076 | 13-Nov-2019 | 08:45 |  |  | 18-Nov-2019 | 15:27 | NOT WORKING | CHECK AND FOUND 24 VDC/500 OHM RELAY NOT WORKING.REPAIRED THE SAME. MACHINE TESTED AND HANDOVER TO USER IN PROPER WORKING CONDITION. | Closed | SC-L/15853 | MEDIUM | 126:42:0 | 6
1548 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00074 | Mobile X-ray machine (HF) | OR SINDRAM | 8806600795 | 37882 | 13-Jan-2020 | 08:38 |  |  | 13-Jan-2020 | 15:03 | NOT WORKING | CHECKED AND FOUND TIMER CARD & 24V RELAY NOT WORKING. REPAIRED TIMER CARD & REPLACED 24V RELAY. MACHINE TESTED AND HANDOVER TO USER IN PROPER WORKING ORDER. | Closed | SC-L/15853 | MEDIUM | 6:25:0 | 1
1549 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00079 | BP APPARATUS ANEROID | PRASHANT | 9404312964 | 36967 | 23-Dec-2019 | 08:16 |  |  | 23-Dec-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS CHARGER IS FAULTY. NEED TO NEW CHARGER . MACHINE IS WORKING IN BATTERY CELL. | Closed | NO INFO | LOW | 8:44:0 | 1
1550 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00079 | BP APPARATUS ANEROID | RAMTEME | 9404312964 | 33055 | 25-Sep-2019 | 09:06 |  |  | 27-Sep-2019 | 11:18 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS HAVING LEAKAGE AND DUE TO WHICH BP COULD NOT BE MEASURED.CHANGED THE OLD DEFECTIVE CUFF WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 50:12:0 | 3
1551 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00079 | BP APPARATUS ANEROID | PRASHANT RAMTEKE | 9404312964 | 23595 | 04-Jan-2019 | 09:02 |  |  | 08-Jan-2019 | 12:46 | NOT WORKING | CHECK AND FOUND MACHINE IS WORKING. | Closed | NO INFO | LOW | 99:44:0 | 5
1552 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00080 | BP APPARATUS ANEROID | PRASHANT | 9404312964 | 36968 | 23-Dec-2019 | 08:17 |  |  | 23-Dec-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS VALVE IS LEAK THEN CHANGE THE NEW VALVE NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | LOW | 8:43:0 | 1
1553 | BHANDARA | SDH SAKOLI | MH/R1/BHA/SDH/SKL/00092 | PULSE OXIMETER LINE POWERED | SONKUSARE NURSE | 7218853925 | 8315 | 25-Oct-2017 | 10:40 |  |  | 27-Oct-2017 | 14:48 | NOT WORKING | REARRANGED THE MAIN PCB SECTION,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 52:8:0 | 3
1554 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00001 | BP APPARATUS ANEROID | KUNDA AGASE | 9823793198 | 8143 | 16-Oct-2017 | 10:14 |  |  | 17-Oct-2017 | 20:31 | NOT WORKING | REPLACE BLADDER,CUFF,TUBES,METAL AND AIR TRANSFER CONNECTOR | Closed | RM0104035 | LOW | 34:17:0 | 2
1555 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00002 | SUCTION PUMP PORTABLE, ELECTRONIC | MANISHA SISTER | 9665426938 | 17176 | 26-Jun-2018 | 11:00 |  |  | 01-Jul-2018 | 13:30 | NOT WORKING | CHECK MACHINE PROBLEM REPORTED IN WIRING, CORRECT THE WIRING , NOW MACHINE IS WORKING SATISFACTORY. | Closed | NO INFO | MEDIUM | 122:30:0 | 6
1556 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00003 | STERILIZER (BIG, MEDIUM, SMALL INSTRUMENTS) | AATKARI NURSE | AATKARI NURSE | 13309 | 28-Feb-2018 | 09:00 |  |  | 02-Mar-2018 | 15:49 | NOT WORKING | REPLACED NEW 16A POWER WIRE,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 54:49:0 | 3
1557 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00004 | BINOCULAR MICROSCOPE | SHAILESH MANMODE | 7020300118 | 23205 | 20-Dec-2018 | 08:19 |  |  | 27-Dec-2018 | 15:48 | NOT WORKING | CHECK AND FOUND MACHINE SPARE PARTS IS MISSING AND MACHINE IS TOTALLY DAMAGED.PROCESS FOR CONDEMNATION. | Closed | NO INFO | MEDIUM | 175:29:0 | 8
1558 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00006 | BLOOD BANK REFRIGERATOR 2-8°C | POOJA GUPTA | 8055432343 | 40452 | 04-Mar-2020 | 09:13 |  |  | 04-Mar-2020 | 18:13 | NOT WORKING | CHECK AND FOUND MACHINE IS KEPT IS STORE FOR LAST 3 YEARS AND IS FULL OF DUST. TOOK THE MACHINE IN BLOOD STORAGE DEPARTMENT, CLEANED THE INTERNAL AND EXTERNAL BODY PARTS. FIXED THE POWER CABLE WIRE. TURNED THE MACHINE 'ON' AND FIXED THE TEMPERATURE AT 4*C. NOW, BBR MACHINE IS WORKING WITH CUT-OFF OF COMPRESSOR AT SET TEMPERATURE. PROPER COOLING IS BEING DONE NOW. | Closed | 507260 | MEDIUM | 9:0:0 | 1
1559 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00007 | WATER BATH | Pooja Gupta | 8055432343 | 44065 | 11-Jul-2020 | 08:50 |  |  | 11-Jul-2020 | 17:00 | NOT WORKING | CHECKED AND FOUND POWER CABLE IS DAMAGE . THEN MAKE PROPER CONNECTION OF POWER CABLE NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | LOW | 8:10:0 | 1
1560 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00008 | BINOCULAR MICROSCOPE | SHILPA JAMATKAR | 8698829729 | 14823 | 11-Apr-2018 | 09:00 |  |  | 12-Apr-2018 | 09:44 | NOT WORKING | MECHANICAL ARRANGEMENT GET COLLAPSED SO AN EQUIPMENT RECOMMENDED FOR CONDEMNATION PROCESS. | Closed | NO INFO | MEDIUM | 24:44:0 | 2
1561 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00008 | BINOCULAR MICROSCOPE | POOJA GUPTA | 8055432343 | 13616 | 09-Mar-2018 | 09:47 |  |  | 12-Mar-2018 | 15:13 | NOT WORKING | REPLACED 6V 20 WATTS HALOGEN LAMP, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 77:26:0 | 4
1562 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00008 | BINOCULAR MICROSCOPE | GUPTA NURSE | 8055432343 | 14068 | 22-Mar-2018 | 08:07 |  |  | 24-Mar-2018 | 13:15 | NOT WORKING | REPLACED NEW 6V 20 WATTS HALOGEN LAMP AND 100X OIL IMMERSION LENS, NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 53:8:0 | 3
1563 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00009 | BINOCULAR MICROSCOPE | MAHESH MANNADE | 9422834688 | 4389 | 09-Jun-2017 | 08:50 |  |  | 10-Jun-2017 | 03:00 | NOT WORKING | REPLACED NEW 6V 20 WATTS LAMP WITH THE NEW BULB HOLDER,NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 18:10:0 | 2
1564 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00009 | BINOCULAR MICROSCOPE | SHILPA JAMOTKAR | 8698829829 | 18821 | 04-Aug-2018 | 08:51 |  |  | 10-Aug-2018 | 03:00 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED. BULB HOLDER IS ALSO BURNT OUT. CHANGED THE NEW 6V,10W HALOGEN BULB AND THE BULB HOLDER. BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 138:9:0 | 7
1565 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00009 | BINOCULAR MICROSCOPE | MAHESH MANNADE | 8208858008 | 6555 | 29-Aug-2017 | 10:32 |  |  | 02-Sep-2017 | 15:14 | NOT WORKING | REPLACED THE 100 X LENSE | Closed | NO INFO | MEDIUM | 100:42:0 | 5
1566 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00009 | BINOCULAR MICROSCOPE | POOJA GUPTA | 8055432343 | 20866 | 28-Sep-2018 | 08:28 |  |  | 04-Oct-2018 | 12:05 | NOT WORKING | CHECK AND FOUND HOLOGEN DISPLAY OPTIC LAMP IS DEFECTIVE. CHANGE THE ABOVE METION LAMP WITH NEW ONE ,NOW MICROSCOPE IS WORKING GOOD | Closed | NO INFO | MEDIUM | 147:37:0 | 7
1567 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00009 | BINOCULAR MICROSCOPE | SHILPA JAMOTKAR | 8698829729 | 21482 | 19-Oct-2018 | 08:30 |  |  | 29-Oct-2018 | 16:07 | NOT WORKING | CHECK AND FOUND 6V,10W BULB IS FUSED. CHANGED THE NEW 6V,10W HALOGEN BULB. BINOCULAR MICROSCOPE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 247:37:0 | 11
1568 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00009 | BINOCULAR MICROSCOPE | SHILPA JAMATKAR | 8698829729 | 14824 | 11-Apr-2018 | 09:00 |  |  | 12-Apr-2018 | 09:47 | NOT WORKING | REPLACED NEW 5X EYEPIECE INTO A MACHINE NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 24:47:0 | 2
1569 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00009 | BINOCULAR MICROSCOPE | MR MAHESH | 9422834688 | 2178 | 06-Mar-2017 | 11:54 |  |  | 10-Mar-2017 | 02:30 | NOT WORKING | REPLACED 6VOLTS 20 WATTS HALOGEN LAMPS, NOW MACHINE IS WORKING FINE | Closed | NO INFO | MEDIUM | 86:36:0 | 5
1570 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | DR SACHIN | 99600206690 | 2996 | 10-Apr-2017 | 11:00 |  |  | 11-Apr-2017 | 12:30 | NOT WORKING | REPLACED A NEW PRIMARY MAINTENANCE KIT,NOW MACHINE IS WORKING GOOD CONDITION | Closed | 107ESOH02449 | MEDIUM | 25:30:0 | 2
1571 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | SHILPA JAMUTKAR | 8698829729 | 14679 | 07-Apr-2018 | 09:10 |  |  | 10-Apr-2018 | 12:40 | NOT WORKING | DONE CLEANING PROCESS ON THE MACHINE, NOW MACHINE IS WORKING CONDITION. | Closed | 107ESOH02449 | MEDIUM | 75:30:0 | 4
1572 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | SHILESH MANMODE | 7020300118 | 20124 | 06-Sep-2018 | 08:19 |  |  | 12-Sep-2018 | 11:59 | NOT WORKING | CHECK AND FOUND PLATLATE IS VERY HIGH AND HEMOGLOBINE IS VERY LOW READING GIVEN THE MACHINE.GIVE THE HYPOCHLORIDE TO THE MACHINE AS WELL AS DO THE SERVICE AND MAINTENECE. 3 PART HEMOTOLOGY MACHINE IS NOW WORKING | Closed | 107ESOH02449 | MEDIUM | 147:40:0 | 7
1573 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | POOJA GUPTA | 8055432343 | 40080 | 25-Feb-2020 | 09:41 |  |  | 27-Feb-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN GLASS TUBE IS DAMAGE AND VALVE IS FAULTY. THEN CHANGE THE NEW GLASS TUBE AND VALVE NOW MACHINE IS IN GOOD WORKING CONDITION HAND OVER TO USER. | Closed | 107ESOH02449 | MEDIUM | 50:19:0 | 3
1574 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | POOJA GUPTA | 8055432343 | 21302 | 12-Oct-2018 | 09:12 |  |  | 18-Oct-2018 | 11:05 | NOT WORKING | CHECK AND INSTRUMENT FOUND PROBLEM IN LARGE MAD NEEDS ASSEMBLY CARAGE AND NEEDLE ROAD CLEAN BY LUBRICANT AND OVERALL CLEANING DONE NOW MACHINE IS IN WORKING SATISFACTORY | Closed | 107ESOH02449 | MEDIUM | 145:53:0 | 7
1575 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | SAVALE | 9403417711 | 912 | 12-Jan-2017 | 10:01 |  |  | 16-Jan-2017 | 11:36 | NOT WORKING | REMOVING THE TUBE BLOCKAGE | Closed | 107ESOH02449 | MEDIUM | 97:35:0 | 5
1576 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | JAMOTKAR SISTER | 8698829729 | 15643 | 08-May-2018 | 08:52 |  |  | 15-May-2018 | 18:00 | NOT WORKING | DONE THE ROUTINE MAINTENANCE OF THE MACHINE. NOW THE MACHINE IS IN WORKING CONDITION | Closed | 107ESOH02449 | MEDIUM | 177:8:0 | 8
1577 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | REKHA KAVALE | 9403417711 | 3873 | 19-May-2017 | 09:50 |  |  | 20-May-2017 | 10:30 | NOT WORKING | AFTER AUTOMATIC CLEANING AND CONCENTRATED CLEANING PROCESS DONE PROPERLY,NOW MACHINE IS WORKING GOOD CONDITION AND ALSO READING SHOWING PROPERLY. | Closed | 107ESOH02449 | MEDIUM | 24:40:0 | 2
1578 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | SHILPA JAMOTKAR | 8698829729 | 28530 | 13-Jun-2019 | 09:07 |  |  | 17-Jun-2019 | 16:56 | NOT WORKING | CHECK THE INSTRUMENT AND FOUND REAGENTS CONTAINERS IS EXPIRED. CHANGED THE REAGENTS. OVERALL CLEANING IS DONE. PREVENTIVE MAINTENANCE IS DONE. SAMPLES REAGENT ACCEPTABLE. NOW MACHINE IS WORKING FINE. | Closed | 107ESOH02449 | MEDIUM | 103:49:0 | 5
1579 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | POOJA GUPTA | 8055432343 | 8469 | 29-Oct-2017 | 08:00 |  |  | 31-Oct-2017 | 13:50 | NOT WORKING | DONE CLEANING PROCESS ON PRINTER SECTION, NOW MACHINE IS WORKING CONDITION. | Closed | 107ESOH02449 | MEDIUM | 53:50:0 | 3
1580 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | PRATHIBA | 8657369089 | 32183 | 06-Sep-2019 | 09:05 |  |  | 09-Sep-2019 | 15:35 | NOT WORKING | CHECK AND FOUND CLOT IN ASPIRATION PROBE. REMOVED CLOT BY CLEANING SOLUTION THEN RAN SAMPLES. RESULTS ACCEPTABLE. | Closed | 107ESOH02449 | MEDIUM | 78:30:0 | 4
1581 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00010 | 3 PART HAEMATOLOGY ANALYZER | POOJA GUPTA | 8055432343 | 34988 | 11-Nov-2019 | 09:02 |  |  | 11-Nov-2019 | 15:40 | NOT WORKING | CHECK AND FOUND READING OBTAINED IS NOT ACCURATE. THERE IS LEAK FROM DILUTION SYRINGE AND NEEDLE ASSEMBLY.DILUTION SYRINGE AND NEEDLE SYRINGE MAINTENANCE DONE. NOW INSTRUMENT IS WORKING FINE. | Closed | 107ESOH02449 | MEDIUM | 6:38:0 | 1
1582 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | KAWALE | 9403417711 | 4980 | 03-Jul-2017 | 09:45 |  |  | 04-Jul-2017 | 10:33 | NOT WORKING | CLEAN THE FLOW CELL PARTS AND LIGHT PARTS WITH IN HELP OF TISSUE PAPER | Closed | NO INFO | MEDIUM | 24:48:0 | 2
1583 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | ARCHANA CHAVAN | 8421237314 | 25351 | 25-Feb-2019 | 09:06 |  |  | 01-Mar-2019 | 11:02 | NOT WORKING | CHECK AND FOUND MACHINE HALOGEN LAMP IS NOT WORKING NEED TO CHANGE THE LAMP. NOW CHANGE THE HALLOGEN LAMP NOW MACHINE IS WORKING GOOD | Closed | NO INFO | MEDIUM | 97:56:0 | 5
1584 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | POOJA GUPTA | 8055432343 | 40081 | 25-Feb-2020 | 09:41 |  |  | 26-Feb-2020 | 12:50 | NOT WORKING | CHECKED AND FOUND INSERTED OT, PT AND UREA PROGRAMS WRONG THEN DELETE WRONG PROGRAM AND INSERT CORRECT PROGRAM AND CALLIBRATE ABOVE ALL PARAMETER NOW RESULTS ARE EXEPTABLE | Closed | NO INFO | MEDIUM | 27:9:0 | 2
1585 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | POOJA GUPTA | 8055432343 | 32362 | 12-Sep-2019 | 08:08 |  |  | 13-Sep-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS CHECK LIGHT PATH ERROR IS SHOWING. CLEAN THE SENSOR BUT ERROR IS NOT GOING. MACHINE WAS NOT DONE SHOT AND SGP TEST OTHER WISE ALL TEST IS DONE . MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 27:52:0 | 2
1586 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | POOJA GUPTA | 8055432343 | 29837 | 12-Jul-2019 | 08:02 | 13-Jul-2019 | 09:30 | 13-Jul-2019 | 09:30 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM OF DISPLAY IS NOT WORKING NEED TO POST THIS MACHINE TO NAGPUR OFFICE FOR REPAIRING DISPLAY. AFTER REPAIRING THE DISPLAY MACHINE HANDE OVER TO USER IN WORKING CONDITION.805543 | Closed | NO INFO | MEDIUM | 25:28:0 | 2
1587 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | POOJA GUPTA | 8055432343 | 11431 | 09-Jan-2018 | 09:00 |  |  | 10-Jan-2018 | 14:01 | NOT WORKING | CREATED THE URIC & UREA STANDARD MODE IN MACHINE AND SAVED | Closed | NO INFO | MEDIUM | 29:1:0 | 2
1588 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | PRATIBHA VANJARI | 8975637243 | 38851 | 01-Feb-2020 | 09:12 |  |  | 02-Feb-2020 | 12:00 | NOT WORKING | CHECKED AND FOUND PROBLEM IS READING COMES INACCURATE THEN SAVE THE NEW PROGRAM NOW MACHINE IS WORKING OK. | Closed | NO INFO | MEDIUM | 26:48:0 | 2
1589 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | POOJA GUPTA | 8055432343 | 21303 | 12-Oct-2018 | 09:12 |  |  | 29-Oct-2018 | 12:04 | NOT WORKING | CHECK AND FOUND GLUCOSE READING MEASURED IS ABSURD. CLEANED THE FLOW CELL. RANGE HAS BEEN SET AND WASH IS BEING DONE. SEMI BIOCHEMISTRY IS NOW WORKING. | Closed | NO INFO | MEDIUM | 410:52:0 | 18
1590 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | SHAILESH MANMODE | 7020300118 | 30920 | 06-Aug-2019 | 08:46 |  |  | 09-Aug-2019 | 13:20 | READING ISSUE | CHECK AND FOUND DISPLAY CONTROLLER BOARD ISSUE. REPAIRED THE DISPLAY CONTROLLER BOARD. SEMI BIOCHEMISTRY ANALYZER IS NOW WORKING. | Closed | NO INFO | MEDIUM | 76:34:0 | 4
1591 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | POOJA GUPTA | 8055432343 | 33435 | 05-Oct-2019 | 09:03 |  |  | 07-Oct-2019 | 13:12 | NOT WORKING | CHECK AND FOUND HALOGEN BULB IS FUSED AND DUE TO WHICH CHECK LIGHT PATH ERROR IS SHOWN. FIXED THE NEW HALOGEN BULB OF RATING 6V,10W.SEMI AUTOMATED BIOCHEMISTRY IS NOW WORKING. | Closed | NO INFO | MEDIUM | 52:9:0 | 3
1592 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | ARCHANA CHAVHAN | 8698829729 | 26411 | 01-Apr-2019 | 08:12 |  |  | 05-Apr-2019 | 12:03 | NOT WORKING | CHECK AND FOUND PRIMARY BOARD AND SENSOR BOARD IS FAULTY . DEU TO NOT REAPAIRABLE PROPOSE TO CONDEMNATION | Closed | NO INFO | MEDIUM | 99:51:0 | 5
1593 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | POOJA GUPTA | 8055432343 | 24390 | 29-Jan-2019 | 08:14 |  |  | 04-Feb-2019 | 11:20 | NOT WORKING | CHECK AND FOUND MACHINE FLOW CELL IS NOT CLEAN. CLEAN THE FLOW CELL NOW MACHINE IS WORKING GOOD | Closed | NO INFO | MEDIUM | 147:6:0 | 7
1594 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | SHILPA NURSE | 8698829729 | 28786 | 19-Jun-2019 | 09:05 |  |  | 24-Jun-2019 | 12:13 | NOT WORKING | CHECK AND FOUND ERROR IN PROBE AND CHECK LIGHT PATH IS DISPLAYED ON SCREEN. REMOVED THE ERROR AND MACHINE IS NOW WORKING AND HANDED OVER TO USER. | Closed | NO INFO | MEDIUM | 123:8:0 | 6
1595 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | NILESH | 9970656661 | 7186 | 16-Sep-2017 | 09:32 |  |  | 17-Sep-2017 | 12:07 | NOT WORKING | REPLACE THE 6V/10W BULB | Closed | NO INFO | MEDIUM | 26:35:0 | 2
1596 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | SHILPA NURSE | 8698829729 | 17329 | 29-Jun-2018 | 09:15 |  |  | 03-Jul-2018 | 14:15 | NOT WORKING | CHECK AND FOUND, LIGHT IS NOT WORKING. REPLACED LIGHT, NOW MACHINE WORKING SATISFACTORY. | Closed | NO INFO | MEDIUM | 101:0:0 | 5
1597 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | DR SHINDE | 9130423071 | 35611 | 22-Nov-2019 | 08:35 |  |  | 23-Nov-2019 | 10:32 | NOT WORKING | CHECK ANF FOUND SMPS BOARD IS FAULTY. MR HAS BEEN SENT FOR THE SAME AND MATERIAL IS ON TRANSIT. AT PRESENT, THIS BREAKDOWN COMPLAINT IS CLOSED AND AS SOON AS SPARES WILL BE RECEIVED, PARTS WILL BE FIXED AND MACHINE WILL BE MADE FUNCTIONAL BY RAISING ANOTHER COMPLAINT. | Closed | NO INFO | MEDIUM | 25:57:0 | 2
1598 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | POOJA GUPTA | 8055432343 | 9713 | 27-Nov-2017 | 08:20 |  |  | 28-Nov-2017 | 11:35 | NOT WORKING | RECONNECT THE LAMP AND CLEANING PROCESS IS DONE | Closed | NO INFO | MEDIUM | 27:15:0 | 2
1599 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00011 | SEMI AUTOMATED BIOCHEMISTRY ANALYSER | SHILPA JAMOTKAR | 8698829729 | 28531 | 13-Jun-2019 | 09:07 |  |  | 18-Jun-2019 | 15:59 | NOT WORKING | CHECK THE INSTRUMENT AND FOUND CONTRAST PCB FAULTY. REPLACE CONTRAST PCB WITH THE NEW ONE. NOW DISPLAY COMING PROPERLY. SAMPLES READING ACCEPTABLE. | Closed | NO INFO | MEDIUM | 126:52:0 | 6
1600 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00014 | SYRINGE NEEDLE DESTROYER | Pratibha Wanjari | 8975637243 | 44064 | 11-Jul-2020 | 08:50 |  |  | 11-Jul-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS POWER CABLE DAMAGE DUE TO RAT THEN MAKE PROPER POWER CABLE CONNECTION NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | LOW | 8:10:0 | 1
1601 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00015 | CENTRIFUGE (BENCH TOP) | POOJA GUPTA | 8055432343 | 11339 | 05-Jan-2018 | 09:00 |  |  | 06-Jan-2018 | 19:51 | NOT WORKING | REPLACED THE MOTOR CARBAN BRUSH | Closed | NO INFO | MEDIUM | 34:51:0 | 2
1602 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00015 | CENTRIFUGE (BENCH TOP) | SHILPA JAMATKAR | 8698829729 | 14825 | 11-Apr-2018 | 09:00 |  |  | 12-Apr-2018 | 09:52 | NOT WORKING | MOTOR PARTS ARE MISSING AND GET DAMAGED SO WE RECOMMENDED FOR CONDEMNATION | Closed | NO INFO | MEDIUM | 24:52:0 | 2
1603 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00015 | CENTRIFUGE (BENCH TOP) | POOJA GUPTA | 8055432343 | 11696 | 15-Jan-2018 | 09:00 |  |  | 18-Jan-2018 | 11:13 | NOT WORKING | PROPERLY CONNECTED CARBON BRUSH | Closed | NO INFO | MEDIUM | 74:13:0 | 4
1604 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00022 | Anesthesia Work Station | Amita Budhe | 8275397196 | 43360 | 23-Jun-2020 | 08:30 |  |  | 23-Jun-2020 | 13:45 | NOT WORKING | CHECK AND FOUND LEAKAGE OF OXYGEN THROUGH TUBING OF FLOW METER. ENTIRE NEW TUBINGS IS FIXED THROUGH FLOW METER. LEAKAGE PROBLEM IS RESOLVED. ANESTHESIA WORK STATION IS NOW WORKING. | Closed | ASFJ025 | HIGH | 5:15:0 | 1
1605 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00023 | OPERATING MICROSCOPE | AMITA BHURE | 8275397196 | 23970 | 16-Jan-2019 | 08:09 |  |  | 21-Jan-2019 | 14:29 | NOT WORKING | CHECK AND FOUND POWER SUPPLY FAILURE. REPAIRED THE POWER SUPPLY ISSUE. OPERATING MICROSCOPE IS NOW WORKING. | Closed | NO INFO | MEDIUM | 126:20:0 | 6
1606 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00028 | SUCTION PUMP PORTABLE, ELECTRONIC | AMITA SHENDE | 8956378781 | 11281 | 04-Jan-2018 | 09:25 |  |  | 06-Jan-2018 | 19:48 | NOT WORKING | REPLACED THE MOTOR BEARING | Closed | NO INFO | MEDIUM | 58:23:0 | 3
1607 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00029 | MONITOR | AMITA SHENDE | 8956378781 | 11282 | 04-Jan-2018 | 09:25 |  |  | 06-Jan-2018 | 19:44 | NOT WORKING | RESOLDER THE CABLE | Closed | NO INFO | MEDIUM | 58:19:0 | 3
1608 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00029 | MONITOR | AMITA BUDHE | 8275397196 | 25612 | 06-Mar-2019 | 09:12 |  |  | 10-Mar-2019 | 16:00 | NOT WORKING | CHECK AND FOUND LEAKAGE IN NIBP CUFF. CHANGED THE DEFECTIVE WITH THE NEW NIBP CUFF. MONITOR IS NOW WORKING. | Closed | NO INFO | MEDIUM | 102:48:0 | 5
1609 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00030 | DIATHERMY MACHINE (ELECTRIC CAUTERY) | MR ANIL | 8806662525 | 2270 | 08-Mar-2017 | 09:35 |  |  | 11-Mar-2017 | 11:07 | NOT WORKING | AFTER REPLACED FUSE,NOW MACHINE IS WORKING CONDITION | Closed | ASG-M1303663 | MEDIUM | 73:32:0 | 4
1610 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00031 | SHADOWLESS LAMP STAND TYPE | AMITA SHENDE | 8956378781 | 11283 | 04-Jan-2018 | 09:25 |  |  | 06-Jan-2018 | 13:39 | NOT WORKING | REPLACED THE 12 V /50W BULB AND RECONNECT THE SCREWS | Closed | NO INFO | MEDIUM | 52:14:0 | 3
1611 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00032 | SHADOWLESS LAMP STAND TYPE | AMITA BUDHE | 8275397196 | 29102 | 26-Jun-2019 | 09:04 |  |  | 27-Jun-2019 | 17:54 | NOT WORKING | CHECK AND FOUND ONE OF THE DOME IS LOOSE AND FREELY MOVABLE. TIGHTENED THE DOME AND ITS NOW NOT MOVING FREELY.IT IS FIXED AS DESIRED IN ANY POSITION. SHADOWLESS LAMP IS NOW WORKING. | Closed | NO INFO | MEDIUM | 32:50:0 | 2
1612 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00033 | BP APPARATUS ANEROID | SHAILESH MANMODE | 7020300118 | 23566 | 03-Jan-2019 | 09:10 |  |  | 09-Jan-2019 | 14:01 | NOT WORKING | CHECK AND FOUND LEAK IN BLADDER, VALVE AND BULB. CUFF IS ALSO WORN OUT. CHANGED THE NEW VALVE, CUFF, BLADDER AND BULB. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 148:51:0 | 7
1613 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00033 | BP APPARATUS ANEROID | V K WALUNG | 9426618467 | 3941 | 22-May-2017 | 10:27 |  |  | 24-May-2017 | 11:18 | NOT WORKING | REPLACE BLADDER ,BULB WITH CONNECTOR | Closed | NO INFO | LOW | 48:51:0 | 3
1614 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00033 | BP APPARATUS ANEROID | AMITA | 8275397196 | 38694 | 29-Jan-2020 | 08:30 |  |  | 30-Jan-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM CONNECTER AND CUFF IS DAMAGE THEN CHANGE THE NEW CUFF AND CONNECTOR NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | LOW | 27:30:0 | 2
1615 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00034 | BP APPARATUS ANEROID | AMITA | 8275397196 | 43344 | 22-Jun-2020 | 08:55 |  |  | 22-Jun-2020 | 14:30 | NOT WORKING | CHECK AND FOUND LEAKAGE OF BULB AND VALVE. FIXED THE NEW BULB AND VALVE. BP APPARATUS IS NOW WORKING. | Closed | NO INFO | LOW | 5:35:0 | 1
1616 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00035 | ANESTHESIA MACHINE (BOYLES APPARATUS) | AMITA | 8275397196 | 43292 | 19-Jun-2020 | 09:15 |  |  | 20-Jun-2020 | 11:45 | NOT WORKING | CHECK AND FOUND INTERNAL TUBINGS HAVE LEAKAGE. FIXED THE NEW TUBINGS AND LEAKAGE OF OXYGEN PROBLEM IS RESOLVED. ANESTHESIA MACHINE IS NOW WORKING. | Closed | NO INFO | HIGH | 26:30:0 | 2
1617 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00036 | ANESTHESIA MACHINE (BOYLES APPARATUS) | AMITA BUDHE | 8275397196 | 34723 | 06-Nov-2019 | 08:22 |  |  | 08-Nov-2019 | 12:00 | NOT WORKING | CHECK THE EQUIPMENT AND FOUND PROBLEM IN BP CUFF EXTENSION TUBE IS BREAK. RECONNECT THE CONNECTOR AND EXTENSION TUBE . NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | HIGH | 51:38:0 | 3
1618 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00036 | ANESTHESIA MACHINE (BOYLES APPARATUS) | SHAILESH MANMODE | 7020300118 | 23567 | 03-Jan-2019 | 09:10 |  |  | 09-Jan-2019 | 13:27 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS LEAK. CHANGED THE NEW NIBP CUFF. ANESTHESIA MACHINE(WORK STATION) IS NOW WORKING. | Closed | NO INFO | HIGH | 148:17:0 | 7
1619 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00038 | RADIANT WARMER | DR MAHSKAR | 9860051432 | 9276 | 17-Nov-2017 | 09:46 |  |  | 22-Nov-2017 | 11:30 | NOT WORKNG | REPLACED THE BULBS | Closed | NO INFO | HIGH | 121:44:0 | 6
1620 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00038 | RADIANT WARMER | AMITA | 8075397196 | 38334 | 23-Jan-2020 | 08:59 |  |  | 23-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS LAMP ON/OFF SWITCHES IS FAULTY THEN CHANGE ALL ON/OFF SWITCHES AND MAKE INTERNAL LAMP CONNECTION PROPERLY NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | HIGH | 8:1:0 | 1
1621 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00038 | RADIANT WARMER | AMITA | 8275397196 | 27740 | 20-May-2019 | 08:10 |  |  | 24-May-2019 | 15:34 | NOT WORKING | CHECK AND FOUND ALL THE 6 BULB IS FUSE AND ALSO THE ON/OFF SWITCH IS FAULTY. NEW BULB(6) AND ON/OFF SWITCH(6) IS FIXED. RADIANT WARMER IS NOW WORKING. | Closed | NO INFO | HIGH | 103:24:0 | 5
1622 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00038 | RADIANT WARMER | AMITA | 8275397196 | 37967 | 15-Jan-2020 | 08:20 |  |  | 16-Jan-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN ON OFF SWITCHES IS FAULTY NEED TO CHANGE NEW SWITCH. MR IS SEND. CALL RAISED AGAIN WHEN SPARE WILL COME. | Closed | NO INFO | HIGH | 27:40:0 | 2
1623 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00044 | RADIANT WARMER | DR MASKAR | 9860051432 | 9225 | 16-Nov-2017 | 10:55 |  |  | 17-Nov-2017 | 11:40 | NOT WORKING | REPLACED THE BULB | Closed | NO INFO | HIGH | 24:45:0 | 2
1624 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00048 | BP APPARATUS ANEROID | DHANIKAR NANDESHWAR | 8605883470 | 32968 | 24-Sep-2019 | 09:06 |  |  | 24-Sep-2019 | 15:32 | NOT WORKING | CHECK AND FOUND PROBLEM IN CUFF AND VALVE. CHANGED THE NEW CUFF AND VALVE. NOW MACHINE IS IN GOOD WORKING CONDITION AND HANDED OVER TO USER. | Closed | RH10096 | LOW | 6:26:0 | 1
1625 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00048 | BP APPARATUS ANEROID | MANMODE | 7020300118 | 31795 | 26-Aug-2019 | 08:18 |  |  | 27-Aug-2019 | 14:35 | NOT WORKING | CHECKED THE EQUIPMENT AND CHANGED BULB AND EXTENSION TUBE. NOW, MACHINE IS IN WORKING CONDITION. HANDED OVER TO THE USER. | Closed | RH10096 | LOW | 30:17:0 | 2
1626 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00048 | BP APPARATUS ANEROID | KUNDA AGASE | 9823793198 | 8144 | 16-Oct-2017 | 10:16 |  |  | 17-Oct-2017 | 20:21 | NOT WORKING | REPLACE THE BLADDER,CUFF,TUBES,METAL AND AIR TRANSFER CONNECTORS | Closed | RH10096 | LOW | 34:5:0 | 2
1627 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00056 | AUTOCLAVE HP HORIZONTAL | AMITA BHUDHE | 8275397196 | 24272 | 24-Jan-2019 | 09:06 |  |  | 30-Jan-2019 | 14:39 | NOT WORKING | CHECK AND FOUND HINGE OF DOOR IS MISSING AND ALSO THE CONTRACTOR SWITCH IS MISSING.IT IS PROPOSED FOR RW.NOW THE CALL IS BEING CLOSED AND WHEN THE PARTS WILL BE ARRANGED AGAIN THE CALL WILL BE RAISED AND MACHINE WILL BE MADE OK. | Closed | MH-44/04 | MEDIUM | 149:33:0 | 7
1628 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00057 | AUTOCLAVE HP HORIZONTAL | DR MAHASKAR | 9860051432 | 17891 | 13-Jul-2018 | 10:45 |  |  | 17-Jul-2018 | 17:30 | NOT WORKING | CHECK THE MACHINE PROBLEM REPORTED IN DOOR HANDLE AND WIRING OF MACHINE. DONE HANDLE SERVICING AND PROPER WIRING DONE. NOW MACHINE IS IN WORKING CONDITION. | Closed | MH-71/04 | MEDIUM | 102:45:0 | 5
1629 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00057 | AUTOCLAVE HP HORIZONTAL | AMITA NURSE | 8275397196 | 36732 | 18-Dec-2019 | 08:36 |  |  | 19-Dec-2019 | 15:30 | NOT WORKING | CHECK AND FOUND LEAKAGE FROM BOILER CHAMBER. ENTIRE SHEETS OF BOILER IS DEGRADED. NEW SHEETS AND ITS SCREWS TO BE FIXED. MR IS TO BE SEND. NOW THIS BREAKDOWN CALL IS CLOSED. AS SOON AS SPARES WILL BE RECEIVED AGAIN THE COMPLIANT WILL BE RAISED AND MACHINE WILL BE MADE FUNCTIONAL. | Closed | MH-71/04 | MEDIUM | 30:54:0 | 2
1630 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00057 | AUTOCLAVE HP HORIZONTAL | AMITA | 8275397196 | 32522 | 16-Sep-2019 | 08:18 |  |  | 18-Sep-2019 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS HEATING COIL IS BURN. CHANGE THE COIL NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | MH-71/04 | MEDIUM | 51:42:0 | 3
1631 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00057 | AUTOCLAVE HP HORIZONTAL | S Y NANMUDE | 9657208602 | 12544 | 05-Feb-2018 | 08:20 |  |  | 07-Feb-2018 | 12:26 | NOT WORKING | THIS MACHINE ACTUAL ASSET NAME IS AUTOCLAVE HORIZONTAL NOT AN RADIANT WARMER CHECKED AND REPLACED NEW 22.5 GASKIT INTO AN MACHINE NOW MACHINE IS WORKING CONDITION. | Closed | MH-71/04 | MEDIUM | 52:6:0 | 3
1632 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00060 | SUCTION PUMP PORTABLE, ELECTRONIC | SHALINI SAMUNDARE | 9922869996 | 17878 | 13-Jul-2018 | 09:17 |  |  | 16-Jul-2018 | 13:30 | NOT WORKING | CHECK THE MACHINE THERE WAS WIRING AND CONNECTION PROBLEM IN MACHINE, PLACED THE WIRING PROPERLY NOW MACHINE WORKING FINE. HANDOVER TO THE USER IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 76:13:0 | 4
1633 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00060 | SUCTION PUMP PORTABLE, ELECTRONIC | DR MHASKAR | 9860051432 | 9646 | 24-Nov-2017 | 09:40 |  |  | 29-Nov-2017 | 12:44 | NOT WORKING | BY THIS ASSET AN ACTUAL MACHINE IS SUCTION MACHINE THEN WE DONE MOTOR REWINDING NOW MACHINE HANDOVER TO USER AS WORKING CONDITION. | Closed | NO INFO | MEDIUM | 123:4:0 | 6
1634 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00062 | NEBULIZER ELECTRIC | Shalini Samudre | 9922869996 | 41579 | 10-Apr-2020 | 08:20 |  |  | 10-Apr-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS MOTOR IS NOT WORKING. THEN SHORT OUT THE PROBLEM AND FAN CONNECT PROPERLY. NOW MACHINE IS IN GOOD CONDITION. | Closed | NO INFO | MEDIUM | 8:40:0 | 1
1635 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00064 | BP APPARATUS ANEROID | Amol D Sapate | 9284465389 | 44247 | 15-Jul-2020 | 09:14 |  |  | 16-Jul-2020 | 09:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS MERCURY IS MISSING AND VALVE IS LEAKED THEN FILL THE MERCURY AND CHANGE NEW VALVE NOW MACHINE IS IN GOOD WORKING CONDITION. | Closed | RH004033 | LOW | 23:46:0 | 2
1636 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00066 | PULSE OXIMETER LINE POWERED | Shalini Samudre | 9922869996 | 41578 | 10-Apr-2020 | 08:20 |  |  | 10-Apr-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS SPO2 PROB IS NOT WORKING THEN SHORT OUT THE PROBLEM AND MAKE PROPER CONNECTION OF SPO2 PROB NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO | MEDIUM | 8:40:0 | 1
1637 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00069 | DENTAL CHAIR UNIT | DR LONARE | 9021156047 | 6033 | 11-Aug-2017 | 10:15 |  |  | 16-Aug-2017 | 10:13 | NOT WORKING | REQUEST THE TECHNICIAN,TO REINSTALLATION THE MACHINE | Closed | NO INFO | LOW | 119:58:0 | 6
1638 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00070 | DENTAL CHAIR UNIT | DR LONARE | 9021156047 | 2605 | 18-Mar-2017 | 10:05 |  |  | 23-Mar-2017 | 11:55 | NOT WORKING | REPLACED THE NOZZLE, AIR REGULATOR AND ON OFF SWITCH | Closed | NO INFO | LOW | 121:50:0 | 6
1639 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00070 | DENTAL CHAIR UNIT | DR LONARE | 9021156047 | 6034 | 11-Aug-2017 | 10:18 |  |  | 16-Aug-2017 | 11:56 | NOT WORKING | REMOVE THE MOISTURE FROM COMPRESSOR PLACE | Closed | NO INFO | LOW | 121:38:0 | 6
1640 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00070 | DENTAL CHAIR UNIT | DR JYOTSHNA SHINDE | 9130423071 | 36387 | 10-Dec-2019 | 09:18 |  |  | 10-Dec-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN COMPRESSOR IS NOT WORKING. SHORT OUT THE PROBLEM NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | LOW | 7:42:0 | 1
1641 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00070 | DENTAL CHAIR UNIT | SAVALE | 9403417711 | 913 | 12-Jan-2017 | 11:00 |  |  | 17-Jan-2017 | 11:30 | NOT WORKING | POWER BOARD GET DAMAGED FULLY | Closed | NO INFO | LOW | 120:30:0 | 6
1642 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00070 | DENTAL CHAIR UNIT | DR JYOTISHNA | 9130423071 | 36018 | 02-Dec-2019 | 08:08 |  |  | 02-Dec-2019 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS ALL TUBING CONNECTION IS DAMAGED BY RAT AND CHANGE JACKPOT WALL AND CHANGE ALL TUBING CONNECTION. NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | LOW | 8:52:0 | 1
1643 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00072 | SUCTION PUMP PORTABLE, ELECTRONIC | HEENA NURSE | 9561501340 | 25837 | 14-Mar-2019 | 08:00 |  |  | 14-Mar-2019 | 16:04 | NOT WORKING | SUCTION MACHINE IS NOT PROPERLY WORK. BECAUSE SUCTION MACHINE MOTOR CAPACITOR IS FALUTY (DAMGE). SO NEW CAPACITOR IS APPLY AND MACHINE IS WORKING. AND MACHINE IS WELL AND PROPERLY WORK AT PRESENT. | Closed | NO INFO | MEDIUM | 8:4:0 | 1
1644 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00072 | SUCTION PUMP PORTABLE, ELECTRONIC | REENA LONARE | 9561501340 | 37423 | 02-Jan-2020 | 08:32 |  |  | 02-Jan-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN SUCTION MOTOR. THEN FEEL OIL IN THE MOTOR NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | NO INFO | MEDIUM | 8:28:0 | 1
1645 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00076 | CARDIOTOCOGRAPHY MACHINE | SHAILESH MANMODE | 7447685806 | 5515 | 20-Jul-2017 | 09:30 |  |  | 26-Jul-2017 | 13:40 | NOT WORKING | RECOMMEND FOR CONDEMNATION | Closed | JA3C100256 | MEDIUM | 148:10:0 | 7
1646 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00077 | PULSE OXIMETER LINE POWERED | BAVANE NURSE | 9405766510 | 7263 | 18-Sep-2017 | 10:00 |  |  | 19-Sep-2017 | 12:04 | NOT WORKING | REPLACE THE 12.5V ADAPTER | Closed | NO INFO | MEDIUM | 26:4:0 | 2
1647 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00083 | BP APPARATUS ANEROID | MANMODE | 7020300118 | 31635 | 22-Aug-2019 | 09:00 |  |  | 23-Aug-2019 | 11:05 | NOT WORKING | CHECKED THE EQUIPMENT AND CHANGED BLADDER AND CUFF. NOW MACHINE IS IN WORKING CONDITION AND HANDED TO USER. | Closed | RM0104027 | LOW | 26:5:0 | 2
1648 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00084 | 300 MA XRAY | Saroj Kadav | 9405245129 | 44020 | 10-Jul-2020 | 08:30 |  |  | 10-Jul-2020 | 10:00 | NOT WORKING | CHECK AND FOUND COLLIMETER BULB IS NOT WORKING. REPAIRED THE WIRING AND TOGGLE SWITCH. NOW, COLLIMETER BULB IS WORKING. DONE THE BASIC MAINTENANCE AND CALIBRATION. X-RAY MACHINE IS NOW WORKING. | Closed | 4k078 | MEDIUM | 1:30:0 | 1
1649 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00084 | 300 MA XRAY | MISS SEEMA | 9405245129 | 16438 | 05-Jun-2018 | 09:00 |  |  | 09-Jun-2018 | 14:30 | NOT WORKING | X RAY ASSEMBLY AND POLIMER TUBE IS NOT WORKING. | Closed | 4k078 | MEDIUM | 101:30:0 | 5
1650 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00091 | BP APPARATUS ANEROID | Meena Lonare | 9561501340 | 44067 | 11-Jul-2020 | 08:54 |  |  | 11-Jul-2020 | 17:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IS BULB AND VALVE IS LEAK THE CONNECT THE NEW VALVE AND BULB NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 8:6:0 | 1
1651 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00091 | BP APPARATUS ANEROID | DR NEENA LONARE | 9021156047 | 26497 | 03-Apr-2019 | 09:05 |  |  | 03-Apr-2019 | 14:01 | NOT WORKING | BP APPRATUS GLASS TUBE IS BROKEN. SO I WILL NEW GLASS TUBE IS APPLY. AND APPRATUS IS WELL AND PROPERLY WORK AT PRESENT. | Closed | NO INFO |  | 4:56:0 | 1
1652 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00091 | BP APPARATUS ANEROID | MANMODE | 7020300118 | 31636 | 22-Aug-2019 | 09:00 |  |  | 23-Aug-2019 | 11:13 | NOT WORKING | CHECK AND FOUND BLADDER IS LEAK AND CUFF IS WORN OUT. CHANGED THE OLD DEFECTIVE PART WITH TH NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 26:13:0 | 2
1653 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00092 | BP APPARATUS ANEROID | MANMODE | 7020300118 | 31637 | 22-Aug-2019 | 09:00 |  |  | 23-Aug-2019 | 11:18 | NOT WORKING | CHECK AND FOUND VALVE IS LEAK AND CUFF IS WORN OUT. CHANGED THE OLD DEFECTIVE PART WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 26:18:0 | 2
1654 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00094 | BP APPARATUS ANEROID | Meena Lonare | 9561501340 | 44068 | 11-Jul-2020 | 08:54 |  |  | 11-Jul-2020 | 17:00 | NOT WORKING | CHECKED AND FOUND BLADDER AND BULB IS LEAK THEN CONNECT THE NEW BLADDER AND BULB NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 8:6:0 | 1
1655 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00095 | CENTRIFUGE (BENCH TOP) | POOJA GUPTA | 8055432343 | 21304 | 12-Oct-2018 | 09:12 |  |  | 29-Oct-2018 | 10:00 | NO WORKING | CHECK AND FOUND SPEED REGULATOR IS NOT WORKING. CHANGED THE NEW SPEED REGULATOR. CENTRIFUGE MACHINE IS NOW WORKING. | Closed | NO INFO |  | 408:48:0 | 18
1656 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00095 | CENTRIFUGE (BENCH TOP) | SHILPA JAMALKAR | 8698829729 | 23179 | 19-Dec-2018 | 09:10 |  |  | 27-Dec-2018 | 14:46 | NOT WORKING | CHECK AND FOUND CARBON BRUSH IS DEGRADED. FIXED THE NEW CARBON BRUSH. CENTRIFUGE MACHINE IS NOW WORKING. | Closed | NO INFO |  | 197:36:0 | 9
1657 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00095 | CENTRIFUGE (BENCH TOP) | SHILPA JAOTKAR | 8698829729 | 18036 | 17-Jul-2018 | 09:56 |  |  | 21-Jul-2018 | 17:30 | NOT WORKING | CHECK THE MACHINE PROBLEM REPORTED IN WIRING AND NEED TO REPLACE CARBON BRUSH. DID REPLACEMENT OF CARBON BRUSH AND DID PROPER WIRING. NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 103:34:0 | 5
1658 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00095 | CENTRIFUGE (BENCH TOP) | SHILPA JAMATKAR | 8698829729 | 14527 | 04-Apr-2018 | 09:08 |  |  | 05-Apr-2018 | 13:00 | NOT WORKING | REPLACED NEW CARBON BRUSH INTO A MACHINE NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 27:52:0 | 2
1659 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00095 | CENTRIFUGE (BENCH TOP) | SHILPA JAMOTKAR | 8698829729 | 19947 | 03-Sep-2018 | 09:08 |  |  | 07-Sep-2018 | 11:59 | NOT WORKING | CHECK AND FOUND SPEED CONTROL REGULATOR WAS NOT FIXED TO POWER SUPPLY.FIXED THE POWER SUPPLY TO THE SPEED CONTROL REGULATOR MACHINE IS WORKING SMOOTHALY HANDOVER TO THE USER | Closed | NO INFO |  | 98:51:0 | 5
1660 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00095 | CENTRIFUGE (BENCH TOP) | SHILPA NURSE | 8698829729 | 28787 | 19-Jun-2019 | 09:05 |  |  | 24-Jun-2019 | 12:20 | NOT WORKING | CHECK AND FOUND MOTOR IS UNBALANCED AND LOOSE & CUVETTE FOR TEST TUBE IS BROKEN. FIXED THE MOTOR. AS CUVETTE IS CONSUMABLE ITEM, IT MUST BE ARRANGED BY HOSPITAL AUTHORITY AND WON'T BE PROVIDED BY COMPANY. CENTRIFUGE MACHINE IS NOW WORKING AND HANDED OVER TO THE USER. | Closed | NO INFO |  | 123:15:0 | 6
1661 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00099 | BP APPARATUS ANEROID | AMITA BUDHE | 8275397196 | 31631 | 22-Aug-2019 | 09:02 |  |  | 23-Aug-2019 | 11:23 | NOT WORKING | CHECK AND FOUND BLADDER AND BULB IS LEAK. CHANGED THE OLD DEFECTIVE BLADDER AND BULB WITH THE NEW ONE. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 26:21:0 | 2
1662 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00112 | DIGITAL BP APPARATUS | DR MASKAR | 9860051432 | 18539 | 28-Jul-2018 | 08:00 |  |  | 31-Jul-2018 | 17:00 | NOT WORKING | CHECK AND FOUND ERROR IS SHOWING WHILE INFLATING THE CUFF. B.P.CUFF IS FOUND LEAK. REPLACED WITH NEW ONE. NOW MACHINE IS WORKING. | Closed | NO INFO |  | 81:0:0 | 4
1663 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00112 | DIGITAL BP APPARATUS | SUSHMA BANSOD | 8623033907 | 19646 | 27-Aug-2018 | 09:05 |  |  | 28-Aug-2018 | 11:57 | NOT WORKING | CHECK AND FOUND BP CUFF WAS LEAK CHANGE THE BP CUFF WITH NEW ONE .NOW MACHINE IS WORKING | Closed | NO INFO |  | 26:52:0 | 2
1664 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00112 | DIGITAL BP APPARATUS | Archna gadge | 9764887196 | 44066 | 11-Jul-2020 | 08:52 |  |  | 11-Jul-2020 | 17:00 | NOT WORKING | CHECKED AND FOUND EXTENSION TUBE IS LEAK THEN CHANGE THE NEW EXTENSION TUBE NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 8:8:0 | 1
1665 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00113 | DIGITAL BP APPARATUS | SHUSHMA BANSOD | 8623033907 | 18594 | 30-Jul-2018 | 09:15 |  |  | 06-Aug-2018 | 10:11 | NOT WORKING | CHECK AND FOUND BP APPARTUS BP CUFF WAS LEAK ,BP CUFF WAS REPLACE WITH NEW ONE, NOW MACHINE IS WORKING GOOD | Closed | NO INFO |  | 168:56:0 | 8
1666 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00113 | DIGITAL BP APPARATUS | SUSHMA NURSE | 8623033907 | 23770 | 09-Jan-2019 | 08:35 |  |  | 15-Jan-2019 | 13:31 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS LEAK. REPLACED THE NEW NIBP CUFF. B. P APPARATUS IS NOW WORKING | Closed | NO INFO |  | 148:56:0 | 7
1667 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00113 | DIGITAL BP APPARATUS | SUSHMA BANSODE | 8623033907 | 24221 | 23-Jan-2019 | 08:37 |  |  | 24-Jan-2019 | 17:22 | NOT WORKING | CHECK AND FOUND POWER ADAPTER IS BROKEN. REPLACED THE OLD WITH THE NEW. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 32:45:0 | 2
1668 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00119 | DIGITAL BP APPARATUS | SUSHMA NURSE | 8623033907 | 23771 | 09-Jan-2019 | 08:35 |  |  | 15-Jan-2019 | 17:34 | NOT WORKING | CHECK AND FOUND NIBP CUFF IS LEAK. REPLACED THE NEW NIBP CUFF. B. P APPARATUS IS NOW WORKING. | Closed | NO INFO |  | 152:59:0 | 7
1669 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00119 | DIGITAL BP APPARATUS | DR DHAKITE | 8275324181 | 13512 | 07-Mar-2018 | 08:22 |  |  | 08-Mar-2018 | 15:46 | NOT WORKING | REPLACED NEW BLADDER, CUFF, BULB WITH CONNECTOR. NOW MACHINE IS WORKING CONDITION. | Closed | NO INFO |  | 31:24:0 | 2
1670 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00119 | DIGITAL BP APPARATUS | SHUSHMA BANSOD | 8623033907 | 18593 | 30-Jul-2018 | 09:15 |  |  | 31-Jul-2018 | 17:00 | NOT WORKING | CHECK AND FOUND LEAKAGE IN B.P. CUFF CHANGE WITH NEW ONE. NOW B.P. APRATUS WORKING SATISFACTORY. | Closed | NO INFO |  | 31:45:0 | 2
1671 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00119 | DIGITAL BP APPARATUS | SUSHMA BANSOD | 8623033907 | 19645 | 27-Aug-2018 | 09:05 |  |  | 28-Aug-2018 | 11:55 | NOT WORKING | CHECK AND FOUND BP CUFF IS LEAK CHANGE THE BP CUFF SAME NOW MACHINE IS WORKING | Closed | NO INFO |  | 26:50:0 | 2
1672 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00120 | NEBULIZER ELECTRIC | SHALINI SAMUNDARE | 9922869996 | 17877 | 13-Jul-2018 | 09:12 |  |  | 17-Jul-2018 | 17:50 | NOT WORKING | CHECK THE MACHINE PROBLEM REPORTED IN MOTOR, DUE TO RUST. SO THE MOTOR IS NOT MOVING PROPERLY. CLEAN THE RUST AND PUT LUBRICANT IN REQUIRED PARTS AND CHANGE THE TUBINGS . NOW MACHINE WORKING FINE . HANDOVER TO USER. | Closed | NO INFO |  | 104:38:0 | 5
1673 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00126 | Fumigation Machine | AMITA | 8275397196 | 24099 | 19-Jan-2019 | 08:22 |  |  | 21-Jan-2019 | 15:22 | NOT WORKING | CHECK AND FOUND POWER SUPPLY CORD IS DEFECTIVE. CHANGED THE NEW POWER CORD. FUMIGATION MACHINE IS NOW WORKING. | Closed | NO INFO |  | 55:0:0 | 3
1674 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00126 | Fumigation Machine | AMITA | 8275397196 | 40306 | 02-Mar-2020 | 08:16 |  |  | 03-Mar-2020 | 10:00 | NOT WORKING | CHECK AND FOUND CHEMICAL COULD NOT BE FUMIGATED SINCE FUMIGATING NOZZLE IS BLOCKED AND ALSO THE INTERNAL TUBING. NOZZLE AND INTERNAL TUBING IS CLEANED. FUMIGATION MACHINE IS NOW WORKING. | Closed | NO INFO |  | 25:44:0 | 2
1675 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00126 | Fumigation Machine | ANITA | 8275397196 | 27410 | 08-May-2019 | 08:50 |  |  | 09-May-2019 | 09:47 | NOT WORKING | CHECK AND FOUND CHEMICAL COULD NOT BE FUMIGATED SINCE FUMIGATING NOZZLE IS BLOCKED AND ALSO THE INTERNAL TUBING. NOZZLE AND INTERNAL TUBING IS CLEANED. FUMIGATION MACHINE IS NOW WORKING. | Closed | NO INFO |  | 24:57:0 | 2
1676 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00134 | BP APPARATUS ANEROID | MANMODE | 7020300118 | 31796 | 26-Aug-2019 | 08:18 |  |  | 27-Aug-2019 | 03:35 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND BULB IS DAMAGE. CHANGED THE NEW BULB. NOW MACHINE IS IN WORKING CONDITION. HANDED OVER TO THE USER. | Closed | CM/L-01986238 |  | 19:17:0 | 2
1677 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00174 | SUCTION PUMP PORTABLE, ELECTRONIC | SHEETAL | 8668717454 | 37741 | 10-Jan-2020 | 08:08 |  |  | 10-Jan-2020 | 11:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN POWER CABLE IS FAULTY THEN CHANGE THE NEW POWER CABLE NOW MACHINE IS IN WORKING CONDITION. | Closed | NO INFO |  | 2:52:0 | 1
1678 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00174 | SUCTION PUMP PORTABLE, ELECTRONIC | DHANIKAR NANDESHWAR | 8605883470 | 32969 | 24-Sep-2019 | 09:06 |  |  | 24-Sep-2019 | 17:35 | NOT WORKING | CHECK AND FOUND PRESSURE IS LOW. FILLED THE HYDRAULIC OIL IN THE COMPRESSOR. NOW DESIRED PRESSURE IS GENERATED. SUCTION MAHCHINE IS NOW WORKING. | Closed | NO INFO |  | 8:29:0 | 1
1679 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00189 | BP APPARATUS ANEROID | SHEETAL | 8668717454 | 37742 | 10-Jan-2020 | 08:08 |  |  | 10-Jan-2020 | 11:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER AND CUFF. THEN CONNECT THE NEW BLADDER AND CUFF NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | 51085 |  | 2:52:0 | 1
1680 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00190 | BP APPARATUS ANEROID | SHEETAL | 8668717454 | 37743 | 10-Jan-2020 | 08:08 |  |  | 10-Jan-2020 | 11:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN BLADDER AND CUFF. THEN CONNECT THE NEW BLADDER AND CUFF NOW MACHINE IS IN WORKING CONDITION HAND OVER TO USER. | Closed | 50733 |  | 2:52:0 | 1
1681 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00194 | MONITOR | SHEETAL | 8668717454 | 37429 | 02-Jan-2020 | 08:43 |  |  | 03-Jan-2020 | 12:00 | NOT WORKING | CHECKED THE EQUIPMENT AND FOUND PROBLEM IN DISPLAY IS FAULTY. RAISED CALL AGAIN WHEN MACHINE IS REPAIRED. | Closed | NO INFO |  | 27:17:0 | 2
1682 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00207 | BP APPARATUS ANEROID | Amol D Sapate | 9284465389 | 44248 | 15-Jul-2020 | 09:14 |  |  | 16-Jul-2020 | 09:00 | NOT WORKING | CHECKED AND FOUND BLADDER, CUFF AND VALVE IS DAMAGED THEN CHANGE THE NEW BLADDER,CUFF AND VALVE NOW MACHINE IS IN WORKING CONDITION. | Closed | RM0104031 |  | 23:46:0 | 2
1683 | BHANDARA | SDH TUMSAR | MH/R1/BHA/SDH/TMR/00209 | ECG UNIT 3 CHANNEL PORTABLE | Amol D Sapate | 9284465389 | 44249 | 15-Jul-2020 | 09:14 |  |  | 16-Jul-2020 | 09:00 | NOT WORKING | CHECKED AND FOUND POWER CABLE IS CUT AND PRINTER ERROR IS OCCURE THEN MAKE PROPER POWER CABLE CONNECTION AND CLEAN PRINTER ROLLER NOW MACHINE IS IN GOOD WORKING CONDITION | Closed | NO INFO |  | 23:46:0 | 2